## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | VENTILATOR TRANSPORT ADULT & PEDIATRIC
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | CART CRASH
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | CYLINDER O2 SIZE E WITH REGULATOR
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | REGULATOR SUCTION HIGH WALL MOUNT
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | STAND IV
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | DEFIBRILLATOR UNIT
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | ECG 12 LEAD WITH CART
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | MATTRESS AIR MED RISK
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | NEBULIZER ULTRASONIC
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | OXIMETER PULSE HAND HELD
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | PUMP INFUSION
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | PUMP SYRINGE
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | PURIFIER AIR SMALL AREA
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | THERMOMETER INFRARED
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | SUCTION UNIT PORTABLE
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | CRRT UNIT
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | DIALYSIS HAEMOFILTRATION
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | ANALYZER BLOOD GAS BENHCTOP
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | DOCKING PUMP INFUSION 4
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | VENTILATOR CPAP ADULT
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | VENTILATOR ICU ADULT & PEDIATRIC
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | ULTRASOUND POC GENERAL PORTABLE
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | HIGH FLOW NASAL CANNULA : HFNC
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | SURFACE DECONTAMINATION SYSTEM H2O2
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | PAPR RESPIRATOR BIOLOGICAL COMPLETE
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | LAMP DECONTAMINATION ULTRAVIOLET
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | CAMERA THERMAL HUMAN FIXED
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | CAMERA THERMAL HUMAN PORTABLE
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | X-RAY DIGITAL MOBILE
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | HYPERBARIC OXYGEN THERAPY
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | INTUBATION DIFFICULT VIDEO
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | LARYNGOSCOPE SET COMPLETE
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | ECMO MACHINE
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | BED ELECTRICAL
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | BED INTENSIVE CARE
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | STRETCHER TRANSPORT
ITEM 42 | This sheet contains the details of the bid with each technical specification for | ITEM 42 | VENTILATOR BIPAP IN-PATIENT

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | MAL20015 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | Medical Eqpt | 11 | [10 empty] | 0
2 | MGE10009 | CART CRASH |  | Medical Eqpt | 5 | [10 empty] | 0
3 | MGE20007 | CYLINDER O2 SIZE E WITH REGULATOR |  | Medical Eqpt | 50 | [10 empty] | 0
4 | MGE20014 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | Medical Eqpt | 150 | [10 empty] | 0
5 | MGE20020 | REGULATOR SUCTION HIGH WALL MOUNT |  | Medical Eqpt | 50 | [10 empty] | 0
6 | MGE20022 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER |  | Medical Eqpt | 50 | [10 empty] | 0
7 | MGE40004 | STAND IV |  | Medical Eqpt | 250 | [10 empty] | 0
8 | MGE50001 | DEFIBRILLATOR UNIT |  | Medical Eqpt | 5 | [10 empty] | 0
9 | MGE50005 | ECG 12 LEAD WITH CART |  | Medical Eqpt | 13 | [10 empty] | 0
10 | MGE50017 | MATTRESS AIR MED RISK |  | Medical Eqpt | 50 | [10 empty] | 0
11 | MGE50020 | NEBULIZER ULTRASONIC |  | Medical Eqpt | 27 | [10 empty] | 0
12 | MGE50024 | OXIMETER PULSE HAND HELD |  | Medical Eqpt | 107 | [10 empty] | 0
13 | MGE50027 | PUMP INFUSION |  | Medical Eqpt | 31 | [10 empty] | 0
14 | MGE50029 | PUMP SYRINGE |  | Medical Eqpt | 120 | [10 empty] | 0
15 | MGE50031 | PURIFIER AIR SMALL AREA |  | Medical Eqpt | 74 | [10 empty] | 0
16 | MGE50037 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | Medical Eqpt | 50 | [10 empty] | 0
17 | MGE50045 | THERMOMETER INFRARED |  | Medical Eqpt | 144 | [10 empty] | 0
18 | MGE50055 | SUCTION UNIT PORTABLE |  | Medical Eqpt | 39 | [10 empty] | 0
19 | MHD10001 | CRRT UNIT |  | Medical Eqpt | 5 | [10 empty] | 0
20 | MHD10004 | DIALYSIS HAEMOFILTRATION |  | Medical Eqpt | 5 | [10 empty] | 0
21 | MHD10005 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO |  | Medical Eqpt | 5 | [10 empty] | 0
22 | MIC10002 | ANALYZER BLOOD GAS BENHCTOP |  | Medical Eqpt | 5 | [10 empty] | 0
23 | MIC10025 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH |  | Medical Eqpt | 181 | [10 empty] | 0
24 | MIC10039 | DOCKING PUMP INFUSION 4 |  | Medical Eqpt | 100 | [10 empty] | 0
25 | MIC10044 | VENTILATOR CPAP ADULT |  | Medical Eqpt | 2 | [10 empty] | 0
26 | MIC10046 | VENTILATOR ICU ADULT & PEDIATRIC |  | Medical Eqpt | 131 | [10 empty] | 0
27 | MIC10055 | ULTRASOUND POC GENERAL PORTABLE |  | Medical Eqpt | 1 | [10 empty] | 0
28 | MIC10064 | HIGH FLOW NASAL CANNULA : HFNC |  | Medical Eqpt | 508 | [10 empty] | 0
29 | MIF40004 | SURFACE DECONTAMINATION SYSTEM H2O2 |  | Medical Eqpt | 21 | [10 empty] | 0
30 | MIF50016 | PAPR RESPIRATOR BIOLOGICAL COMPLETE |  | Medical Eqpt | 500 | [10 empty] | 0
31 | MIF50017 | LAMP DECONTAMINATION ULTRAVIOLET |  | Medical Eqpt | 4 | [10 empty] | 0
32 | MIF50018 | CAMERA THERMAL HUMAN FIXED |  | Medical Eqpt | 24 | [10 empty] | 0
33 | MIF50019 | CAMERA THERMAL HUMAN PORTABLE |  | Medical Eqpt | 4 | [10 empty] | 0
34 | MMI110012 | X-RAY DIGITAL MOBILE |  | Medical Eqpt | 1 | [10 empty] | 0
35 | MOP100002 | HYPERBARIC OXYGEN THERAPY |  | Medical Eqpt | 1 | [10 empty] | 0
36 | MOR20017 | INTUBATION DIFFICULT VIDEO |  | Medical Eqpt | 10 | [10 empty] | 0
37 | MOR20019 | LARYNGOSCOPE SET COMPLETE |  | Medical Eqpt | 50 | [10 empty] | 0
38 | MOR30006 | ECMO MACHINE |  | Medical Eqpt | 2 | [10 empty] | 0
39 | MWD10003 | BED ELECTRICAL |  | Medical Eqpt | 140 | [10 empty] | 0
40 | MWD10006 | BED INTENSIVE CARE |  | Medical Eqpt | 20 | [10 empty] | 0
41 | MWD10023 | STRETCHER TRANSPORT |  | Medical Eqpt | 20 | [10 empty] | 0
42 | MWD10034 | VENTILATOR BIPAP IN-PATIENT |  | Medical Eqpt | 100 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | MAL20015 | 0 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PATIENT TYPE | Adult, Pediatric
2 | Built-in Turbine ( Compressor ) | Yes
3 | CONTROLS
4 | Monitoring | , Ventilation time ratio, Gas consumption, Battery capacity.
5 | Monitoring | Total expiratory minute volume, Spontaneous minute volume, Peak pressure, Mean pressure, Plateau pressure, Positive end expiratory pressure, Expiratory tidal volume, Inspiration time, Frequency, Spontaneous breathing frequency, Oxygen concentration
6 | Tidal volume, L | 50-2000 mL
7 | Breath rate, br/min | 0 - 60 or more
8 | Inspiratory time, sec | 0.2-10 SEC
9 | Inspiratory flow, L/min | 60 up to 100
10 | Pressure level, cm H2O | 0-60 CM H2O
11 | Pressure ramp | Specify
12 | Pressure support | 0 - 30 CM H20
13 | Trigger Mechanism | Flow or pressure or both, Specify.
14 | Sensitivity, cm H2O | Flow or pressure trigger, specify
15 | FiO2, Percentage | 21 % - 100 %
16 | I:E ratio | 1:4 to 3:1
17 | Adjustable PEEP either Internal or External, specify | 0-20 cm H2O
18 | OPERATING MODES
19 | Intermittent Positive Pressure Ventilation (VC-CMV) | Yes
20 | Assisted Intermittent Positive Pressure Ventilation (VC-AC) | Yes
21 | Synchronized Intermittent Mandatory Ventilation (VC-SIMV) | Yes
22 | Spontaneous ventilation (SpnCPAP) | Yes
23 | Pressure Support to assist spontaneously breathing patients (SpnCPAP+PS) | Yes
24 | Mask ventilation with leakage compensation in the modes SpnCPAP (SpnCPAP+PS)(NIV) & BIPAP | Yes
25 | Display Type | Colored , Built-in with the unit
26 | MONITOR PARAMETERS
27 | Pressure, PIP, MAP, PEEP. | Yes
28 | Volume, Tidal, Minute | Yes
29 | Inspiratory/Expiratory Time | Yes
30 | I:E Ratio | Yes
31 | Respiratory Rate | Yes
32 | PATIENT ALARMS
33 | FiO2, Low/High minute volume, Low Inspiratory pressure, High pressure, Loss of PEEP, Apnea, Inverse IE, High continuous pressure occlusion, High respiratory rate, Others. | Specify
34 | EQUIPMENT ALARMS
35 | Gas-supply failure, Power failure, Vent inoperative, Low battery, Self Diagnostic, Others. | Specify
36 | Silence | 2 minutes
37 | CONTROL TYPES | Dial, soft keys, toggle switch
38 | POWER, VAC | 220, 60Hz.
39 | Breathing circuit Adult & Pediatric | 3 set
40 | BUILT-IN BATTERY WITH CHARGER | 2 HOURS AT LEAST
41 | THE ABOVE UNIT IS COMPATIBLE WITH AMBULANCE CAR | Yes
42 | ALL ACCESSORIES NEEDED ,INCLUDING PATIENT BREATHING CIRCUIT LOUPES & FILTERS with each VENTILATOR | Yes
43 | CERTIFICATES | APPROVED FROM REGULATORY BODY
44 | LANGUAGE | ENGLISH
45 | PLUGS | BRITISH STANDARD
46 | UNIT OF MEASUREMENT | SI : INTERNATIONAL STANDARD
47 | MOBILE STAND | INCLUDED ON CASTORS WITH BASKET
48 | BED RAIL HANGER | INCLUDED
49 | HANDLE | YES
50 | WEIGHT | 7 Kg APPROX
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | MGE10009 | 0 | CART CRASH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CRASH CART
2 | CARDIO-RESPIRATORY RESUSCITATION SYSTEM | Mobile cart with lockable drawers and defibrillator transport tray ( bracket )
3 | MATERIAL | HEAVY DUTY QULAITY TYPE
4 | ACCESSORIES
5 | Oxygen supply sys. | Yes, 1 cylinder
6 | Suction unit portable, battery. | Yes
7 | Oxygen flow meter | Yes
8 | Laryngoscope blades set | Yes, Adult/Pedia/Neonate
9 | Resuscitation bag | Yes with mask for adult/pedia/neonate
10 | ALL ACCESS NEEDED | SPECIFY
11 | Tourniquet | Yes, Velcro fastener
12 | Forceps | Yes, 2 sets
13 | Magill | Yes
14 | Artery | Yes
15 | Dressing | Yes
16 | Suture | Yes
17 | Penlight | Yes
18 | Sphygmomanometer ( aneroid ) | Yes, with cuffs, (adult, & pedia.)
19 | Infusion pole | Yes, Telescopic
20 | CART
21 | Wheels | 4 anti-static wheel, > 12.5 cm dia.
22 | Brakes | > 2
23 | Waste container | Yes
24 | Heart board | Yes
25 | Oxygen tank holder | Yes
26 | Medication tray | Yes
27 | Infusion pole receptacle | Yes
28 | Push handle | Yes
29 | Bumpers | Yes
30 | Top guard rail | Yes
31 | Drawers | Yes
32 | OTHER SPECIFICATION | Work surface and size shall accommodate the Defibrillator and all accessories mentioned.
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | MGE20007 | 0 | CYLINDER O2 SIZE E WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | OXYGEN CYLINDER E-SIZE | YES
2 | ALUMINUM CYLINDER | YES
3 | Nominal contents (litres) | 680L
4 | PIN INDEX | YES
5 | CAPACITY | SPECIFY
6 | ALUMINUM TROLLEY | INCLUDED
7 | OXYGEN REGULATOR | INCLUDED WITH FLOW METER 15 L/Min
8 | OTHER SPECIFICATION | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | MGE20014 | 0 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Flow Meter | Wall Mounted
2 | Flow Range | 0 - 15 L / Min
3 | Color Coded | specify
4 | Pressure Compensated | Yes
5 | Glass Envelop | Included with indicator ball
6 | Metal Nipple | Yes
7 | Probe | As Per Site
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | MGE20020 | 0 | REGULATOR SUCTION HIGH WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE(S) OF SUCTION | All common procedures and GI
2 | MOUNTING | Wall mounted
3 | FUNCTION | Continuous and intermittent
4 | ADJUSTABLE CYCLE | Independent off/on
5 | Increments | 3 - 30 sec
6 | VACUUM ADJUST LOCK | Yes
7 | VACUUM GAUGE | Digital or analog
8 | Pin stop | Yes
9 | Range, mm Hg | 0 - 300, full line
10 | Graduations, mm Hg | Five
11 | Disinfection | Specify
12 | Sterilization | Specify
13 | Safty trape | Yes
14 | Probe as per site | Yes
15 | Collection bottle | Disposable System
16 | Material of bottle | Poly carbonate
17 | OTHER SPECIFICATIONS | Color-coded gauge
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | MGE20022 | 0 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CYLINDER SIZE H | FOR MEDICAL OXYGEN
2 | CONSTRUCTION | STEEL
3 | CAPACITY | 7,700 LITERS APPROX
4 | REGULATORS / VALVE | CGA 540
5 | FLOW METER 15L/MIN | INCLUDED WITH CYLINDER
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | MGE40004 | 0 | STAND IV |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FRAME | Stainless steel 304 preferred
2 | TYPE | Telescopic
3 | HOOKS | 4 - 5 Approx
4 | HEIGHT | Adjustable
5 | BASE | Mobile heavy base to withstand the hanged Items weights
6 | Castors | Qty : 5
7 | CAPACITY | 5 kg per hook
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | MGE50001 | 0 | DEFIBRILLATOR UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DEFIBRILLATOR ENERGY SELECTION, Joules. | Yes
2 | Internal with Lead | 5-50 Joules
3 | External Biphasic , Adult | 2-200 Joules
4 | External Biphasic , Pediatric / Neonatal | 2-200Joules
5 | PADDLE CONTROLS | Charge, Discharge
6 | Synchronizer | Yes
7 | Pediatric Paddles / Neonatal | Yes
8 | Disposable PADDLE | Yes , 25 pairs , validity two years
9 | ECG Monitor | Yes
10 | Type | LCD colored
11 | Screen display size, | 5 to 7 Inch
12 | Sweep Speed, mm/sec. | 20-25, Other specify.
13 | Lead Configuration | Yes 3 - 5 -12 leads
14 | THROUGH THE PADDLE MONITORING | Yes
15 | HR Display | Yes
16 | SPO2 | Yes
17 | HR Alarms | Yes
18 | Frequency Response, Hz. | Specify
19 | Lead fault indicator | Yes
20 | EXTERNAL PACEMAKER | YES
21 | Pacing Mode | Demand, fixed rate, Specify.
22 | Pacing Rate, ppm | 50-150, Specify
23 | Output current, mA | 0-140 or wider
24 | Pulse width, m sec. | > 20
25 | ECG RECORDER | Yes
26 | Paper Speed, mm/sec. | 25, Others specify.
27 | Auto/Manual Print | Specify
28 | Annotation | Time, date, lead, gain, heart rate, operating mode, other specify.
29 | BATTERY/LINE POWER OPERATION | Both
30 | Integral or Removable | Specify
31 | Operating Time, Hr. | > OR = Two Hours continuous ECG monitoring - OR > / = 20 discharges.
32 | Charging Time, Hr. | Specify
33 | TO BE PLACED ON CRASH CART | YES
34 | AC POWER | 220Volt - 60Hz
35 | OTHER SPECIFICATIONS | FDA or CE, ISO approved.
36 | Options Itemized Prices | Yes
37 | CPR quality monitoring | Yes
38 | Trace freeze | specify
39 | CO2 MODULE | YES
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | MGE50005 | 0 | ECG 12 LEAD WITH CART |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile
2 | Cart | Yes
3 | LEADS
4 | Lead switching | Automatic/Manual
5 | Sensitivity, mm/mV | 5, 10, 20., other specify.
6 | Calibration signal | Automatic
7 | FREQUENCY RANGE, HZ
8 | Diagnostic | 0.05-150
9 | Filtered | Notch, EMG, Baseline, Wander, filters
10 | INPUT IMPEDANCE,
11 | Mega ohms | one hundred
12 | CMRR @ 60 Hz, dB | 100-140 approx
13 | Leads-off indicator | Yes
14 | RECORDER
15 | Channels | twelve
16 | Recording | YES A4 PAPER
17 | Lead marker | Automatic
18 | Timing marker | Yes
19 | Event marker | Yes
20 | Chart speed, mm/sec | 5, 10, 25, 50.
21 | LCD MONITOR 10 OR MORE | Yes ADJSTABLE
22 | No. of traces | 3, 6, 12
23 | PREVIEW SCREEN | Yes
24 | No. Waveforms stored | Min. 30
25 | ECG transmission | Yes
26 | Interpretation | Yes
27 | ECG measurements | Yes
28 | Auxiliary output | Yes
29 | Auxiliary input | Yes
30 | DEFIBRILLATOR SYRORNAIZ | YES
31 | SINGLE AVREGING | YSE
32 | Overload protection | Yes
33 | BATTERY OPERATION | Yes
34 | Battery type | Built in rechargeable
35 | Operating time, hr | Specify
36 | ACCESSORIES | All standard accessories
37 | PRINTER | Built-in
38 | A4 paper | Yes
39 | SAFTY AND PERFORMANCE | IEC 60601-2-51
40 | POWER REQUIREMENTS | 220V, 60Hz
41 | OTHER SPECIFICATIONS | Spare patient cable, chest electrodes, clamp electrodes for limb leads, with ICU Analysis. SOFT WARE Patient cable holder.FDA, CE, ISO, Approved. HL7
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | MGE50017 | 0 | MATTRESS AIR MED RISK |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA approved and/or CE Marked | Yes, Please specify
2 | Type | Powered Air Mattress
3 | Indications | Medium Risk pressure relief therapy surface (Prevention and Treatment)
4 | Bed Type | Adult Bed
5 | Alternating Therapy | Yes
6 | Cycle Time | Yes Specify
7 | Static therapy mode | Yes
8 | Adjustable wieght capcity | Yes
9 | Integrated Pressure sensor | Yes, which monitors internal pressure 24 hours.
10 | Minimum of 16 air cylinders individual replaceable cells | Yes, please specify
11 | Mattress Dimensions (LxWxD) cm | 200 x 90 x 18 (Inflated)
12 | Cover Material , Water Proof , Bacterial Barrier , Stain Resistant , Air breathable& moist , Soft for friction elmination , Shear resistant , Washable | yes
13 | Maximum weight capacity | 150Kg
14 | Manual CPR function from the mattress | Yes
15 | Transport mode | Yes , specify
16 | Alarms , Visual , Audiable | Yes
17 | Low Pressure Alert | Yes
18 | Voltage | 220 – 240 V 50/60 Hz
19 | Flexible Hangers in the rear side | Yes
20 | Whisper quite | Yes
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | MGE50020 | 0 | NEBULIZER ULTRASONIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | NEBULIZER
2 | Warm-up time, min
3 | Mist temperature, 0C | 33-35 aprox.
4 | Mist output, mL/min | 0.2-3.15
5 | Noise level, dbA,
6 | max @ 1 m distance
7 | O2 conc settings, % | 40-100 % approx.
8 | HEATING UNIT
9 | Location | Tube
10 | Type | Processor controlled/heater block
11 | Power-on indicator | Yes
12 | Temperature control | Yes
13 | Safety shutoff | Yes
14 | Line voltage, VAC | 220-240 V, 60 Hz
15 | RESERVOIR
16 | Refillable | YES
17 | Usable volume, mL | >200
18 | Water ml | >200
19 | Saline (0.45/0.9%) | >150
20 | STERILIZATION METHOD
21 | Nebulizer head and Heating unit | Chemical and autoclaveable
22 | Reservoir | Chemical and autoclaveable
23 | OTHER SPECIFICATIONS | ALL ACCESS. NEEDED TABLE TOP
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | MGE50024 | 0 | OXIMETER PULSE HAND HELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Portable / Stand-alone/transport/HAND HELD FOR ADULT/PEDIATRIC/INFANT
2 | DISPLAYS | SpO2, pulse rate, pulse strength and/or signal, low battery, Volumes, alarm limits.
3 | Type | LCD/ with backlit LED SHOWING PARAMETERS & READINGS
4 | Sp02 RANGE, % | 0-100
5 | Accuracy, ,% | < ± 2 %
6 | PULSE RATE, bpm | 30-240
7 | PERFUSION INDEX AND/OR SIGNAL STRENGTH INDICATOR | YES
8 | RESPONSE TIME, sec | Pulse to Pulse or time ( Specify)
9 | SETTLING TIME | Pulses or time ( Specify)
10 | START-UP TIME, sec or pulses
11 | ALARMS:
12 | Audiovisual | Yes
13 | Visual | Yes, Low battery
14 | ALARM OVERRIDE | Yes
15 | Reactivation method | Yes, after 2 mins.
16 | Volume control | Yes
17 | SELF-TEST MODE | Yes
18 | Memory: up to 72 hours of data storage | YES
19 | PROBE TYPES | YES REUSABLE (1 ADULT, 1 PEDIATRIC, AND 1 NEONATE)
20 | Patient range | Adult / Neonate/ Pediatric
21 | Cable length, m | Not less than1.5 m
22 | BATTERY BACKUP :
23 | Rechargeable battery | Yes, Specify the type
24 | Low battery notice | Yes
25 | Rechargeable time, hr | Specify
26 | Battery life, hrs | Specify
27 | The unit should operate on batteries as well as AC electricity, with automatic battery charging while in use. | YES
28 | The unit shall operate on batteries | YES
29 | battery type | Specify
30 | VAC | 220 V
31 | CURRENT | 13 A
32 | FREQUENCY | 60 Hz.
33 | PLUG TYPE | 3 Pin British
34 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
35 | IMPORTANT NOTE: | Please specify whether accessories are standard or optional, otherwise they will be considered standard.
36 | MDMA | Yes
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | MGE50027 | 0 | PUMP INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PROGRAMMING PATINET ID / NURSE ID ON THE PUMP DURING ADMINISTRATION | YES ( PLEASE SPECIFY )
2 | QUICK START OF INFUSION | YES
3 | FLOW RANGE ML/HR | 0.1 - 999 ML/HR
4 | ACCURACY | 0.05
5 | DRUG/DOSE CALC | YES
6 | VTBI, ML | 0.1 TO 9999
7 | MAX PRESSURE, PSI | SPECIFY MAX
8 | PUMPING MECHANISM | SPECIFY MECHANISM
9 | IV SET | YES
10 | FREE-FLOW PROTECTION | YES
11 | MANUAL AND PROGRAMMABLE BOLUS | PLEASE SPECIFY BOLUS RATE
12 | ALARMS & INDICATORS
13 | EMPTY BAG | YES
14 | FLOW ERROR | YES
15 | DOOR OPEN | YES
16 | CONTINUOUS PRESSURE MONITORING PSI (ON THE FRONT SCREEN ALL THE TIME) | SPECIFY
17 | DRUG DOSE CALCULATION | ML/H, µG/H, µG/MIN, µG/KG/H, µG/KG/MIN, MG/H, MG/MIN,MG/KG/MIN, MG/KG/H, IU/H, IU/MIN, IU/KG/H, IU/KG/MIN.
18 | FLEXIBLE PRESSURE ALARM SETTINGS | IN LINE PRESSER
19 | HIGH PRESSURE/OCCLUSION | YES
20 | PRE- OCCLUSIONALARM | YES
21 | CIRCUIT MALFUNCTION | YES
22 | INFUSION NEAR END | YES
23 | INFUSION COMPLETE | YES
24 | LOW BATTERY | YES
25 | DEPLETED BATTERY | YES
26 | USER PROMPT | YES
27 | AUDIBLE ALARM | YES
28 | VOLUME CONTROL | YES
29 | MOMENTARY SILENCE | YES
30 | PPM DUE DATA REMINDER | YES
31 | FLUID RESISTANT | YES
32 | NURSE CALL JACK | YES
33 | DATA LOG | YES
34 | EVENT LOG | YES
35 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED)
36 | NUMBER OF EVENTS | 1500
37 | PRINT OUT | YES
38 | DISPLAY | LED AND LCD WITH COLOR CABLITIES IS A PLUS & BIG SCREEN INTERFACE IS ALSO A PLUS
39 | INFUSED VOLUME DISPLAY | YES , DIGITAL
40 | WEIGHT WITH BATTERY KG | SPECIFY WIGHT
41 | TIME /DATE | YES
42 | POWER SOURCE, VAC | 220VOLT / 60 HZ.
43 | POWER SUPPLY | INTERNAL OR EXTERNAL (IF EXTERNAL MUST BE LIFETIME WARRANTY)
44 | POWER CABLE | DURABLE POWER CABLE
45 | BATTERY | YES, RECHARGEABLE BUILT IN
46 | BATTERY | NI-MH
47 | OPERTING TIME , HR | SPECIFY HOURS @ 10 ML/HR
48 | RECHARGE TIME , HR | PREFERRED LOWER CHARGING HOURS
49 | POLE MOUNTING | PREFERRED
50 | OTHER SPECIFICATIONS | FDA / CE
51 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED).
52 | STACKABLE IN THE SAME DOCKING STATION | INFUSION PUMPS / SYRINGE PUMPS / ENTRAL FEEDING PUMPS.
53 | DOSE ERROR REDUCTION SYSTEM (SMART TECHNOLOGY) | YES
54 | LIBRARY SIZE | YES CONFIGURABLE
55 | NO. OF DRUG ENTITIES/CARE AREA | YES PLEASE SPECIFY
56 | WIRELESS CONNECTIVITY | UPGRADABLE
57 | TIME RETAINED | FIRST IN FIRST OUT
58 | ACCESSORIES FOR THE PUMP | CLOSED SYSTEM
59 | ALL SETS HAVE SELF -SEALING NEEDLE FREE INJECTION SITES | YES
60 | ALL SETS SHOULD BE LATEX DEHP & PVC FREE | YES
61 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
62 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
63 | ACCESSORIES FOR INFUSION PUMP | CLOSE SYSTEM
64 | THE VENDOR MUST PROVIDE A FIXED PRICE FOR THE ALL ABOVE MENTIONED ACCESSORIES VALID FOR 3 YEARS | YES
65 | PUMP HL7 READY | YES ( PLEASE SPECIFY )
66 | PROVIDE CONSUMABLE DEAL ( PUMPS FREE AND PRICE ON SETS ) | OPTIONAL DEAL SECTOR MIGHT CHOOSE
67 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
68 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
69 | THE WORKSTATION SHOULD BE MOUNTABLE , WALL UNIT (FIXATION TO WALL RAILS OF DIFFERENT MANUFACTURER).MOVABLE POLE-BASED UNIT.
70 | SMART ALARM SYSTEM BY AUDIBLE, VISUAL AND DISPLAYED MESSAGE. | OCCLUSION, END OF INFUSION, EMPTY BATTERY, NO RATE SET, DEFECT ON PUMP. PRE-ALARMS: LOW BATTERY, NEAR-END OF INFUSION , PLEASE SPECIFY
71 | MOBILE STAND TO FIX THE PUMP ON INSTEAD OF PENDANT OR WALL FOR SPECIAL SITE NEEDS | ITEMIZED PRICE OPTIONAL
72 | PRIMARY INFUSION SET (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 120 SETS WITH EACH PUMP ITEMIZED
73 | LIGHT SENSITIVE PRIMARY INFUSION SET WITH NO INJECTION PORT FOR NEONATAL TPN (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 40 SETS WITH EACH PUMP ITEMIZED
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | MGE50029 | 0 | PUMP SYRINGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled with digital LCD alphanumeric display of parameters and alarms
2 | Variable rate ranging from 0.1 to 500 mL/hr or better, with 0.1 mL/hr increments
3 | 3 % accuracy or better
4 | Variable volume-to-be-infused from 1 to 1,000 mL or similar
5 | Digitally displayed parameters to include:
6 | Infusion rate
7 | Battery / AC operation
8 | Running indicator
9 | Alarming condition when active, with indication of alarm type or code
10 | Back pressure monitor / indicator
11 | Event history
12 | Capability to accept different syringe types and sizes with automatic syringe detection and identification
13 | Syringe compatibility and auto detection shall include but not be limited to all sizes of the following (3 to 60 mL):
14 | BD
15 | Terumo
16 | Monoject
17 | Braun
18 | Fresenius
19 | Variable bolus rate
20 | Up to maximum flow rate
21 | Bolus infused volume indicator during bolus activation
22 | Protected access
23 | Audiovisual alarms shall include but not be limited to the following:
24 | Syringe installation and integrity (detection)
25 | Line disconnection (sudden drop in back pressure)
26 | Occlusion pressure pre-alarm
27 | Occlusion pressure
28 | Near end of infusion alarm
29 | End of infusion
30 | Volume limit pre-alarm
31 | Volume limit
32 | KVO (1 ml/hr; if other, specify)
33 | Low battery pre-alarm
34 | Discharged battery
35 | Internal malfunction
36 | Syringe unlocked
37 | Plunger disengaged
38 | Empty syringe
39 | Data log capability and data port for data transmission, display and printing. Any required software for such function shall be included.
40 | Logged data to include ( ≥1 year) :
41 | Settings
42 | Alarms
43 | Errors
44 | Key presses
45 | Error codes
46 | Amount infused
47 | Safety features shall include but not be limited to:
48 | Self test at start-up
49 | Nurse call interfacing capability
50 | Splash proof design
51 | Auto priming
52 | Adjustable alarm volume. No permanent silencing shall be possible.
53 | Keypad lock
54 | Impossibility to improperly install infusion set
55 | Free flow prevention system
56 | Last parameter setting retention
57 | IV stand mounting accessory shall be included
58 | Battery autonomy of 3 hrs or more when fully charged. specify:
59 | Battery type and characteristics (voltage and current capacity)
60 | Autonomy at 10 mL/hr
61 | Recharging time from depleted to 90%
62 | Compliance with standards & legislation:
63 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
64 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
65 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
66 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
67 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | MGE50031 | 0 | PURIFIER AIR SMALL AREA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To provide clean, airborne contaminant free air in individual ICU rooms, isolation rooms
2 | TYPE | Floor / wall mounted , specify
3 | ELECTRONIC CONTROLS | Yes
4 | GRAPHIC DISPLAY | LCD
5 | AIR QUALITY SENSING SYTEM | Yes
6 | FILTER MODULE
7 | Dust filter | Yes
8 | Bacteria filter | Yes
9 | Odour filter | Yes
10 | Germs filter | Yes
11 | Hepa filter | Yes
12 | Carbon filter | Yes
13 | Plasma or UV filtration | Yes
14 | FILTER STATUS INDICATION | Yes
15 | AIR FLOW RATE IN CFM | Variable 350 - 500 APPROX
16 | AIR PURIFIER CLEANING CAPACITY | 40 - 50 m2 approx.
17 | REGULATORY COMPLIANCE | CE, CSA, ISO or any other, specify
18 | SAFETY STANDARD | UL, IEC or any other, specify
19 | POWER SUPPLY | 220 V, 60 Hz
20 | OTHER SPECIFICATION | Vendor is responsible for complete installation at site.
Company Stamp |  |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | MGE50037 | 0 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | for Adult , Pediatric and neonatal .
2 | TYPE | Electronic, Mobile
3 | METHOD OF INFLATION | Oscillometric or any other, specify
4 | MAXIMUM INFLATION PRESSURE | 300 mmHg
5 | AUTO INFLATE & DEFLATE | Yes
6 | CUFF SIZES for Adult , Pediatric and neonatal . | 2 cuffs for each
7 | AUTOMATIC ZERO POSITION | Yes
8 | DISPLAY PARAMETERS | LCD
9 | SYSTOLIC PRESSURE | Yes
10 | DIASTOLIC PRESSURE | Yes
11 | PULSE RATE | Yes
12 | SPO2 | Included
13 | ALARMS, AUDIBLE & VISIBLE | Yes
14 | POWER SUPPLY | 220v, 60 Hz
15 | BATTERY | Yes, Built in re-chargeable batteries with integral charger.
16 | ACCESSORIES | Basket for cuffs
17 | OTHER SPECIFICATION | yes.
18 | SPO2 Probe | For Adult, Pediatric and Neonatal ( Disposable Box )
19 | Temp module | Included
20 | Oral Temp. Probe | Included
21 | Rectal Temp. Probe | Optioanl
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | MGE50045 | 0 | THERMOMETER INFRARED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Non-contact thermometer uses infrared technology to take temperature
2 | Effective distance: 5-8 cm
3 | Range: 35-43˚C
4 | Accuracy: ± 0.2 ˚C
5 | Resolution: 1˚C
6 | Battery operated
7 | Indication for low battery power
8 | Auto recall of last reading
9 | LCD back lit display
10 | Automatic shut off after 10-15 second
11 | Approx. Dimension: 150 H x 40 W x 40 D mm
12 | Approx. Weight: 100 g
13 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | MGE50055 | 0 | SUCTION UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | PORTABLE FOR PATIENTS EXCESSIVE SECRETIONS AND PATIENTS WITH WEEK COUGH OR ARTIFICIAL AIRWAY
2 | TRACHEAL USE: | YES
3 | PREHOSPITAL (EMS) | YES
4 | INTRAHOSPITAL (CRASH CARTS) | YES
5 | PUMP TYPE | SPECIFY
6 | COLLECTION CANISTER | DISPOSABLE
7 | CAPACITY MEASURED, ML | 1,000 ML
8 | GRADUATIONS | YES
9 | OVERFLOW PROTECTION / BACK FLOW PROTECTION | YES
10 | CONNECTOR LABELS | YES
11 | FILTER TYPE | BACTERIAL, HEPA
12 | TUBING:
13 | INTERNAL DIAMETER | 8
14 | LENGTH | 1.8
15 | VACUUM:
16 | RANGE, MM HG | 50 TO 550
17 | TIME TO 300 MM HG, SEC | <3
18 | FLOW RATE AT MAX VACUUM, L/MIN | 27 APPROX
19 | VACUUM-LEVEL CONTROL | REGULATOR, VARIABLE
20 | VACUUM GAUGE | ANALOG / GAUGE INDICATORS / DIGITAL
21 | DIAMETER, CM (IN) | LINEAR LED ARRAY
22 | SCALE, MM HG | 50 TO 550
23 | BATTERY: | YES
24 | TYPE | SLA , NIMH , SPECIFY
25 | LIFE, MIN | 45 TO 60
26 | LOW-BATTERY SIGNAL | YES, LED
27 | CHARGER | INTEGRAL
28 | RECHARGE TIME, HR | ≤3
29 | INDICATOR, BATTERY CHARGING | YES, ON SUCTION UNIT
30 | INDICATOR, BATTERY FULLY CHARGED | YES, ON SUCTION UNIT
31 | WEIGHT, KG (LB) | ≤ 2 APPROX KGS
32 | CAR KIT CHARGER | OPTIONAL ITEMIZED
33 | ASPIRATOR
34 | H X W X D, CM (IN) | SPECIFY
35 | ACCESSORIES | EXTERNAL CHARGER, WIRE CANISTER BRACKET, RECHARGEABLE 12V DC, AC AND DC POWER CORDS, AC/DC ADAPTER CHARGER, DISPOSABLE CANISTERS , SPLASH-PROOF (PROTECTION CLASSS IP12 )
36 | VAC | 220 V 60 HZ
37 | CURRENT | 13 A
38 | PLUG TYPE | 3 PIN BRITISH
39 | UNIT COMPLETE WITH FULL ACCESSORIES | YES
40 | FILTERS | YES 5, FILTERS EXTRA
41 | IMPORTANT NOTE: | PLEASE SPECIFY WHETHER ACCESSORIES ARE STANDARD OR OPTIONAL, OTHERWISE THEY WILL BE CONSIDERED STANDARD.
42 | CERTIFICATE | APPROVED REGULATORY BODY
43 | VENDOR MUST PROVIDE ITEMIZED PRICE FOR ALL SETS AND CANISTER
44 | DISPOSABLE CANISTER / SETS | 300 CANISTER & TUBING
Company Stamp |  |  |  | Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | MHD10001 | 0 | CRRT UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CRRT Machine
2 | Continuous Hemodialysis (CRRT) with the following treatment modalities
3 | SCUF: Slow Continuous Ultra-Filtration | YES
4 | CVVH: Continuous Veno-Venous Hemofiltration | YES
5 | CVVHD: Continuous Veno-Venous Hemodialysis | YES
6 | CVVHDF: Continuous Veno-Venous Hemodiafiltration | YES
7 | HV-CVVH: High Volume Continuous Veno-Venous Hemofiltration | YES
8 | TPE: Therapeutic Plasma Exchange for adult and pediatric applications including consumables availability in KSA market | YES
9 | HP: Hemoperfusion program and consumables ( coated charcoal filters) | YES
10 | Pediatric CRRT / Neonatal CRRT | YES
11 | The Unit should be able to change therapy (between SCUF, CVVH, CVVHD & CVVHDF) during treatment without line disconnection/addition nor changing the set | YES
12 | The Unit should use one treatment set which can perform all CRRT therapies (SCUF, CVVH, CVVHD & CVVHDF) | YES
13 | The Unit should contain at least the following 5 operating & software controlled pumps: | YES
14 | Blood Pump | YES
15 | Substitution Pre Pump | YES
16 | Substitution Post Pump | YES
17 | Dialysate Pump | YES
18 | Drain Pump | YES
19 | The Unit should contain Heparin (syringe) pump with the ability to perform both Continuous & Bolus delivery types | YES
20 | The Unit should have the Flexibility of performing Pre or Post / Pre and Post fluid substitution at the same time with all CRRT therapies (SCUF & CVVH & CVVHD and CVVHDF) | YES
21 | The Unit must also contain the following integrated parts:
22 | Deareation chamber | YES
23 | Air Detector | YES
24 | Clamp system to stop blood return when detecting air. | YES
25 | Blood leak detector. | YES
26 | Pressure pods for measuring:
27 | Arterial line pressure ( -250 to +200 mmHg) | YES
28 | Venous line pressure ( 0 to + 300 mmHg) | YES
29 | Pre filter pressure ( 0 to +450 mmHg) | YES
30 | Drain line pressure ( -350 to +350 mmHg) | YES
31 | Fluid Control units (Scales): | YES
32 | At least 4 independent scales & scale range from 0 to 11kg at least | YES
33 | Measuring principle: | Gravimetric
34 | Screen with the following specifications:
35 | On-screen service menu preferably via touch screen control | YES
36 | Screen size not less than 10.4” Color, TFT-LCD | YES
37 | User Interface with graphical treatment parameters (Curves), Continuous TMP and filter pressure drop monitoring
38 | The unit preferably should have the ability to recognize the type of the set, allows traceability & Automatic set parameters setting to avoid any human mistakes | YES
39 | Anti electrostatic device (ASD) to avoid ECG interferences | YES
40 | The unit must have citrate anti-coagulation module with the capability of performing Citrate Anticoagulation in all CRRT therapies (CVVH & CVVHD and CVVHDF) | YES
41 | The unit should have heating system to heat the blood or fluid to maintain patient temperature | YES
42 | The unit should have back up battery, not less than 15 min | YES
43 | Additional specifications :
44 | Flexibility of extracorporeal circuit flow-path management | YES
45 | Combined simultaneous Pre & Post fluid Replacement | YES
46 | To be used for citrate anticoagulation
47 | The unit should have Fully Pre-connected complete treatment set to allow fast installation & avoid human mistakes | YES
48 | The unit should have Pump protection faceplates to protect both operator and pumps and ensure safety operation | YES
49 | The unit preferably shall have the ability to record all the treatment sessions & events on the Unit for almost complete one week. It can be loaded on computer to build analysis & statistics | YES
50 | The unit should be able to be upgradeable for any future applications. | YES
Company Stamp |  |  |  | Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | MHD10004 | 0 | DIALYSIS HAEMOFILTRATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Haemodialysis unit to be used with chronic patients Adult and Pediatric and Neonatal in the kidney unit, and shall be capable of performing the following procedures: | YES
2 | Haemodialysis (HD) | YES
3 | Isolated UF | YES
4 | Sodium profiling & Bicarbonate Profiling and UF profiling | YES
5 | Unit shall have One blood pumps , and provide a continuous range of flow from 50 up to 500 ml/min. | YES
6 | Unit should have Single Needle System ( SN/SP). | YES
7 | Unit shall be connected to a centralised water treatment plant supply or a portable reverse Osmosis machine. | YES
8 | The unit shall have an integrated heparin syringe pump able to supply a continuous flow | YES
9 | range from 0.5 up to 10 ml/hour and give measurable stat doses. | YES
10 | Unit shall be equipped with BP monitor to measure systolic & diastolic & mean and heart rate, that can be controlled within the software of the unit and the measurement intervals Shall be adjustable. | YES
11 | BP cuffs must be provided in sizes for pediatric up to extra large adult, which are compatible with the standard of cuff used by .
12 | Unit shall use powder / dry bicarbonate, and provide a connection to all electrolyte solutions Requested. | YES
13 | Unit shall be able to perform a true sodium profiling without affecting minor electrolytes. | YES
14 | Unit should have ionic clearance measurement device which allow the operator to enter the Target KT/V and give a forecast based on the current clearance parameters. | YES
15 | Unit shall be provided with large display : specify size provided | YES
16 | All parameters has to be clearly displayed providing operator with an easy interact with. | YES
17 | Unit shall be capable to store dialysis records. | YES
18 | Unit extracorporeal circuit shall be equipped with blood sensor that will trigger in case of blood leakage rate exceeds 0.50 ml/min | YES
19 | Unit shall have a sophisticated alarming mechanism for all machine’s parameters, such as conductivity, temperature, blood volume, blood leakage, water supply flow and pressure etc, | YES
20 | the unit shall block the patient circuit in case of a potential risk alarm is triggered.
21 | Machine shall be provided with alarming even if the power is cut or power supply fails, the alarm shall be audiovisual and the machine shall go to a safe condition and with certain High risk alarms, clamp the venous return line. | YES
22 | Machine shall be equipped with back-up battery that can operate the machine for at least 15 - 30 minutes in case of power failure | YES
23 | The blood pump should be equipped with emergency hand crank to return the blood to the patient in case of power failure | YES
24 | Unit shall have procedure timer on the front panel; the timer shall show the elapsed time Since the procedure starts or remaining time till finish. | YES
25 | Unit shall be easy to operate and shall have intuitive user interface. | YES
26 | Unit shall have integrated spillage proof key pad | YES
27 | Unit should be equipped with the following sensors for the blood circuit | YES
28 | Dialysate Delivery | YES
29 | Comfort Control ˚C : ( 35 – 39 ). | YES
30 | Conductivity Range : 13 – 15 mS/cm | YES
31 | UF Removal Rate : 0.1 – 3 L/h. | YES
32 | Temp. Alarm Limits : 34 - 40 | YES
33 | Unit should be equipped with the following sensors for the blood circuit
34 | Artrial pressure sensor, range ( -350 to at least +150 mmHg) | YES
35 | Venous pressure sensor (-100 to at least 350 mmHg) | YES
36 | Ultrasonic air sensor for the dripping chamber | YES
37 | A blood volume monitor to measure the change in blood volume DBv to allow early detection of volume depletion. | Preferred
38 | A mechanism to stop the blood returned to patient in case of alarm. | YES
39 | The dialysis machines have ability to provide adjustable dialysis fluid flow rates i.e. bypass (off) and up to 700 ml/min : Specify your range & Steps. | YES
40 | Unit should provide Ultrapure Conventional dialysis by using built in Ultra dialyzer filter for helping achieve sterile substitution and disinfection fluids. | YES
41 | The machines blood pump being easily adjustable by operator to provide blood flow rates Down to 50ml/min via pediatric blood lines. | YES
42 | Unit should have chemical and heat disinfections cycles (preferably open system) with temperature range 90-95 C degrees in all internal parts of the machine to destroy any organism at this temperature | YES
43 | Unit shall be equipped with audiovisual alarm, alarm should not be defeated indefinitely, And the volume shouldn’t be adjustable below the hearing level. | YES
44 | Unit shall have service routines for checking various machine parameters. | YES
45 | Unit shall have service menu for monitoring and controlling parameters and to perform calibrations and service for the various machine’s functions | YES
46 | Unit shall be compatible with all consumables stated. | YES
47 | Unit shall have integrated network connectivity and serial communication protocols. | YES
48 | Unit shall have a safety class I. | YES
49 | Vendor machine be compatible with hospital information system and connectable to Lab / dialysis/ dr. clinic. | YES
50 | Vendor shall conduct comprehensive training for all operators; training shall be conducted by application specialist approved by the manufacture. | YES
51 | Mechanism that can be unlocked by machine operator. | YES
Company Stamp |  |  |  | Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | MHD10005 | 0 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Haemodialysis unit to be used with chronic patients in the kidney unit, and shall be capable of performing the following procedures: | YES
2 | Haemodialysis | YES
3 | Isolated UF | YES
4 | Sodium profiling & Bicarbonate Profiling and UF profiling | YES
5 | Unit shall have One blood pumps , and provide a continuous range of flow from 50 up to 500 ml/min. | YES
6 | Unit should have Single Needle System ( SN/SP). | YES
7 | Unit shall be connected to a centralised water treatment plant supply or a portable reverse Osmosis machine. | YES
8 | The unit shall have an integrated heparin syringe pump able to supply a continuous flow | YES
9 | range from 0.5 up to 10 ml/hour and give measurable stat doses. | YES
10 | Unit shall be equipped with BP monitor to measure systolic & diastolic & mean and heart rate, that can be controlled within the software of the unit and the measurement intervals Shall be adjustable. BP cuffs must be provided in sizes for pediatric | YES
11 | up to | YES
12 | Unit shall use powder bicarbonate, and provide a connection to all electrolyte solutions Requested. | YES
13 | Unit shall be able to perform a true sodium profiling without affecting minor electrolytes. | YES
14 | Unit should have ionic clearance measurement device which allow the operator to enter the Target KT/V and give a forecast based on the current clearance parameters. | YES
15 | Unit shall be provided with large display : specify size provided | YES
16 | All parameters has to be clearly displayed providing operator with an easy interact with. | YES
17 | Unit shall be capable of automatically adjusting UF & Na levels to prevent IDH episodes : Specify the IDH prevention mechanism. | YES
18 | Unit shall be capable to store dialysis records. | YES
19 | Unit extracorporeal circuit shall be equipped with blood sensor that will trigger in case of blood leakage rate exceeds 0.50 ml/min | YES
20 | Unit shall have a sophisticated alarming mechanism for all machine’s parameters, such as conductivity, temperature, blood volume, blood leakage, water supply flow and pressure etc, the unit shall block the patient circuit in case of a potential ris | YES
21 | k alarm | YES
22 | Machine shall be provided with alarming even if the power is cut or power supply fails, the alarm shall be audiovisual and the machine shall go to a safe condition and with certain High risk alarms, clamp the venous return line. | YES
23 | Machine shall be equipped with back-up battery that can operate the machine for at least 15 - 30 minutes in case of power failure | YES
24 | The blood pump should be equipped with emergency hand crank to return the blood to the patient in case of power failure | YES
25 | Unit shall have procedure timer on the front panel; the timer shall show the elapsed time Since the procedure starts or remaining time till finish. | YES
26 | Unit shall be easy to operate and shall have intuitive user interface. | YES
27 | Unit shall have integrated spillage proof key pad | YES
28 | Unit should be equipped with the following sensors for the blood circuit | YES
29 | Dialysate Delivery | YES
30 | Comfort Control ˚C : ( 35 – 39 ). | YES
31 | Conductivity Range : 13 – 15 mS/cm | YES
32 | UF Removal Rate : 0.1 – 3 L/h. | YES
33 | Temp. Alarm Limits : 34 - 40 | YES
34 | Unit should be equipped with the following sensors for the blood circuit
35 | Artrial pressure sensor, range ( -350 to at least +150 mmHg) | YES
36 | Venous pressure sensor (-100 to at least 350 mmHg) | YES
37 | Ultrasonic air sensor for the dripping chamber | YES
38 | A blood volume monitor to measure the change in blood volume DBv to allow early detection of volume depletion. : Specify | YES
39 | A mechanism to stop the blood returned to patient in case of alarm. | YES
40 | Preferably Access flow meter to show patient access flow rate (Qa in ml/min). | YES
41 | The dialysis machines have ability to provide adjustable dialysis fluid flow rates i.e. bypass (off) and up to 700 ml/min : Specify your range & Steps. | YES
42 | Unit should provide Ultrapure Conventional dialysis ( preferably sterile ) by using built in Ultra dialyzer filter. | YES
43 | The machines blood pump being easily adjustable by operator to provide blood flow rates Down to 50ml/min via pediatric blood lines. | YES
44 | Unit should have chemical and heat disinfections cycles (preferably open system) with temperature range 90-95 C degrees in all internal parts of the machine to destroy any organism at this temperature | YES
45 | Unit shall be equipped with audiovisual alarm, alarm should not be defeated indefinitely, And the volume shouldn’t be adjustable below the hearing level. | YES
46 | Unit shall have service routines for checking various machine parameters. | YES
47 | Unit shall have service menu for monitoring and controlling parameters and to perform calibrations and service for the various machine’s functions | YES
48 | Unit shall be compatible with all consumables stated. | YES
49 | Unit shall have integrated network connectivity and serial communication protocols. | YES
50 | Unit shall have a safety class I. | YES
51 | Vendor machine be compatible with hospital information system and connectable to Lab / dialysis/ dr. clinic. | YES
52 | Vendor shall conduct comprehensive training for all operators; training shall be conducted by application specialist approved by the manufacture. | YES
53 | Mechanism that can be unlocked by machine operator. | YES
54 | Vendor must provide an itemized price for the following : | YES
55 | Qty : 1000 Blood Line Tubes Sets | YES
56 | Qty : 1000 Cartridges Bicarbonate 650 g. | YES
57 | PORTABLE RO WATER | YES
58 | WATER SAVING FUNCTION | Yes
59 | PROGRAMMABLE REMINDERS FOR CLEANING, FILLING OF SALT TABLETS CONTAINER IN THE SOFTENER | Yes
60 | PRODUCT WATER CONDUCTIVITY DISPLAY | Yes
61 | INLET WATER CONDUCTIVITY DISPLAY | Yes
62 | REJECTION RATE % DISPLAY | Yes
63 | DATE, TIME, TIME SINCE LAST DISINFECTION AND TOTAL RUN TIME DISPLAY | Yes
64 | AUO /MANUAL START UP FOR HEAT DISINFECTION | Yes
65 | CUSTOMIZED CLEANING PROGRAM FOR DIFFERENT NEEDS | Yes
66 | MACHINE REJECTION RATE IS MORE THAN 95% FOR TOTAL DISSOLVED SALTS | Yes
67 | MACHINE REJECTION RATE IS MORE THAN 99% FOR BACTERIA AND PYROGEN | Yes
68 | PC PRESET AND SOFTWARE UPGRADING | Yes
69 | PROGRAMMABLE TUBING FLUSH DURING STANDBY TO PREVENT BACTERIA GROWTH | Yes
70 | PC LOGGING ATO MONITOR THE FLOW DIAGRAM DURING OPERATION FOR TROUBLESHOORING | Yes
71 | PRODUCT WATER
72 | out put water minimum 1.11 l/min | Yes
73 | total dissolved salt more than 96% | Specify
74 | bacteria and pyrogens more than 99% | Specify
75 | FEED WATER SUPPLY
76 | input : 3.0l/min | Specify
77 | pressure : 1 - 8 bar | Specify
78 | TDS below than 1500mg/l | Specify
79 | chlorine total below 0.1mg/l | Specify
80 | TEMPERATURE RANGE : 1-100 C | Specify
81 | Size | Specify
82 | PERFERABLY CONTAIN INTERNAL RESERVIOR | Specify
83 | POWER SUPPLY 220 VAC / 60 HZ | Yes
84 | FDA and CE approved | Yes
Company Stamp |  |  |  | Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | MIC10002 | 0 | ANALYZER BLOOD GAS BENHCTOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Fully automated, microprocessor controlled, table-top blood gas and electrolytes analyzer for whole blood, plasma and serum samples , PLEURAL FLUID & DIALYSATE | YES
2 | To incorporate a LCD screen and alphanumeric keyboard for user interfacing. Provide detailed specs (touch control, size, color or mono, resolution, etc.) | YES
3 | To incorporate a built-in thermal printer. Provide detailed specs (paper size and type (roll, Z-fold, etc.), resolution, printing format and information (data, tables, patient demographics, normal values, abnormal flags, etc.) | YES
4 | Measured blood gas parameters to include (Specify measurement range, units, accuracy as well as any other relevant information): | YES
5 | PCO2 / PO2 (total oxygen concentration in blood ) | YES
6 | CO-Oximetry panel shall be integrated. | YES
7 | SO2 (Oxygen saturation of hemoglobin in blood ) | YES
8 | HCO 3 (Concentration of hydrogen carbonate in plasma ) | YES
9 | Aa DpO2 (Alveolo-arterial oxygen tension difference for arterial blood) | YES
10 | SBE & ABE (standard and actual base excess) | YES
11 | SBC (standard bicarbonate) | YES
12 | Plasma tCO2 (Concentration of total CO2 in plasma) | YES
13 | PT PH (pH of plasma at patient temperature) | YES
14 | PT PCO2 (CO2 tension in blood at patient temperature) | YES
15 | PT H+ (Concentration of hydrogen ions in blood at patient temperature | YES
16 | PT PO2 (O2 tension in blood at patient temperature) | YES
17 | Measured electrolytes (Specify measurement range, units, accuracy as well as any other relevant information): | YES
18 | User entered data shall include (provide details): | YES
19 | Patient demographics (name, ID, age, sex, height, weight) | YES
20 | Technician ID | YES
21 | Physician name | YES
22 | Location (outpatient, in-hospital department or bed, etc.) | YES
23 | Sample type and time taken | YES
24 | Patient temperature CORRECTION | YES
25 | Hb | YES
26 | FiO2 | YES
27 | Automatic washing capability. Specify parameters. | YES
28 | Hands free sample administration (syringe or capillary tube): | YES
29 | Fully automated aspiration, transport, analysis and processing of sample | YES
30 | State sample volume in all applicable modes (normal, micro, etc.) | YES
31 | State type, stability and accuracy of temperature control | YES
32 | Specify test throughput (analysis time) from “READY” state as well as start up time from off to ready | YES
33 | Automatic calibration. State levels, schedule (or time interval) and duration. Calibration interruption for STAT samples - LIQUID OR GAS CALIBRATION | YES
34 | On-board QC management | YES
35 | Barcode reading capability for sample and reagent handling. Automatic reagent tracking (level detection, expiry date, etc.) | YES
36 | Built-in automatic maintenance program, preferably zero user maintenance | YES
37 | State type of required gas mixture and method of delivery | YES
38 | Internal troubleshooting (self diagnostic) software capability is an asset. Specify details. | YES
39 | Computer connectivity for Bi-directional data, results and report transmission from and to the Hospital Information System including archiving capabilities | YES
40 | The following features of the offered system are preferred. State any that apply to your system. If optional, Price separately: | YES
41 | Sealed waste container | YES
42 | On-board help function | YES
43 | Remote diagnostics | YES
44 | Sodium / Potassium / Calcium / Chloride | YES
45 | PH | YES
46 | LACTATE / HCT
47 | PROVIDE ITEMIZED PRICE FOR 1,000 TESTS | SPECIFY ITEMIZED PRICE EACH TEST DEVICE CAN DO
48 | LANGUAGE | ENGLISH
49 | PLUG | BRITISH
50 | SAMPLE VOLUME | 200 uL APPROX
51 | ANALYSIS TIME | LESS THAN 100 SEC. APPROX
52 | ELECTRODE | LONG LIFE ELECTRODE FOR AT LEAST 1 YEAR
53 | CERTIFICATION | FDA , CE OR APPROVED BY REGULATORY BODY
54 | UPS | 1 HOUR BACKUP THAT MATCH DEVICE POWER REQUIREMENT
Company Stamp |  |  |  | Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | MIC10025 | 0 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The system should be based on the latest technology in Monitors, patient transport that cover all applications needed in ICU. | Yes
2 | The monitor should have Intensive care software for Adult, Pediatric. | Yes
3 | The monitor should provide a high level of accuracy of monitored parameters. | Yes
4 | The monitor should be mobile stand with lockable wheels (minimum of 5), with basket, | Yes
5 | The monitor should provide a high level of accuracy (e.g. 2% approximate for ECG) of all monitored parameters. Vendors should provide parameters accuracy details. | Yes
6 | The monitor should be modular with single or multi-parameters per module. | Yes
7 | The display should be color LCD and should have an approximate size of 15 inches with a high level of contrast, brightness, resolution, and visibility from a distance and at a wide angle. Monitored parameters values should be large enough to be | Yes
8 | read from a distance. | Yes
9 | The monitor should be a touchscreen. | Yes
10 | The monitor should be capable of displaying six (6) physiological waveforms simultaneously as a minimum. | Yes
11 | The monitor should be capable of interfacing with the Hospital and Bedside Clinical Information System (HCIS). It should be able to download any recorded parameter, waveforms and numerical data, to the HCIS. The vendor should submit a list of | Yes
12 | compatible HCIS. (Please Provide it as itemized price) | Yes
13 | The vendor should provide state of the art technology for information technology used in the monitor applications. | Yes
14 | The vendor should provide an itemized price for all clinical software that is used on the monitor (vendor should support documents with evidence that prove the benefit of the clinical programs) | Yes
15 | The unit should be well constructed with durable materials to withstand typical abuse, heavy used environment. The vendor should provide support document that proves the durability of his product (International reference site should be mentioned). | Yes
16 | The monitor should be capable of providing an auxiliary ECG output signal for defibrillator/IABP synchronization, and an auxiliary BP output signal for IABP synchronization (preferable). | Yes
17 | The monitor should provide user-adjustable visual and audible alarms tailored for individual parameters. | Yes
18 | The monitor should be capable of providing graphical and numerical 24-hours trends of all monitored parameters. | Yes
19 | The monitor should be software-upgradeable to include future updates and/or additions of software application packages. | Yes
20 | The vendor should provide detailed technical information for the unique feature in the monitor. | Yes
21 | The vendor should provide detailed technical information about the following: | Yes
22 | HR Accuracy | Yes
23 | ECG HR range, bpm | Yes
24 | Interpretation | Yes
25 | Arrhythmia detect technique and number of lead analyzed | Yes
26 | ST technique and No. of leads analyzed | Yes
27 | Respiration method and Waveform displayed feature | Yes
28 | DISPLAY Size, (cm) in | Yes
29 | Display User Interface feature | Yes
30 | Remote display availability | Yes
31 | Trending Parameters | Yes
32 | Networking and Centralization. | Yes
33 | Hardwired/wireless | Yes
34 | Communication protocols | Yes
35 | Alarm Types and Feature for the machine | Yes
36 | The monitor should be capable of recording and display in waveforms and numerical values the following parameters (please provide it as itemized price): | Yes
37 | ECG, with 5 leads display, detection of pacemaker spikes, and arrhythmia. | Yes
38 | Respiration, with waveform display and alarms. | Yes
39 | Invasive Pressure (IP). | Yes
40 | Pulse Oximetry (SpO2), with Nellcor-Oxymax or Massimo RD Set cables and finger probes. | Yes
41 | Non-Invasive Blood Pressure (NIBP), and quick-connect hoses. | Yes
42 | Temperature, in two channels for two simultaneous probes, skin and rectal. | Yes
43 | End-Tidal CO2 (EtCO2), with waveform display, rapid warm-up and calibration times, and low dead-space connectors. | Yes
44 | Cardiac Output. | Yes
45 | Level of Consciousness. | Yes
46 | The vendor should provide an itemized price for different modules. | Yes
47 | The monitor should have a module that can measure at the same time: | Yes
48 | ECG, RESP. | Yes, QTY: 1
49 | Pulse Oximetry (SpO2), Nellcor-Oxymax or Massimo. | Yes, QTY: 1
50 | Non-Invasive Blood Pressure (NIBP). | Yes, QTY: 1
51 | Invasive Blood Pressure (IBP). | Yes, QTY: 2
52 | Temperature. | Yes, QTY: 1
53 | All patient input should be defibrillator-protected. | Yes
54 | All cables, probes, and transducers (except for disposables) for all required parameters should be provided with the monitor, for adults and Pediatric. | Yes
55 | All Hardware and Accessories for each module should be designed for all type of patient use (provide itemized price). | Yes
56 | The vendor should provide an itemized price for all option, accessories on the machine. | Yes
57 | The manufacturer should guaranty that spare parts and technical support will be provided for at least ten years. | Yes
58 | Hardware and software should be industry standard with the latest specification; the vendor should provide a specification of both hardware and software used in the system. | Yes
59 | The vendor should provide an itemized price for all options in the system. | Yes
60 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and should be secured with adequate strain reliefs. | Yes
61 | The chassis should be grounded and grounding resistance should not exceed 0.15 ohm. | Yes
62 | If the unit is double insulated, it should be so labeled. | Yes
63 | Electrical leakage current from the chassis of the system should not exceed [0.5 mA per IEC 601-1 or 0.3 mA in the U.S. per NFPA 99-1993].Operator safety and system performance should not be adversely affected by fluid spills. | Yes
64 | The system performance should not be affected by EMI radiated or conducted through the power lines from another device. | Yes
65 | The power cord should be automatic retractable. | Yes
66 | The vendor should provide a full warranty of at least three (3) years. | Yes
67 | General Specifications and Requirements: | Yes
68 | Power Supply: 230 VAC 60 Hz. | Yes
69 | Hospital Grade British type Power Plugs, Line Cords, and Strain Relief. | Yes
70 | The unit must meet the applicable requirements and standards of the Underwriters Laboratories (UL), CE mark, IEC, ISO, and/or FDA Approval. | Yes
71 | The unit should be manufactured according to GMP guidelines. | Yes
72 | Labeled with all appropriate operation and safety tags and symbols. | Yes
73 | Switches, knobs, and other controls should be visible and clearly identified, and their functions should be self-evident. | Yes
74 | Switches, knobs, and other controls should be protected against accidental setting changes and should be sealed against penetration of liquids. | Yes
75 | The unit should be simple to operate and easy to clean and disinfect with Isopropanol alcohol 70%. | Yes
76 | The exterior should be well constructed with durable fluid/shock resistant materials to withstand typical abuse and cleaning. | Yes
77 | Rugged to handle in the rough transport environment. | Yes
78 | The unit should have no sharp edges. All external components should be securely mounted. | Yes
79 | All available options, consumables, or disposables deemed necessary for the intended function of the equipment and not listed in the Technical Specifications should be included and specified. | Provide Itemized Price
80 | Wall mount brackets as itemized price | Provide Itemized Price
81 | ETCO2 itemized | Provide Itemized Price
82 | Central station itemized (license of 20 beds) | Provide Itemized Price
83 | Itemized: The monitor should have a module that can measure 8 waves at the same time as follows : | Provide Itemized Price
84 | QTY: 1 ECG, RESP. | Provide Itemized Price
85 | QTY: 1 Pulse Oximetry (SpO2), Nellcor-Oxymax or Massimo. | Provide Itemized Price
86 | QTY: 1 Non-Invasive Blood Pressure (NIBP). | Provide Itemized Price
87 | QTY: 3 Invasive Blood Pressure (IBP). | Provide Itemized Price
88 | QTY: 2 Temperature. | Provide Itemized Price
Company Stamp |  |  |  | Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | MIC10039 | 0 | DOCKING PUMP INFUSION 4 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | INFUSION PUMP | QTY : 4 PUMPS
2 | QUICK START OF INFUSION | YES
3 | FLOW RANGE ML/HR | 0.1 - 999 ML/HR
4 | ACCURACY | 0.05
5 | DRUG/DOSE CALC | YES
6 | VTBI, ML | 0.1 TO 9999
7 | MAX PRESSURE, PSI | SPECIFY MAX
8 | PUMPING MECHANISM | SPECIFY MECHANISM
9 | IV SET | YES
10 | FREE-FLOW PROTECTION | YES
11 | MANUAL AND PROGRAMMABLE BOLUS | PLEASE SPECIFY BOLUS RATE
12 | ALARMS & INDICATORS
13 | EMPTY BAG | YES
14 | FLOW ERROR | YES
15 | DOOR OPEN | YES
16 | CONTINUOUS PRESSURE MONITORING PSI (ON THE FRONT SCREEN ALL THE TIME) | SPECIFY
17 | DRUG DOSE CALCULATION | ML/H, µG/H, µG/MIN, µG/KG/H, µG/KG/MIN, MG/H, MG/MIN,MG/KG/MIN, MG/KG/H, IU/H, IU/MIN, IU/KG/H, IU/KG/MIN.
18 | FLEXIBLE PRESSURE ALARM SETTINGS | IN LINE PRESSER
19 | HIGH PRESSURE/OCCLUSION | YES
20 | PRE- OCCLUSIONALARM | YES
21 | CIRCUIT MALFUNCTION | YES
22 | INFUSION NEAR END | YES
23 | INFUSION COMPLETE | YES
24 | LOW BATTERY | YES
25 | DEPLETED BATTERY | YES
26 | USER PROMPT | YES
27 | AUDIBLE ALARM | YES
28 | VOLUME CONTROL | YES
29 | MOMENTARY SILENCE | YES
30 | PPM DUE DATA REMINDER | YES
31 | FLUID RESISTANT | YES
32 | NURSE CALL JACK | YES
33 | DATA LOG | YES
34 | EVENT LOG | YES
35 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED)
36 | NUMBER OF EVENTS | 1500
37 | PRINT OUT | YES
38 | DISPLAY | LED AND LCD WITH COLOR CABLITIES IS A PLUS & BIG SCREEN INTERFACE IS ALSO A PLUS
39 | INFUSED VOLUME DISPLAY | YES , DIGITAL
40 | WEIGHT WITH BATTERY KG | SPECIFY WIGHT
41 | TIME /DATE | YES
42 | POWER SOURCE, VAC | 220VOLT / 60 HZ.
43 | POWER SUPPLY | INTERNAL OR EXTERNAL (IF EXTERNAL MUST BE LIFETIME WARRANTY)
44 | POWER CABLE | DURABLE POWER CABLE
45 | BATTERY | YES, RECHARGEABLE BUILT IN
46 | BATTERY | NI-MH
47 | OPERTING TIME , HR | SPECIFY HOURS @ 10 ML/HR
48 | RECHARGE TIME , HR | PREFERRED LOWER CHARGING HOURS
49 | POLE MOUNTING | PREFERRED
50 | OTHER SPECIFICATIONS | FDA / CE
51 | STACKABLE IN THE SAME DOCKING STATION | YES
52 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED).
53 | STACKABLE IN THE SAME DOCKING STATION | INFUSION PUMPS / SYRINGE PUMPS / ENTRAL FEEDING PUMPS.
54 | DOSE ERROR REDUCTION SYSTEM (SMART TECHNOLOGY) | YES
55 | LIBRARY SIZE | YES CONFIGURABLE
56 | NO. OF DRUG ENTITIES/CARE AREA | YES PLEASE SPECIFY
57 | WIRELESS CONNECTIVITY | UPGRADABLE
58 | TIME RETAINED | FIRST IN FIRST OUT
59 | ACCESSORIES FOR THE PUMP | CLOSED SYSTEM
60 | WIRE CONNECTIVITY | FOR DOCKING STATION
61 | DOCKING STATION | INCLUDED TO OCCUPY 4 PUMPS
62 | ALL SETS HAVE SELF -SEALING NEEDLE FREE INJECTION SITES | YES
63 | ALL SETS SHOULD BE LATEX DEHP & PVC FREE | YES
64 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
65 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
66 | ACCESSORIES FOR INFUSION PUMP | CLOSE SYSTEM
67 | THE VENDOR MUST PROVIDE A FIXED PRICE FOR THE ALL ABOVE MENTIONED ACCESSORIES VALID FOR 3 YEARS | YES
68 | PUMP HL7 READY | YES ( PLEASE SPECIFY )
69 | PROVIDE CONSUMABLE DEAL ( PUMPS FREE AND PRICE ON SETS ) | OPTIONAL DEAL SECTOR MIGHT CHOOSE
70 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
71 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
72 | THE WORKSTATION SHOULD BE MOUNTABLE , WALL UNIT (FIXATION TO WALL RAILS OF DIFFERENT MANUFACTURER).MOVABLE POLE-BASED UNIT.
73 | SMART ALARM SYSTEM BY AUDIBLE, VISUAL AND DISPLAYED MESSAGE. | OCCLUSION, END OF INFUSION, EMPTY BATTERY, NO RATE SET, DEFECT ON PUMP. PRE-ALARMS: LOW BATTERY, NEAR-END OF INFUSION , PLEASE SPECIFY
74 | MOBILE STAND TO FIX THE PUMP ON INSTEAD OF PENDANT OR WALL FOR SPECIAL SITE NEEDS | ITEMIZED PRICE OPTIONAL
75 | PRIMARY INFUSION SET (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 120 SETS WITH EACH PUMP ITEMIZED
76 | LIGHT SENSITIVE PRIMARY INFUSION SET WITH NO INJECTION PORT FOR NEONATAL TPN (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 40 SETS WITH EACH PUMP ITEMIZED
77 | PROGRAMMING PATINET ID / NURSE ID ON THE PUMP DURING ADMINISTRATION | YES ( PLEASE SPECIFY )
Company Stamp |  |  |  | Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | MIC10044 | 0 | VENTILATOR CPAP ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Country Of origin | Specify
2 | Model | Specify
3 | Manufacture | Specify
4 | VENTILATOR | CPAP
5 | Gas supply pressure | 40,6 - 87 psi
6 | PANUMATIC OPERATION | yes
7 | Breathes per Minute / Accuracy / Range 1 second inspiretor / Range 2 second inspiretory | 10% / 0 - 30 bpm / 0 - 20 bpm
8 | Tidal Volume / Accuracy / Range 1 / Range 2 | 10 % / 200 - 600 ml / 400 - 1500 ml
9 | Adjustable tidal volume | yes
10 | Adjustable BpM | yes
11 | Chilled and Adult settings | yes
12 | Internal demand valve | yes
13 | Antibreth stochking | yes
14 | Manometer readout | yes
15 | Adjustable relife with alarm | yes
16 | Low source gas pressure alarm | yes
17 | Accessories
18 | Ventilator Circuit | yes
19 | Oxygen House | yes
20 | Oxygen Regulator | yes
21 | Internal air filter | yes
22 | CPAP ventilator circuit Disp. | yes ,50 set to be included with each unit
23 | OTHER SPECIFICATION | FDA , CE , APPROVED.
Company Stamp |  |  |  | Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | MIC10046 | 0 | VENTILATOR ICU ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SYSTEM SELF CHECK | YES
2 | Controlled / Assist modes | Pressure, Volume & Volume Guaranteed
3 | Intermittent modes | SIMV-VC, SIMV-PC , SIMV-VG
4 | Two level of pressure mode | BiPAP Invasive OR Equivelent BiPAP , NIV-ST ( with Mask )
5 | Spontaneous mode | CPAP/PS mode
6 | NIV , NIV-ST | REQUIRED
7 | Mode Settings & controls
8 | Tidal Volume | 50 - 2000 ml
9 | Respiratory Rate | 3 - 80 b/m
10 | FiO2 | 21% - 100%
11 | Inspiratory Pressure | 1 - 60 cmH2O
12 | PEEP | 0 - 40 cmH2O
13 | Pressure Support | 0 - 60 cmH2O
14 | I:E ration | 1:9 to 4:1
15 | Inspiratory Time | 0.25 - 15 sec
16 | Flow Trigger | 1 - 9 l/m
17 | Pressure Trigger | -10 to -0.25
18 | Inspiratory Pause | Yes, Specify
19 | Pressure slope/rise time | Yes, Specify
20 | Bias Flow | Yes, Specify
21 | Tube Compensation | Yes
22 | Leak Compensation | Yes
23 | Trigger Compensation | Yes
24 | Apnea backup / rate | Yes
25 | Manual Breath | Yes
26 | Inspiratory Hold | Yes
27 | Expiratory Hold | Yes
28 | AutoPEEP measurement | Yes
29 | Weaning Parameters (P0.1, NIF & RSBI) | Yes, Specify
30 | Lung Compliance & Airway Resistance | Yes
31 | Monitored Parameters
32 | PIP, Pplat, MAP & PEEP | Yes
33 | TV & MV (controlled & Spontaneous) | Yes
34 | Respiratory Rate | Yes
35 | FiO2 | Yes
36 | Clinical Alarms (High/Low)
37 | Peak Pressure | Yes
38 | TV & MV | Yes
39 | Patient Disconnect | Yes
40 | Leak | Yes
41 | Technical
42 | O2 sensor | BUILT IN
43 | Active Expiratory Valve (Autoclavable is preferred) | Yes
44 | electronic micro pump nebulizer | OPTIONAL ITEMIZED PRICE
45 | Display & Control | Colored Touch Screen not less than 12 INCH
46 | Patient circuit supported arm | ADJUSTABLE
47 | Backup battery | not less than 2 hrs
48 | Mobile base/ cart with locking castors | Yes
49 | Central Gas inlets (O2 & Air) | BRITISH STANDARD BOC
50 | Electricity 220v, 60 Hz or Auto. | Yes
51 | Air inlet Filter | Yes
52 | Technical Alarms
53 | Gas supply or power Failure, Vent inoperative & Low battery | Specify
54 | Data Displayed | Numerical, Graph/waveforms & Loops
55 | Trends | not less than 72 hrs for all patient data, Specify
56 | WEANING MODE | yes
57 | Humedifier with Cahmber & Heater Wire : Approved standard by sector | INCLUDED WITH UNIT
58 | MOBILE CART | INCLUDED WITH BASKET
59 | PLUG STANDARD | BRITISH
60 | LANGAUGE | ENGLISH LANGUAGE
61 | UNIT OF MEASUREMENT | INTERNATIONAL STANDARD SI
62 | APRV or equivalent | included
63 | Additional Modes: PAV, ASV, PPS, NAVA & BiLevel-VG or equivalent & Smartcare with Capno | OPTIONAL
64 | P/V tool, Low flow PV, Spirodynamic, Static PV loop, or equivalent & others | OPTIONAL
65 | EtCO2, EtO2, VCO2, VO2, EE & RQ | OPTIONAL
66 | Open Lung tool, Lung Recruitment, FRC with PEEP inview or equivalent & others | OPTIONAL
67 | BREATHING CIRCUIT | 100 WITH EACH UNIT , OPEN STANDARD
68 | DEVICE CERTIFICATION | APPROVED FROM REGULATORY BODY
69 | AIR COMPRESSOR INTEGRATED WITH THE DEVICE | ITEMIZED PRICE
70 | PROVIDE WHERE THE DEVICE & TECHNOLOGY WERE USED | PROVIDE REF HOSPITALS
71 | Modes of Ventilation
72 | BI-LEVEL MODE ( APRV ) | OPTIONAL ITEMIZED
73 | Inverse Ratio mode : IRV | optional
Company Stamp |  |  |  | Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | MIC10055 | 0 | ULTRASOUND POC GENERAL PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Regulation Standard :
2 | FDA 510 (K) Clearance / CE Mark & SFDA Registration | Yes
3 | Validated Up to Date Certificates Must be Submitted with the Offer | Yes, Failure of Submission and Comply Will Disqualify the Offer
4 | Product Description :
5 | Range | Laptop Ultrasound
6 | Purpose | Used for point of care Critical Care
7 | Console Design :
8 | Laptop / Hand Carried on height adjusted mobility Cart | Cart Should Include the Following :Integrated Power Supply , Probe Ports , wheels with locking braking system
9 | Total Weight - (incl. battery) | ≤ ( 5 ) kg .
10 | Easy Acess mode keys | Yes, (Specify)
11 | Screen Monitor | ≥ (10) inch. LCD with High Resolution ( maximum active screen size ratio )
12 | Security Lock | Yes
13 | Built-In Battery | Yes
14 | Battery scanning time | 1 hour
15 | Power Requirement | 220-240 V / 60 Hz
16 | Computing Capabilities :
17 | Operation System | Windows or equivalent
18 | Integrated Hard Desk type | Flash memory≥8 GB for better shock absorbence
19 | Fast System boot up | Yes, <=30 sec
20 | DVD /CD burner/USB port | Yes
21 | Clinical Applications : | Analysis Packages Standard
22 | Vascular Access and Nerve Block | Yes
23 | Other Available Applications | Yes
24 | Multi Frequency Probes : | One Each , With dedicated Operational Software
25 | Linear array | 5-10 MHz ,
26 | Convex array | 2-5 MHz,
27 | Phased array | 1-5 Mhz,
28 | Active Probe Ports | 2 Active Ports
29 | Operating Modes :
30 | Brightness Mode (B-Mode) (2D) | Yes
31 | Motion Mode (M-Mode) | Yes
32 | Color Doppler mode | yes
33 | Pulsed wave Doppler (PW) with HPRF | Yes
34 | High PRF Doppler Mode | Yes
35 | Color Flow Doppler Mode (CFM) | Yes
36 | Power Doppler Mode | Yes
37 | Continuous Wave Doppler | Yes
38 | Scanning Parameters :
39 | Imaging Depth, cm | Not less than 30 cm
40 | Dynamic Range | Up to 150 dB
41 | TGC Control | Yes
42 | Frame Rate 2D | Specify
43 | Frame rate Color Doppler | Specify
44 | GrayScale Levels | Specify
45 | Scanning Automatic Optimization | Capability for Auto optimization in 2D & Doppler
46 | Digital Calipers | Yes
47 | Preprocessing | Yes
48 | Postprocessing | Yes, (Measurement, adjust the Dynamic Range, TGC, color map & Gain) on stored images
49 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
50 | Doppler :
51 | Type | PW in all probes
52 | Frequency Display | Yes
53 | Doppler baseline shift after freeze | Yes
54 | Velocity Display | Yes
55 | User – defined annotation | Yes
56 | GrayScale Levels | Yes
57 | Transmit focal zones | Specify
58 | Digital reconstructed zoom | Specify
59 | Gain Control after freeze | Yes
60 | Adaptive Gain Compensation | Yes
61 | Directional color power Doppler | Yes
62 | Cine Features :
63 | Cine Storage | Yes, (Still images / Clips / Volume)
64 | Length / Capacity | 140 MB (up to 3000 images)
65 | Scanning Features :
66 | Speckle Reduction Imaging | Yes, Organ Specific
67 | Trapeziodal Imaging / Virtual Convex | Yes
68 | B-Mode Steering | Yes
69 | Compound Imaging | Yes
70 | Tissue Harmonic Imaging | Yes
71 | TGC (Time Gain Compensation) | Yes
72 | Color flow mapping | Standard
73 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
74 | Automatic Doppler Calculations | Yes
75 | Adjustable Transmit Focus | Yes
76 | Dynamic Receive Focus | Yes
77 | Selectable Dynamic Range | Yes
78 | Zoom / Pan :
79 | Real- time image | Yes
80 | Frozen image | Yes
81 | Connectivity :
82 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicable
83 | Wireless Connection | Yes
84 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
85 | Storage and Playback of M-mode, PW Spectral Doppler. | yes
86 | Printing of Patient Reports Via a USB Connection. | Yes
87 | Recalculation for stored images | Yes
88 | Security / Privacy :
89 | Accessories :
90 | Secondary Battery | Yes
91 | Carrying Case | Yes
92 | MOH Requirement :
93 | All the vendors should supply all available options of their machine not stated in the above specifications as an optional items with separate price. | Yes
94 | Supplier is an authorized representative of the manufacturer by SFDA | Yes (Attach official documents)
95 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | MIC10064 | 0 | HIGH FLOW NASAL CANNULA : HFNC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 29 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit should be a High/Low Oxygen Flow Generator with an integrated humidifier. | YES , SPECIFY
2 | The unit should have a high and low flow of warmed and humidified respiratory gases, delivered to the patient through a variety of nasal, tracheostomy and mask interface. | YES , SPECIFY
3 | The unit should be designed to be used in Adult and Pediatric and Neonatal. | YES , SPECIFY
4 | The unit should be based on the Oxygen Flow Generator latest technology. | YES , SPECIFY
5 | The unit should have a flow rate range 1-50 L/min approximately. | YES , SPECIFY
6 | The unit should have a temperature range of 33-43 ˚C approximately. | YES , SPECIFY
7 | Oxygen and Air can be added from wall supply or cylinder. | YES , SPECIFY
8 | The unit must be capable of delivering FiO2 up to 90%. | YES , SPECIFY
9 | The unit should have a graphical display showing menu setting and protected lock against accidental setting changes. | YES , SPECIFY
10 | The unit should have an integrated Oxygen/Air Mixer/Blender. | YES , SPECIFY
11 | The unit should have an Oxygen sensor. | YES , SPECIFY
12 | The unit should be small in size and lightweight. | YES , SPECIFY
13 | The unit should be mobile in pole stand with caster wheel, hanger for water irrigation and basket for accessories. | YES , SPECIFY
14 | Switches, knobs, and control must be visible and functions should be self-evident and clearly identified and designed for the condition of heavy long-term use and sealed again penetration of liquid. | YES , SPECIFY
15 | The unit must be provided with the operator and technical/service manual. | YES , SPECIFY
16 | The unit should have no sharp edges. | YES , SPECIFY
17 | All external components should be securely mounted. | YES , SPECIFY
18 | The unit should be well constructed with durable materials to withstand typical abuse and cleaning. | YES , SPECIFY
19 | The unit should be simple to learn to use, operate, and maintain. | YES , SPECIFY
20 | The chair should be easy to clean, disinfect, and/or sterilize, as appropriate. | YES , SPECIFY
21 | The unit should be designed for easy access to serviceable parts. | YES , SPECIFY
22 | The unit should comply with the respiratory international standard. | YES , SPECIFY
23 | The vendor should provide a clear technical description of the offered items. | YES , SPECIFY
24 | The vendor should attach a copy of the catalog/brochure of the item, containing its technical specifications. | YES , SPECIFY
25 | All Medical gas hose connector should come with British standard | YES , SPECIFY
26 | The vendor should provide a full warranty of at least three (3) years for all solution components. | YES , SPECIFY
27 | The vendor should provide onsite training for customized based on MNG-HA environment deployment model, and equivalent to the level of manufacturer training center. | YES , SPECIFY
28 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and should be secured with adequate strain reliefs. | YES , SPECIFY
29 | The chassis should be grounded and grounding resistance should not exceed 0.15 ohm. | YES , SPECIFY
30 | If the unit is double insulated, it should be so labeled. | YES , SPECIFY
31 | CERTIFICATION | YES CE / FDA , SPECIFY IF YOU HAVE ANY OTHER APPROVAL FROM REGULATORY BODY
32 | Power Requirements, 220V at 60 Hz with molded British standard Plug, If the unit uses a Power adapter, it should be of the same standard | YES , SPECIFY
33 | Electrical leakage current from the chassis of the system should not exceed [0.5 mA per IEC 601-1 or 0.3 mA in the U.S. per NFPA 99-1993]. | YES , SPECIFY
34 | All type of consumables needed to functional the unit with all available sizes. | Provide Itemized Price
35 | All available options, consumables, or disposables deemed necessary for the intended function of the equipment and not listed in the Technical Specifications should be included and specified. | Provide Itemized Price
36 | LANGUAGE | ENGLISH
37 | UNIT OF MEASURE | INTERNATIONAL SI
38 | CIRCUIT | HEATED WITH HUMIDIFIER
39 | HIGH FLOW CIRCUIT COMPLETED SET (SHOULD BE COMPATIBLE WITH EACH HOSPITAL STANDARD MACHINE) | QTY : 40
40 | HIGH FLOW CIRCUIT COMPLETED SET (SHOULD BE COMPATIBLE WITH EACH HOSPITAL STANDARD MACHINE) | PROVIDE ITEMIZED PRICE IF REQUIRED QTY MORE THAN SPECS # 39
41 | HIGH FLOW NASAL CANNULA, LARGE SIZE, FOR ADULT, (EACH HOSPITAL SHOULD BE COMPATIBLE WITH HEATED VENTILATOR MACHINE) | QTY : 20
42 | HIGH FLOW NASAL CANNULA, MEDIUM SIZE, FOR ADULT, (EACH HOSPITAL SHOULD BE COMPATIBLE WITH HEATED VENTILATOR MACHINE) | QTY : 20
43 | HIGH FLOW NASAL CANNULA, SMALL SIZE, FOR ADULT (EACH HOSPITAL SHOULD BE COMPATIBLE WITH HEATED VENTILATOR MACHINE) | QTY : 20
Company Stamp |  |  |  | Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | MIF40004 | 0 | SURFACE DECONTAMINATION SYSTEM H2O2 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SURFACE DECONTAMINATION SYSTEM - H2O2 , Hydrogen Peroxide Sterilant space up to 250 m3
2 | Features :
3 | Microprocessor Controlled
4 | Hydrogen peroxide vapor
5 | Specifications :
6 | System shall have User Configuration Programs: Specify
7 | Blower speed : 10-40 m3/h
8 | Pump Injection rates :1-16 grams/min.
9 | System shall have Audio visual Alarm.
10 | System shall include all needed Hoses and connections .
11 | Compliance wh standards & legislation:
12 | The system must comply wh the Electrical safety standards for electrical safety IEC-60601
13 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
14 | All electrical connections and plugs should be hospal grade and follow international, local and hospal requirements.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and condions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | MIF50016 | 0 | PAPR RESPIRATOR BIOLOGICAL COMPLETE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 31 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Provides higher filtration efficiency. | YES FOR BIOLOGICAL HAZARDOUS SUCH AS VIRUSES
2 | The blower must have MAIN FILTRATION INLET | YES
3 | The blower should be light weight for easy carry or wear. | YES
4 | The blower should have adjustable west belt. | YES
5 | The blower should have integrated battery not less 4 hour after full charger. | YES
6 | The device display screen must provide warning indicators for the safety status always and be visible for airflow performance / performance filters / battery chargin or similer warmings. | YES
7 | Provides high airflow to maintain positive pressure. | YES, Approx (150 - 200) liters per minute
8 | Provides bump-cap—like head protection; full-helmet design provides more comfortable head room and convenience in donning and doffing. | YES
9 | Requires less than a minute to disassemble or assemble. | YES
10 | The blower must have a lower noise-level system compatible with stethoscope use. | YES
11 | The system should have Long Loose-Fitting Hood w/Cape, ( polyethylene barrier film), double layer. | QTY : 6
12 | PAPR filter for biohazard ( PRE FILTER + HEPA FILTER ) | QTY : 12
Company Stamp |  |  |  | Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | MIF50017 | 0 | LAMP DECONTAMINATION ULTRAVIOLET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 32 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Shall be easily fits in confined spaces | YES
2 | Shall be reduce shadows and deliver more UVC does delivery in less time | YES
3 | Shall be chemical free treatment provides immediate post treatment access | YES
4 | Shall be preferred Shutter Protected Emitter Modules minimize potential for lamp breakage | Specify
5 | Shall he simple automatic operation | YES
6 | Type | Mobile
7 | Type of lamps and useful life | Amalgam germicidal ultraviolet lamps / 16000 hrs or better
8 | No. of lamps | Specify
9 | total UV-C watts | Specify
10 | Door motion sensors | YES
11 | Shall be preferred Safeguards protecting lamps pneumatically driven shutters stem | Specify
12 | No. of Emitters | Specify
13 | Casters with Lock | YES
14 | Motion sensors | Infrared motion sensors shut down unit if motion is detected
15 | Wireless remote control | YES
16 | Remote control | Rechargeable with charging cradle
17 | Password protection | YES
18 | Barcode scanner | YES
19 | Room treatment data replication | Data is transmitted to software wirelessly over 4G network or USB cable
20 | Disinfection documentation | Replicated data generates reports to assist with JCI
21 | Monitoring of service intervals | YES
22 | System must use only UV-C spectrum. UV-A and UV-B are not for safety concern | YES
23 | Device must be real-time UV-C dosage monitoring | YES
24 | Disinfection cycle time, minutes | Specify
25 | Provide the UV-C test equipment option | YES
26 | The Device Should Integrate with the current and future HIS. | Specify
27 | VAC | 220V
28 | CURRENT | 13A
29 | FREQUENCY | 60HZ
30 | PLUG TYPE | 3 Pin British
31 | UNIT COMPLETE WITH FULL ACCESSORIES | YES
Company Stamp |  |  |  | Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | MIF50018 | 0 | CAMERA THERMAL HUMAN FIXED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 33 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Distance Rang | Specify
2 | Distance Rang | Specify
3 | Distance Rang | Specify
4 | Distance Rang | Specify
5 | Sensors | Thermal sensor RGB sensor 3D sensor
6 | Thermal Network Bullet Camera: | Thermal sensor RGB sensor 3D sensor
7 | uncooled VOx thermal sensor technology | Approx. ± 0.5 °C
8 | Detection accuracy of body temperature | Approx. ± 0.5 °C
9 | uncooled VOx thermal sensor technology | Approx. 400 × 300
10 | Detection accuracy of body temperature | Approx. 400 × 300
11 | Measuring speed | YES
12 | Athermalized lens (thermal), focus-free | YES
13 | Measuring speed | Approx. 0.7 - 2 s
14 | Athermalized lens (thermal), focus-free | Approx. 0.7 - 2 s
15 | Weight | Approx. 130 g
16 | Support body temperature measurement | YES
17 | Weight | YES
18 | Support body temperature measurement | Approx. 130 g
19 | Dimensions | YES
20 | Measurement Accuracy: ±0.3°C, with blackbody | YES
21 | Dimensions | Approx. 120 mm x 37 mm x 23 mm
22 | Measurement Accuracy: ±0.3°C, with blackbody | Approx. 120 mm x 37 mm x 23 mm
23 | Active deterrence with white light & siren | USB 3.0 and USB 2.0
24 | Interface | USB 3.0 and USB 2.0
25 | Active deterrence with white light & siren | YES
26 | Interface | YES
27 | Ambient temperature during operation | YES
28 | Support Various lens options up to 13 mm | YES
29 | Ambient temperature during operation | Approx. 10 °C to 30 °C
30 | Support Various lens options up to 13 mm | Approx. 10 °C to 30 °C
31 | 2/2 alarm in/out | Approx. 0 °C to 40 °C
32 | Ambient temperature and when storing or transporting the device | Approx. 0 °C to 40 °C
33 | 2/2 alarm in/out | YES
34 | Ambient temperature and when storing or transporting the device | YES
35 | System Requirements | YES
36 | Micro SD memory | YES
37 | Detector Type | Vanadium oxide uncooled focal plane detector
38 | Processor | Vanadium oxide uncooled focal plane detector
39 | Detector Type | Intel i3/9th generation – Intel i5/ 8th generation (similar or equal)
40 | Processor | Intel i3/9th generation – Intel i5/ 8th generation (similar or equal)
41 | RAM | 16 GB DDR4
42 | Effective Pixels | 16 GB DDR4
43 | RAM | Approx. 400 × 300
44 | Effective Pixels | Approx. 400 × 300
45 | Pixel Pitch | Approx. 17 μm
46 | Interfaces | Approx. 17 μm
47 | Pixel Pitch | USB 3.0 SuperSpeed and USB 2.0 FullSpeed, Power supply & current consumption: USB 3.0: 5 Volt ± 5%, 700 mA max. USB 2.0: 5 Volt ± 5%, 220 mA max.
48 | Interfaces | USB 3.0 SuperSpeed and USB 2.0 FullSpeed, Power supply & current consumption: USB 3.0: 5 Volt ± 5%, 700 mA max. USB 2.0: 5 Volt ± 5%, 220 mA max.
49 | Graphics card | Recommended
50 | Spectral Range | Recommended
51 | Graphics card | Approx. 8 μm–14 μm
52 | Spectral Range | Approx. 8 μm–14 μm
53 | Thermal Sensitivity (NETD) | Approx. ≤40 mK
54 | Operating system | Approx. ≤40 mK
55 | Thermal Sensitivity (NETD) | Windows 10
56 | Operating system | Windows 10
57 | Focal Length | Approx. 7.5 mm; 13 mm
58 | Field of View | Approx. 7.5 mm: horizontal: 30° - 53°, vertical: 22° - 39°
59 | Focus Mode | Fixed-focal
60 | Digital Detail Enhancement (DDE) | YES
61 | Thermal Image Stabilization | Electronic Image Stabilization
62 | AGC | Auto; manual
63 | Noise Reduction | 2D NR; 3D NR
64 | Image Flip | 90°/180°/mirror
65 | Temperature | Approx. 30°C ~+45°C Measurement Range
66 | Temperature Accuracy | Max ±0.3°C with blackbody, Max ±1°C without blackbody
67 | Focal Length | Approx. 4mm - 8mm
68 | Max Aperture | More than F1.5
69 | Angle of View | Approx. H: 84° V: 45°, H: 40° V: 22°
70 | Network Connection | YES
71 | SD card recording | YES
72 | Alarm | Motion detection/privacy mask/audio detection/SD card abnormality/network abnormality/anti-burn warning
73 | Operating Temperature | Approx. 10°C to +30°C
74 | Black Box:
75 | Effective radiant surface | Approx. 70mm×70mm
76 | Temperature resolution | Approx. 0.1
77 | Temperature accuracy | Approx. ±0.2℃
78 | Temperature stability | Approx. ±( 0.10.2 )℃/30min
79 | Effective emissivity | Approx. 0.97
80 | Power supply: | 220VAC
81 | Ambient temperature | Approx. 0℃~40 ℃
82 | Network Connection | YES
83 | Storage | Internal HDD, Up to 8 TB capacity
84 | Operating Conditions | -10°C ~ +55°C
Company Stamp |  |  |  | Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | MIF50019 | 0 | CAMERA THERMAL HUMAN PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 34 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Distance Rang | Specify
2 | Distance Rang | Specify
3 | Distance Rang | Specify
4 | Distance Rang | Specify
5 | Thermal Network Bullet Camera: | Thermal sensor RGB sensor 3D sensor
6 | Sensors | Thermal sensor RGB sensor 3D sensor
7 | Detection accuracy of body temperature | Approx. ± 0.5 °C
8 | uncooled VOx thermal sensor technology | Approx. ± 0.5 °C
9 | Detection accuracy of body temperature | Approx. 400 × 300
10 | uncooled VOx thermal sensor technology | Approx. 400 × 300
11 | Athermalized lens (thermal), focus-free | YES
12 | Measuring speed | YES
13 | Athermalized lens (thermal), focus-free | Approx. 0.7 - 2 s
14 | Measuring speed | Approx. 0.7 - 2 s
15 | Support body temperature measurement | Approx. 130 g
16 | Weight | YES
17 | Support body temperature measurement | YES
18 | Weight | Approx. 130 g
19 | Dimensions | YES
20 | Dimensions | Approx. 120 mm x 37 mm x 23 mm
21 | Measurement Accuracy: ±0.3°C, with blackbody | YES
22 | Measurement Accuracy: ±0.3°C, with blackbody | Approx. 120 mm x 37 mm x 23 mm
23 | Active deterrence with white light & siren | USB 3.0 and USB 2.0
24 | Active deterrence with white light & siren | YES
25 | Interface | USB 3.0 and USB 2.0
26 | Interface | YES
27 | Ambient temperature during operation | YES
28 | Ambient temperature during operation | Approx. 10 °C to 30 °C
29 | Support Various lens options up to 13 mm | YES
30 | Support Various lens options up to 13 mm | Approx. 10 °C to 30 °C
31 | 2/2 alarm in/out | Approx. 0 °C to 40 °C
32 | 2/2 alarm in/out | YES
33 | Ambient temperature and when storing or transporting the device | Approx. 0 °C to 40 °C
34 | Ambient temperature and when storing or transporting the device | YES
35 | Micro SD memory | YES
36 | System Requirements | YES
37 | Processor | Vanadium oxide uncooled focal plane detector
38 | Detector Type | Vanadium oxide uncooled focal plane detector
39 | Processor | Intel i3/9th generation – Intel i5/ 8th generation (similar or equal)
40 | Detector Type | Intel i3/9th generation – Intel i5/ 8th generation (similar or equal)
41 | Effective Pixels | 16 GB DDR4
42 | RAM | 16 GB DDR4
43 | Effective Pixels | Approx. 400 × 300
44 | RAM | Approx. 400 × 300
45 | Interfaces | Approx. 17 μm
46 | Pixel Pitch | Approx. 17 μm
47 | Interfaces | USB 3.0 SuperSpeed and USB 2.0 FullSpeed, Power supply & current consumption: USB 3.0: 5 Volt ± 5%, 700 mA max. USB 2.0: 5 Volt ± 5%, 220 mA max.
48 | Pixel Pitch | USB 3.0 SuperSpeed and USB 2.0 FullSpeed, Power supply & current consumption: USB 3.0: 5 Volt ± 5%, 700 mA max. USB 2.0: 5 Volt ± 5%, 220 mA max.
49 | Spectral Range | Recommended
50 | Graphics card | Recommended
51 | Spectral Range | Approx. 8 μm–14 μm
52 | Graphics card | Approx. 8 μm–14 μm
53 | Operating system | Approx. ≤40 mK
54 | Thermal Sensitivity (NETD) | Approx. ≤40 mK
55 | Operating system | Windows 10
56 | Thermal Sensitivity (NETD) | Windows 10
Company Stamp |  |  |  | Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | MMI110012 | 0 | X-RAY DIGITAL MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 35 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | GENERAL STANDARD
2 | (FDA 510 (K) Clearance OR CE Mark) & SFDA Registration | Yes
3 | MOBILE X-RAY UNIT
4 | Microprocessor Controlled | Yes
5 | Forward & Reverse Motor Driven | Yes, Self Propelled Handle Drive Control
6 | Max incline, degree | ( 5-7 ) Degree is Preferable
7 | Speed | Adjustable from ( 0.5 - 1.3 ) m/s
8 | AC line | Yes
9 | Line-voltage Compensator | Yes / Automatic
10 | Type of Power Sources | Both AC & Battery, The Generator Should Work on AC in Case of Battery Failure (Preferable)
11 | Power Requirements | 220 Volt / 60 Hz.
12 | Single or Dual Battery System | Two Independent Systems Provide Power for Driving Motors & Imaging Exposure (Preferable)
13 | Battery Type | Lead acid sealed lead / NiCad
14 | Battery Charging Time - ( From Empty to Full Charged ) | Not more than ( 10 ) hrs
15 | Low-Battery Indicator | Yes
16 | Battery Power Storage | Specify number of maximum oprating hours
17 | Power Cord Length, m | Not less than ( 2.5-5 ) m or More
18 | Wired Handheld Switch & Cordless Remote control Switch | Yes, Exposure By Using Remote Control for Optimal Radiation Protection
19 | Manual Movement in Case of Battery or Motor Failure | Yes
20 | Remote Control Operating range | Not Less than ( 7 ) m
21 | Anti-Collision Protection System or Equivalent | Yes
22 | System Weight | Specify
23 | Mobile X-ray can be Moved Forward & Backward by using the Bedside Drive Controls Located on the Collimator or on the Articulated Arm | Preferred
24 | X-RAY TUBE
25 | Rotating Anode | Yes
26 | Anode Heat Storage Capacity | Not Less Than ( 120 ) kHU
27 | Maximum output Voltage, kVp | Approx. ( 140 ) KV , ( ±10 )
28 | Nominal Focal Spot Value, mm | ( 0.8 / IEC ) or Dual Focal Spots
29 | Nominal Anode Input power | Not Less Than ( 30 ) KW
30 | Tube Movement & Angulation | Yes
31 | Clear view while moving (telescopic movement of the column or dropdown) | Yes
32 | Shortest Exposure Time | 2 ms or less
33 | Working Column Movement Range | Rotating ( 180 ) Degree
34 | X-RAY GENERATOR
35 | High Frequency X-Ray Generator | Yes
36 | Nominal power Rating | Not Less Than ( 30 ) KW
37 | Kv Range | Approx. ( 40 - 140 ) Kv ( ±10 ) step of 5 (as minimum)
38 | mA range | Not Less Than ( 300 ) mA
39 | mAs range | Min.: ≤ 0.4 , Max.: ≥ 320
40 | Increments | Step Up / Step Down / Continuous
41 | Digital Indicator meters | Yes, (If applicable) For KVp, mAs, mA, Dose
42 | X-RAY COLLIMATOR
43 | Type | Manual Collimator
44 | Rotation | Approx. ( ±90 ) Degree
45 | Aluminum filter, mm | ≤ (3.5) mm
46 | SID Range, cm | Min.: ≤ 68 , Max.: ≥ 200
47 | Centering indicator (light /laser) | Yes, for Quick & Easy Indication of SID
48 | LARGE DETECTOR ASSEMBLY
49 | Waterproof | Yes
50 | Shock Sensor | Yes
51 | Portable Digital Detector | Wireless
52 | Semiconductor material | Amorphous silicon ( A-Si )
53 | Scintillator | Cesium iodide ( Csl )
54 | Detector Active Area Size | ≥ ( 35 x 43 ) cm, Approx.
55 | Image Depth / A/D Conversion / Acquisition Depth | ( 14 ) Bit
56 | Pixel size | 150 Micron or less is Preferred
57 | Matrix Size / Resolution | Not Less Than ( 2022 X 2022 )
58 | Detective Quantum Efficiency ( DQE ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
59 | Modulation Transfer Function ( MTF ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
60 | Detector Weight Including Battery | Less Than or equal ( 4 ) kg
61 | Max. Load Capacity On Lying Position | More Than ( 100 ) kg
62 | Rechargeable Battery (Built in or Removable) | Lithium-ion , in Case of Removable Second Battery is Required or equivalent
63 | Charging Time | Not More Than ( 4 ) hours for 100% Charge
64 | Battery Operation Time | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
65 | Charging Station | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
66 | Rigidity of Detector Housing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
67 | Detector Handle | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
68 | Detector Sharing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at (90) images per hour
69 | PROCESING & VIEWING WORKSTATION (Operator Console)
70 | Acquisition Workstation | Built in
71 | LCD Touch Screen Monitor | Not Less Than ( 15 ) inch
72 | Radiation Dose Monitoring ( DAP ) | indicate
73 | Had Disk Storage | Not Less Than ( 3000 ) images
74 | Automatic Programming ( APR ) | Yes
75 | IMAGE POST PROCESSING FUNCTIONS
76 | Window Width & Leveling | Yes
77 | Gray Scale Invert / Annotation | Yes
78 | Image Rotate / Free Rotation | Yes
79 | Electronic L/R Marker | Yes
80 | Patient Edit / Emergency Exam | Yes
81 | NETWORKING & SECURITY
82 | DICOM Compatible | Yes
83 | DICOM Store and DICOM structured dose report | Yes
84 | Image auto transfer | Yes
85 | Query, retrieve, Send, Receive | Yes
86 | Modality worklist | Yes
87 | Storage commitment | Yes
88 | DICOM print | Yes
89 | Patient edit | Yes
90 | DICOM MPPS | Yes
91 | DICOM viewer on CD "Burn exam on CD with DICOM viewer" | Yes
92 | HIPAA Patient Data Security | Yes
93 | DATA PROTECTION AGAINST DELETION & export "Different privilege levels " | Yes
94 | Security Package | Yes
95 | Virus Protection | Yes
96 | The ability to interact with the processing workstation from remote PC to view and process studies within the same hospital | Yes
97 | ACCESSORIES
98 | Two light weight Aprons (0.5) Equivalent of ( Non Leaded Material ) | ( Two ) - (medium & large)
99 | One Moveable Detector holder | ( One ) with Vertical, Horizontal & Swivel Movement
100 | Add-on specifications on top of base specs | Vendor input
101 | Extra detector (small)
102 | Extra detector (large)
103 | Above specs with telescopic column movement for X-Ray tube (if available)
104 | Extra detector (small) with exchange price
105 | Extra detector (large) with exchange price
106 | One Mobile Apron Hanger, and must hold 5 aprons
107 | Please specify Add-on specifications on top of base specs
Company Stamp |  |  |  | Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | MOP100002 | 0 | HYPERBARIC OXYGEN THERAPY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 36 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | HYPERBARIC MEDICAL TREATMENT SYSTEM | 12 Patient Chamber
2 | pressure vessel and the medical monitoring system | YES
3 | Vessel shape and size: | divided in main chamber and ante chamber
4 | Double-Lock
5 | cylindrical | YES
6 | lying pressure vessel | YES
7 | Omega shape with Flat Wall Technology with non-enforced flat steel plates as endwalls, side walls and bottom | YES
8 | Inside chamber diameter | 2500 mm
9 | Inside free height (min) | 1900 mm
10 | Seating capacity | twelve seated patients (arranged in 4 groups of 3 seats mounted on trolleys) or four lying patients on trolleys with stretchers, or 2 lying patients in ICU beds
11 | each seat should have its own breathing system by oxygen mask and oxygen hood. | YES
12 | treatment pressure | shall be 2,0 bar gauge = 3,0 bar absolute
13 | Design temperature | 50°C
14 | Pressure Vessel (Scope of Delivery):
15 | 1 Cylindrical pressure vessel, with flat end and intermediate walls and flat floor as pressure absorption components without enforcements | YES
16 | Extra low entrance step. | YES
17 | Load distribution well spread over bottom of the building | YES
18 | 2 rectangular doors | (1x between ac/mc) 1800 mm free height x 700 mm free width (1x for direct access ac) as flat bottom walk in door, easy to operate
19 | Minimum of 3 observation window in wide angle technique WINDOW | free 250 mm diameter and two rubber lip-seal, sealing with the pressure
20 | 1 medical supply lock with one hand operation technique size | free inner length 300 mm, inner diameter 200 mm
21 | 1 bayonet closure outside , flat door inside | door leave stainless-steel
22 | 1 set equalizing valves | YES
23 | 1 safety devices with pressure indication and one-hand operation | YES
24 | 4 observation-window Extremely wide vision | free diameter 250 mm, Minimum of 3 in each side walls of the main-chamber
25 | three spare flanges free diameter 100 mm, blind leafs exchangeable for medical data (2x in mc / 1x in ac) | YES
26 | 1 set of Hoisting eyes according to hoist standard | YES
27 | 1 Corrosion prevention Corrosion prevention and coat of varnish | (inside and outside) approved for use in hyperbaric systems
28 | Main Chamber Equipment:
29 | Twelve Patient seats | passenger seats style
30 | ergonomically formed high back rests | YES
31 | foldable arm rests | 2 arm rests per seat
32 | upholstery | YES
33 | covered with water resistant material that is resilient to chemicals | YES
34 | fire resistant | YES
35 | easy to clean | YES
36 | silencer Set of high-efficient silencers for air inlet and outlet | noise level will be less than 85 dBA for normal operation
37 | chamber ventilation is connected to the system | YES
38 | one Safety valve adjusted to the max working pressure
39 | minimum 8 Illumination units external cold light lamp/ optical fibre type 50 watt/each | Lamps are non-glaring, chamber is not heated up, absolutely no noise is generated, easy to maintain, with extra long-lasting halogen lamps.
40 | One communication / display | YES
41 | One permanent dynamic telephone | sound powered
42 | howler | as second redundant talking possibility)
43 | One analogue clock | YES
44 | button for emergency signal system | YES
45 | digital indication of working pressure | YES
46 | digital indication of time | YES
47 | digital indication of temperature | YES
48 | digital indication of remaining treatment time | YES
49 | Two mounting rails Suspension bars | integrated in supply rail (for holding infusion bottles etc.)
50 | one stacker plates Set, stainless steel
51 | twelve patient oxygen breathing system O2-/air breathing units | inlet and outlet, integrated in supply rail, with overboard dump system for the exhaled gas, breathing resistance 3 mbar at 1,5 bar chamber pressure and breathing volume 22,5 l/min acc.
52 | Main Chamber -Breathing units shall be able to switch collectively toO2/ Air from the control station | demand-inlet-regulators and the outlet valves are integrated in and covered by the supply rail above the patients
53 | twelve semi-masks High-quality semi-masks | made of anti allergic silicon
54 | with infinitely adjustable head harness Connection of mask and connection block is effected via standard flexible and extremely light-weight one-way corrugated hose | YES
55 | Twelve Hood-couplings | Additional connections for connecting of hood-systems
56 | Twelve Head Tent-System | YES
57 | incl. hoods via the a connection | 12x type with flowmeter
58 | coupling system | YES
59 | overboard dump system for the hoods exhaled gas | can be switched collectively from O2/Air from the control station
60 | One set of microphone and loudspeaker | trans-receiver
61 | set of sockets for entertainment-headphones for each patient | incl. 12 headsets one for each patient
62 | luding set of audio-channel selection switches for each patient | YES
63 | pipe-work and equipment is recessed and hidden behind supply rails or panels | YES
64 | call button for each patient | YES
65 | distribution channels set of distribution channels with outlet gratings for air conditioning system | YES
66 | One emergency evacuation valve | leaded
67 | Sufficient fire extinguisher Portable fire extinguisher filled with water | use in hyperbaric system
68 | Anti-slip floor, easy to clean | wet cleaning
69 | fire resistant and antistatic material proved for use under hyperbaric conditions | YES
70 | painting Inner painting | non-toxic, hardly inflammable/fire retardant and non-smelling
71 | Paint colours light white | YES
72 | cabinet to store medical appliances, rugs etc. | cover made of stainless steel
73 | Ante Chamber Equipment:
74 | Two single seats with arm and back rests for seating persons | YES
75 | silencer Set of high-efficiency silencers for air inlet and outlet | noise level will be less than 85 dBA for normal operation
76 | The chamberventilation | connected to the system
77 | emergency evacuation valve | leaded
78 | Two Illumination unit external cold light lamp/ optical fibre type | 50 watt/each, Lamps are non-glaring, chamber is not heated up, absolutely no noise is generated, easy to maintain, with extra long-lasting halogen lamps.
79 | communication / display system with one set of microphone and loudspeaker (transreceiver) System | YES
80 | permanent dynamic telephone (sound powered) with howler (as second redundant talking possibility) | YES
81 | analogue clock | 100
82 | depth gauge, scale 0-60 msw to indicate the depth inside ac | YES
83 | Two patient oxygen /air breathing system, new O2-breathing-units | inlet and outlet, integrated in supply rail
84 | with overboard dump system for the exhaled gas | YES
85 | breathing resistance 3 mbar at 1,5 bar chamber pressure and breathing volume 22,5 l/min | YES
86 | AC Breathing units can be switched collectively (together with mc) to O2/ Air from the control station | YES
87 | the demand-inlet-regulators and the outlet valves | integrated in and covered by the supply rail above the patients
88 | Two semi-masks High-quality semi-masks | made of anti allergic silicon
89 | with infinitely adjustable head harness | YES
90 | Connection of mask and connection block is effected via flexible and extremely light weight one way corrugated hose | YES
91 | One supply rail above the patient in which are integrated | One set of microphone and loudspeaker (trans-receiver) System
92 | One fire extinguisher Portable fire extinguisher filled with water | specially use in hyperbaric system
93 | piping, wiring Set of material for piping and wiring | YES
94 | stacker plates (equipment tray) Set, stainless steel | YES
95 | floor Anti-slip floor, easy to clean | wet cleaning
96 | fire resistant and antistatic material | proved for use under hyperbaric conditions
97 | CHAMBER CONTROL SYSTEM | YES
98 | Medical Monitoring system | important operating and health related parameters for each patient and for the chamber itself.
99 | TV-MONITORING SYSTEM:
100 | four TV cameras CCD-colour TV cameras with wide angle lens for video monitoring | 3x MC and x AC) In the main chamber 4 patients are observed by one camera. The wide angle lens is of fixed/adjustable type
101 | four camera- adjustment Fixing bows with adjustment possibility | YES
102 | 1 monitor Color-TV-TFT Flat 17” | Screen Monitor for display of one selected camera-picture
103 | 1 monitor Color-TV-TFT Flat 17” | Screen Monitor for Quad display
104 | max 4 camera pictures on one screen as well as selectable large image display | YES
105 | 1 Quad-processor Set E- and control technique for TV-monitoring system | YES
106 | 1 DVD-Recorder as video recorder | records the pictures which are displayed on a selected monitor.
107 | 1 set E- and control technique for TV-monitoring system | YES
108 | TRANSFER TROLLEY | Two TRANSFER-TROLLEY for lying patients, inclusive 2 berth with mattress and upholstery, up-foldable head rest
109 | COOL SYSTEM / HEAT SYSTEM | Environment Control System to control the temperature inside the chamber shall be installed to regulate temperature during operation as well the relative humidity inside the chamber
110 | FIRE EXTINGUISHING SYSTEM | Fire extinguishing system shall be installed to meet the highest standard for multiplace hyperbaric chamber systems.
111 | The main-chamber and the ante-chamber must be equipped with a stationary built in fire-fighting system | YES
112 | COMPRESSED-AIR SUPPLY (LP+HP-System) | The compressed-air supply system SHOULD meet the requirements for pressure chambers for hyperbaric treatment and is thus designed according to the valid safety standards.
113 | HP-COMRESSED AIR STORAGE system | consist of 2 cylinder-battery-halves of minimum 16 pieces each with 80 l/200 bar compressed air cylinders, should be arranged in one supply rack. HP-batteries and low pressure supply are united in a control panel to which also the high pressure
114 | HP-COMRESSED AIR STORAGE system | compressor
Company Stamp |  |  |  | Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | MOR20017 | 0 | INTUBATION DIFFICULT VIDEO |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 37 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Video Laryngoscope For difficult airway intubation
2 | Design | light weight & portable
3 | The system shall be based on the latest technological advances in difficult intubations | YES
4 | The unit shall be designed to provide a clear view of the vocal cords during intubation | YES
5 | The unit shall simplify intubation of the difficult airway | YES
6 | Trauma Airway | Yes
7 | Range in blade sizes allow coverage patient weights | Yes
8 | It shall have 60° angled blades | YES
9 | It should have pre-shaped stylet for quick intubation process | YES
10 | It should utilize minimum force for intubation | YES
11 | It should have anti-fog feature | YES
12 | The unit shall have integrated high resolution camera at good Position on blade to protect from blood and secretion | YES
13 | The unit shall mount on mobile carts and castors | YES
14 | Intubation of cervical spine immobilization | Yes
15 | Difficult Airway management and routine intubation | Yes
16 | Screen Size not less than 7 inch | LCD OR TFT
17 | Blades Size
18 | SIZE 2 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
19 | SIZE 2 SINGLE USE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
20 | SIZE 3 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
21 | SIZE 3 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
22 | SIZE 4 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
23 | SIZE 4 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
24 | Made of medical grade plastic or stainless steel for less trauma handlingo teeth and soft tissue and easy | Specify
25 | Built in LED light for illumination | Yes
26 | Rechargable Battery on system | Yes
27 | Stylet | Qty:3
28 | Accessories | Any accessories, options and consumable items necessary to operate the offered system(s) must be clearly identified and priced separately
Company Stamp |  |  |  | Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | MOR20019 | 0 | LARYNGOSCOPE SET COMPLETE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 38 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE, FIBER OPTIC | Yes
2 | Macintosh blades | Yes
3 | Adult | 3, 4
4 | pediatric | 2
5 | Miller blades for sizes 00, 0, 1 for neonatal and infant | Included
6 | Handle | 2
7 | ILLUMINATION | LED
8 | Spare lamps | 5
9 | Materials of handle | Stainless steel
10 | Cleaning and sterilization for handle and blades | specify
11 | Case | Included
12 | Power supply AA battery | Yes
13 | Charger with rechargeable battery | Option
14 | Itemized Price for charger with battery | Yes
15 | Itemized Price for Lamp | Yes
16 | Itemized Price for handle | Yes
17 | Itemized Price for all blades | Yes
18 | All Accessories (itemized price) | Yes
Company Stamp |  |  |  | Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | MOR30006 | 0 | ECMO MACHINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 39 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | ECMO UNIT
2 | LUNG TYPE TECHNOLOGY | SPECIFY
3 | SURFACE AREA (m2) | 2.5 - 4.5
4 | PRIMING VOLUME (ml) | 270 - 450
5 | O2 TRANSFER ( Ml/MIN ) | 375 - 400
6 | BLOOD PHASE PRESSURE DROP AT RATED FLOW (mmHg) | SPECIFY
7 | C02 TRANSFER ( Ml/MIN ) | 200 - 300
8 | MAXIMUM GAS FLOW RATE (L/min) | 14
9 | MAXIMUM BLOOD FLOW (L/min) | 7
10 | COMPLIANCE (Ml /100mmHg) | LOW : SPECIFY
11 | O2 TRANSFER /m2 ( Ml/M2/MIN ) | 90 - 150
Company Stamp |  |  |  | Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | MWD10003 | 0 | BED ELECTRICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 40 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Fully electrically motorized | Yes
2 | Electric variable height | Yes
3 | Electric sliding backrest | Yes
4 | Trend and reverse Trend | Yes
5 | Hand control ( remote wire control ) | Yes
6 | The motor should be heavy-duty type and protected for over heat. | Yes
7 | Removable headboard for CPR | Specify
8 | Instant CPR release at sides | Yes
9 | MATTRESS | Included
10 | The mattress should be supported by steel frame or other | Yes
11 | Shall be fire retardant, conductive and disinfectable layer | Yes
12 | Length: 200 cm Width - 85cm Mattress Thickness 12 cm | Specify
13 | Easy to clean and extremely hygenic | Yes
14 | manufacturer material for mattress | Yes
15 | SIDE RAIL
16 | up and down movement | Yes
17 | Split Siderails with integrated controls | Yes
18 | side rail length | Specify
19 | manufacturer material for side rail | Yes
20 | OVERALL LENGTH
21 | Fully extended -240.0cm approx. | SPECIFY
22 | Fully retracted - 210.0cm approx. | Specify
23 | OVERALL WIDTH
24 | Siderails stowed -90.0cm approx. | Yes
25 | Siderails up -100cm approx. | Yes
26 | SLEEP SURFACE
27 | Length - fully extended -210cm approx. | SPECIFY
28 | Length - fully retracted -190cm approx. | Yes
29 | Head and foot boards | Yes
30 | Night-light | Yes
31 | Adjustable height from 38 to 75 cm apprpx. | Specify
32 | Under bed frame clearances | Yes
33 | IV rod (all 4 bed corners) | Yes
34 | Head elevation gauge | Specify
35 | Drainage bag holders . | Yes
36 | Bed shall hold 220 kg working load approx. | Yes
37 | 4 Corner bumpers and accessory holders | Yes
38 | Battery backup or Manual backup | Yes
39 | The unit should have no sharp edges. All external components should be securely mounted | Yes
40 | The unit should be secure and provide adequate protection against moving and electrically energized parts | Yes
41 | The unit should be well constructed with durable materials to withstand typical abuse and cleaning | Yes
42 | The unit should be simple to use, operate, and maintain | Yes
43 | Bed extender | Yes
44 | 4 Casters | Specify
45 | Central brake and steer | Yes
46 | Brake and steer pedals accessible from all four corners of the bed | Yes
47 | Easy maneuverability | Yes
48 | Power supply 220 VAC / 60 HZ | Yes
49 | FDA and CE approved | Yes
Company Stamp |  |  |  | Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | MWD10006 | 0 | BED INTENSIVE CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 41 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Intensive Care Bed/ ICU adult
2 | PATIENT CONTROLS | YES
3 | Type | specify
4 | Functions | Head High/ Low , Knee up / Down, Trendelenburg & Reverse trendelenburg, electrical flexafoot
5 | NURSE CONTROLS | YES
6 | Patient control lockout | YES
7 | Full low indicator | YES
8 | CPR MODE | Dual sided
9 | CPR control | YES
10 | graphical caregiver interface with touch screen for bed and mattress functions | YES
11 | AUTOMATIC CONTOUR | YES
12 | CHAIR MODE | Full Chair & Full Chair front Egress Position
13 | ALARM SYSTEM | Head angle alarm 30 and 45 degrees, bed Exit alarm, Brake off alarm
14 | TRENDELENBURG GAUGE | YES
15 | BackUp Function | Electro-hydraulic or other specify
16 | BED SURFACE TYPE | specify
17 | C-ARM APPLICATION | YES
18 | FRACTURE FRAME RECEPTACLES | YES
19 | ELECTRIC FOOT EXTENDER / RETRACTABLE | Yes / Electrical
20 | DRAINAGE BAG HOLDER | YES
21 | MAX. PATIENT WEIGHT, Kg | 230 kg or more
22 | OVERALL DIMENSIONS
23 | L x W, cm | L210 235x W93 -102 cm approx.
24 | Height, cm | 45 cm – 92 cm approx
25 | SIDE RAILS LENGTH
26 | Fraction of overall length | YES specify
27 | CASTERS
28 | SIZE (Diameter) cm | 12.5 cm or more specify
29 | FUNCTION | Swivel, brake, steer
30 | CENTRAL BREAK SYSTEM | YES
31 | BUMPERS | Corner/ Full perimeter
32 | REMOVAL HEADBOARD & FOOTBOARD | YES
33 | IV POLE MOUNT | MINIMUM 3 PROVISIONS
34 | INTEGRATED BED SCALE With patient weight statistics for up to7 days or more | YES
35 | Power supply | 220 Volts, 60 Hz
36 | Battery (Back-up) | OPTIONAL ITEMIZED
37 | Double insulation | YES
38 | Electric Shock Protection | Class 1, Type B
39 | Number of motors | 4 or more specify
40 | GROUNDED | YES
41 | Isolated (motor ground | YES
42 | Non-Conductive Side rails | YES
43 | AIR MATTRESS | Anti-bacterial, Anti-static, Moisture proof, fire resist & chemical-resist, tough & high durable
44 | INTEGRATED WITH THE BED | YES
45 | Dimensions 20cm or more height | Compatible with bed dimensions
46 | Controls are imbedded within the side rails | YES
47 | AUTOMATIC DEFLATE | Full inflation once CPR mode is pressed
48 | Max inflate function | YES
49 | X-ray Sleeve inside the mattress | YES
50 | other specifications | FDA approved CE marked
51 | Printout for the item offered with all technical parameters exactly matches the soft copy in the CD. | YES
52 | Supplier is an authorized representative of the manufacturer by SFDA and/or an official agent by Ministry of Trade. | Yes (Attach official documents)
53 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | MWD10023 | 0 | STRETCHER TRANSPORT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 42 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | AREAS OF USE | Emergency / Re-covery bay
2 | FRAME MATERIAL | Epoxy painted steel
3 | STRETCHER TYPE | Hydraulic
4 | FOOT PEDAL OPERATED | Yes, both sides
5 | HEIGHT ADJUSTMENT | Yes
6 | TRENDELENBURG | 15 degree approx.
7 | REVERSE TRENDELENBURG | 15 degree approx.
8 | HEAD ADJUSMENT | Yes
9 | STRETCHER SURFACE | Radio translucent
10 | X-RAY CASSETTE HOLDER | Yes, full length
11 | SIDERAILS | Yes
12 | OXYGEN TANK HOLDER | Yes
13 | RESTRAINING STRAPS | Yes
14 | BUMPERS | Yes, all corners
15 | IV POLE SOCKETS
16 | CASTORS | 4 or 5, lockable
17 | DIAMETER | 20 cm approx.
18 | MAXIMUM PATIENT WEIGHT | 220 Kg
19 | ACCESSORIES | All neceessory accessories
20 | MATTRESS | Yes, durable and easy to clean
21 | PLATFORM DIMENSION IN MM | 2000 X 800 approx.
22 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 42
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
42 | MWD10034 | 0 | VENTILATOR BIPAP IN-PATIENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 43 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit must be the latest technology in noninvasive ventilation. | YES
2 | The unit must be Positive Airway Pressure machine to be used in healthcare facilities. | YES
3 | The unit must be used for adult and pediatric patients. | YES
4 | The unit must have the following controls: | YES
5 | Inspiratory time 0.5 to 3 sec. | YES
6 | Inspiratory pressure 4 to 30 cm H2O. | YES
7 | Expiratory pressure 4 to 25 | YES
8 | Respiratory rate 5 to 40 bpm. | YES
9 | The unit must have the following modes: | YES
10 | CPAP | YES
11 | BiPAP: Spontaneous mode, Timed mode and Spontaneous-Timed mode. | YES
12 | BiPAP with back up rate. | YES
13 | Pressure Assist Control. | YES
14 | Volume Assured Pressure Support with Backup Rate | YES
15 | Volume Assured Pressure Support with Automatic Expiratory Pressure Adjustment. | YES
16 | The unit must have different control for an inspiratory time, expiratory time, PIP, Maximum PIP, Respiratory rate, CPAP and EPAP, trigger sensitivity, and screen lock. | YES
17 | The unit must be equipped with auto patient start and stop. | YES
18 | The unit must have an adjustable ramp time with a range of 0 to 45 min approximately. | YES
19 | The unit should have a built-in filter. | YES
20 | The unit must have low-pressure oxygen inlet. | YES
21 | The unit must have leak compensation for at least 100 L/M. | YES
22 | The unit must have a built-in heated humidifier with temperature adjusted. | YES
23 | They unit should include humidifier blocker to use the device without humidifier. | YES
24 | The unit must have a digital waveform display and waveform software. | YES
25 | The unit should have memory port with memory card included. | YES
26 | The unit should have a carrying case. | YES
27 | The unit must be installed on a movable stand. | YES
28 | The unit should be easy to operate. | YES
29 | The unit should have event log. | YES
30 | The controls, labels, and displays should be clear and visible. | YES
31 | The controls should be protected against accidental setting changes. | YES
32 | The unit should be small in size and lightweight with approximate weight less than 4Kg with humidifier. | YES
33 | The unit should be able to resist tipping over. | YES
34 | Alarms should allow quick assessment and correction of the alarm condition. | YES
35 | The priority of the alarm should be indicated by different audible tones and visual indicators. | YES
36 | Audible alarms and visual indicators must activate when switching from line to battery power. | YES
37 | The unit must also have power failure, overheat and fault alarms. | YES
38 | The following alarms should also be included | YES
39 | Leak alarm | YES
40 | Low and High pressure | YES
41 | Respiratory rate | YES
42 | Apnea alarm is a must | YES
43 | BiPAP units should be able to operate in a variety of environments and should not be affected by electromagnetic interference and electrostatic discharge. | YES
44 | Servicing by a skilled technician should be easy; the operator’s manual should provide adequate information for clinicians, users, and caregivers. | YES
45 | The unit should be well constructed with durable material to withstand typical abuse & cleaning. | YES
46 | The unit should be FDA approved and CE marked. | YES
47 | The unit should comply with International Respiratory Standards. | YES
48 | The unit should be 220VAC at 60Hz. | YES
49 | All available options, consumables, or disposables deemed necessary for the intended function of the equipment and not listed in the Technical Specifications should be included and specified. | Provide Itemized Price
Company Stamp |  |  |  | Signature