## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | BOX COLD 8 LITERS
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | ULTRASOUND POC ADVANCE PORTABLE
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | CHAIR BLOOD DONATION
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | CHAIR BLOOD DONATION PORTABLE
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | REFRIGERATOR BLOOD BANK 360 BAGS
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | REFRIGERATOR BLOOD BANK 500 BAGS
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | BLOOD BAG ROTATING MIXER
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | BINOCULAR MICROSCOPE DUAL OBSERVATION
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | CENTRIFUGE FLOOR
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | CENTRIFUGE REFRIGERATED TABLE TOP
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | CHAIR BLOOD DRAWING
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | ANALYZER AUTOMATED ELISA
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | INCUBATOR CO2 WITH O2
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | SHELVING PHARMACY

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | MGE40012 | BOX COLD 8 LITERS |  | Medical Eqpt | 10 | [10 empty] | 0
2 | MIC10054 | ULTRASOUND POC ADVANCE PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
3 | MLB100004 | CHAIR BLOOD DONATION |  | Medical Eqpt | 2 | [10 empty] | 0
4 | MLB100005 | CHAIR BLOOD DONATION PORTABLE |  | Medical Eqpt | 10 | [10 empty] | 0
5 | MLB100010 | REFRIGERATOR BLOOD BANK 360 BAGS |  | Medical Eqpt | 2 | [10 empty] | 0
6 | MLB100011 | REFRIGERATOR BLOOD BANK 500 BAGS |  | Medical Eqpt | 2 | [10 empty] | 0
7 | MLB100022 | BLOOD BAG ROTATING MIXER |  | Medical Eqpt | 10 | [10 empty] | 0
8 | MLB10019 | BINOCULAR MICROSCOPE DUAL OBSERVATION |  | Medical Eqpt | 1 | [10 empty] | 0
9 | MLB10039 | CENTRIFUGE FLOOR |  | Medical Eqpt | 2 | [10 empty] | 0
10 | MLB10044 | CENTRIFUGE REFRIGERATED TABLE TOP |  | Medical Eqpt | 2 | [10 empty] | 0
11 | MLB10045 | CHAIR BLOOD DRAWING |  | Medical Eqpt | 6 | [10 empty] | 0
12 | MLB40003 | ANALYZER AUTOMATED ELISA |  | Medical Eqpt | 1 | [10 empty] | 0
13 | MLB40023 | INCUBATOR CO2 WITH O2 |  | Medical Eqpt | 1 | [10 empty] | 0
14 | MPH50004 | SHELVING PHARMACY |  | Medical Eqpt | 4 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | MGE40012 | 0 | BOX COLD 8 LITERS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CAPACITY: 8L
2 | BOX DIMENSION WxDxH
3 | Outer Dimension
4 | Inner Dimension
5 | 20 Hours of battery life
6 | External LED temperature display
7 | OUTER MATERIAL
8 | INNER MATERIAL
9 | INSULATION POLYTHERINE
10 | INSULATION STRENGTH (mm)
11 | LID
12 | Carrying Strap
13 | Power Supply 220VAC/60Hz .
14 | Cold Storage Temperature 2℃ - 8℃
15 | FDA or CE approved .
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | MIC10054 | 0 | ULTRASOUND POC ADVANCE PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Regulation Standard :
2 | FDA 510 (K) Clearance / CE Mark & SFDA Registration | Yes
3 | Validated Up to Date Certificates Must be Submitted with the Offer | Yes, Failure of Submission and Comply Will Disqualify the Offer
4 | Product Description :
5 | Range | Laptop Ultrasound
6 | Purpose | Used for point of care General Clinical Routine Applications.
7 | Console Design :
8 | Laptop / Hand Carried on height adjusted mobility Cart | Cart Should Include the Following :Integrated Power SupplyProbe Portswheels with locking braking system
9 | Total Weight - (incl. battery) | ≤ ( 5 ) kg .
10 | Esay Acess mode keys | Yes, (Specify)
11 | Screen Monitor | ≥ (12) inch. LCD with High Resolution
12 | Image To Screen Ratio @ highest Resolution | >=75%
13 | Security Lock | Yes
14 | Built-In Battery | Yes
15 | Battery scanning time | 1 hour
16 | Power Requirement | 220-240 V / 60 Hz
17 | Computing Capabilities :
18 | Operation System | Windows or equivalent
19 | Integrated Hard Desk Type | Flash memory≥15 GB for better shock Absorbance
20 | Fast System boot up | Yes , Less than 30 Sec
21 | DVD /CD burner/USB port | Yes
22 | Clinical Applications :
23 | Radiology, abdominal, Small parts | Yes
24 | OB / Gyn, Breast | Yes
25 | Musculoskeletal | Yes
26 | Vascular | Yes
27 | Urology | Yes
28 | pediatrics Radiology | Yes
29 | Other Available Applications | Yes
30 | Multi Frequency Probes : | One Each , With dedicated Operational Software
31 | Linear array | 3 -12 MHz (±2)
32 | Convex array | 2 - 5 MHz (±2)
33 | Phased array (TCD) | 1 - 5 MHz (±2)
34 | Active Probe Ports | 2-3 Active Ports
35 | Operating Modes :
36 | Brightness Mode (B-Mode) (2D) | Yes
37 | Motion Mode (M-Mode) | Yes
38 | Color Doppler mode | yes
39 | Pulsed wave Doppler (PW) with HPRF | Yes
40 | High PRF Doppler Mode | Yes
41 | Color Flow Doppler Mode (CFM) | Yes
42 | Power Doppler Mode | Yes
43 | Continuous Wave Doppler | Yes
44 | Doplex Mode | Yes
45 | Triplex Mode | Yes
46 | Scanning Parameters :
47 | Imaging Depth, cm | Not less than 30 cm
48 | Dynamic Range | Up to 165 dB
49 | TGC Control | Yes
50 | Frame Rate 2D | Specify Range
51 | Frame rate Color Doppler | Specify Range
52 | GrayScale Levels | 256 ( 8 bits )
53 | Scanning Automatic Optimization | Capability for Auto optimization in 2D & Doppler
54 | Digital Calipers | Yes
55 | Preprocessing | Yes
56 | Postprocessing | Yes, (Measurment, adjust the Daynamic Range, TGC, color map & Gain) on stored images
57 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
58 | Doppler :
59 | Type | PW in all probes
60 | Frequency Display | Yes
61 | Doppler baseline shift after freeze | Yes
62 | Velocity Display | Yes
63 | User – defined annotation | Yes
64 | GrayScale Levels | Yes
65 | Transmit focal zones | Select between 1 to 8
66 | Digital reconstructed zoom | 16 levels
67 | Gain Control after freeze | Yes
68 | Adaptive Gain Compensation | Yes
69 | Directional color power Doppler | Yes
70 | Cine Features :
71 | Cine Storage | Yes, (Still images / Clips / Volume)
72 | Length / Capacity | 140 MB (up to 3000 images)
73 | Scanning Features :
74 | Speckle Reduction Imaging | Yes, Organ Specific
75 | Trapeziodal Imaging / Virtual Convex | Yes
76 | B-Mode Steering | Yes
77 | Compound Imaging | Yes
78 | (Panoramic Imaging) - Free Style Extended Imaging | Yes, Min 30 cm
79 | Tissue Harmonic Imaging | Yes
80 | TGC (Time Gain Compensation) | Yes
81 | Color flow mapping | Standard
82 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
83 | Automatic Doppler Calculations | Yes
84 | Adjustable Transmit Focus | Yes
85 | Dynamic Receive Focus | Yes
86 | Selectable Dynamic Range | Yes
87 | Advance Needle Visualization | Yes
88 | B - Color For Color Coding of Grey Levels | Yes
89 | Automatic Frequency and focal zone optimization with depth adjustment | Yes
90 | Zoom / Pan :
91 | Real- time image | Yes
92 | Frozen image | Yes
93 | Connectivity :
94 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicaple
95 | Wireless Connection | Yes
96 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
97 | Storage and Playback of M-mode, PW Spectral Doppler. | yes
98 | Recalculation for stored images | Yes
99 | Security / Privacy :
100 | Accessories :
101 | Secondary Battery | Yes
102 | Carrying Case | Yes
103 | Requirement :
104 | Supplier is an authorized representative of the manufacturer by SFDA | Yes (Attach official documents)
105 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | MLB100004 | 0 | CHAIR BLOOD DONATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit must be securely balanced in all operational positions.
2 | The unit must be constructed with easily sanitized materials
3 | Suitable for right or left hand use.
4 | The blood-drawing chair shall have two skin foam cradle armrests located at both sides.
5 | Arm rest adjusts to any size/height patient.
6 | The armrest should lock into place at the desired location.
7 | Arm extensions must provide visibility of the elbow region, be large enough to securely support venipuncture, and be easily retracted in case of patient collapse
8 | Adjustable leg cushion with self blocking mechanism
9 | The chair should allow for rapid and easy adjustment from sitting to reclining, supine and Trendelenburg position - in the event of the patient fainting or going into shock
10 | The unit shall have power adjustable (electric motors) movements to allow safe, comfortable and precise patient positioning at the press of a button on the hand control unit.
11 | Capability to carry up to 200 Kg
12 | The unit casters should allow easy repositioning and lock to secure safety in operation
13 | The unit shall have non slip rubber feet
14 | Color shall be coordinated with ID concept / Engineer.
15 | All surfaces should be easily cleaned and sanitized, to have:
16 | Powder-coated frame for durable, high quality, finish
17 | Premium medical grade treated vinyl upholstery
18 | Replaceable clear protective cover over foot area
19 | Upholstery should possess superior bacterial protection making it ideal for healthcare applications.
20 | Compliance with standards & legislation:
21 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
22 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
23 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
24 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
25 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | MLB100005 | 0 | CHAIR BLOOD DONATION PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit must be securely balanced in all operational positions.
2 | The unit must be constructed with easily sanitized materials
3 | Suitable for right or left hand use.
4 | The blood-drawing chair shall have two skin foam cradle armrests located at both sides.
5 | Arm rest adjusts to any size/height patient.
6 | The armrest should lock into place at the desired location.
7 | Arm extensions must provide visibility of the elbow region, be large enough to securely support venipuncture, and be easily retracted in case of patient collapse
8 | Adjustable leg cushion with self blocking mechanism
9 | The chair should allow for rapid and easy adjustment from sitting to reclining, supine and Trendelenburg position - in the event of the patient fainting or going into shock
10 | The unit shall have power adjustable (electric motors) movements to allow safe, comfortable and precise patient positioning at the press of a button on the hand control unit.
11 | Capability to carry up to 150 Kg
12 | The unit casters should allow easy repositioning and lock to secure safety in operation
13 | The unit shall have non slip rubber feet
14 | Color shall be coordinated with ID concept / Engineer.
15 | All surfaces should be easily cleaned and sanitized, to have:
16 | Powder-coated frame for durable, high quality, finish
17 | Premium medical grade treated vinyl upholstery
18 | Replaceable clear protective cover over foot area
19 | Upholstery should possess superior bacterial protection making it ideal for healthcare applications.
20 | Wheel | 4 castors required , 2 lockable
21 | Compliance with standards & legislation:
22 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
23 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
24 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
25 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
26 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | MLB100010 | 0 | REFRIGERATOR BLOOD BANK 360 BAGS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
ert |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | ertyrwty | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Free standing single door refrigerator with sufficient storage for at least 300 450ml blood bags, or about750 L
2 | Temperature:
3 | Independent, adjustable microprocessor-based temperature control system set at 4oC.
4 | Monitor that will display chamber temperature to the tenth degree Celsius.
5 | A seven day temperature chart recorder.
6 | Alarms:
7 | Factory-set high/low alarms that can be reprogrammed. High and Low limits will be displayed.
8 | Audible and visual alarm to alert for power loss and/or temperature deviation beyond set limits.
9 | Audible alarms will not be defeatable or will have a corresponding visual indicator during a temporary alarm silence period.
10 | Power failure alarm, and Remote Alarm contacts.
11 | A method to manually test alarms.
12 | Cabinet Construction:
13 | Heavy gauge stainless steel interior.
14 | Pass-Thru dimensions that fit through standard doorways.
15 | Locking casters.
16 | At least six stainless steel rollout drawers with dividers. Durable, solid bottom and liquid tight stainless steel drawers for containment of spills (four inch depth).
17 | Tube-Guard drawer shields minimize tubing from catching in the drawer slides.
18 | Sure-Seal self-closing door system eliminates the problem of broken door latches (includes standard key lock).
19 | Heated dual-pane, insulated glass door removes fogging and condensation.
20 | Interior cabinet bottoms formed to contain spills for easy clean up.
21 | Interior fluorescent light.
22 | Security lock.
23 | Refrigeration System :
24 | Hermetically sealed, air-cooled compressor.
25 | Temperature probes monitor condenser and ambient temperature.
26 | Non-CFC R134a refrigerant.
27 | Forced air circulation maintains chamber uniformity of +/-1°C and provides quick recovery.
28 | Interior fans shut down when door is opened.
29 | Automatic condensate evaporation system.
30 | Refrigerator should meet the standards established by the AABB, ANRC, FDA and BSI for whole blood storage.
31 | Operator safety and system performance should not be adversely affected by fluid spills.
32 | The unit should have no sharp edges. All external components should be securely mounted.
33 | The unit should be secure and provide adequate protection against moving and electrically energized parts.
34 | The unit should be well constructed with durable materials to withstand typical abuse and cleaning.
35 | Switches, knobs, and other controls should be designed for conditions of heavy use.
36 | Connections should be secure to resist accidental disconnection and should maintain sterility, when appropriate.
37 | The controls (e.g., switches, knobs) should be visible and clearly identified, and their functions should be self-evident.
38 | Device design should prevent misinterpretation of displays and control settings.
39 | Switches and controls should be protected against accidental setting changes (i.e., due to someone brushing against the panel).
40 | Controls should be sealed against penetration of liquids.
41 | The unit should be simple to use, operate, and maintain.
42 | The unit should be easy to clean, disinfect, and/or sterilize, as appropriate
43 | Compliance with standards & legislation:
44 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
45 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
46 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
47 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
48 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | MLB100011 | 0 | REFRIGERATOR BLOOD BANK 500 BAGS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Free standing double door refrigerator with sufficient storage for at least 360 x 2 (720 total) units of blood, or about 1500 L
2 | Refrigeration System
3 | Hermetically sealed, air-cooled compressor.
4 | Temperature probes monitor condenser and ambient temperature.
5 | Non-CFC R134a refrigerant.
6 | Forced air circulation maintains chamber uniformity of +/-1°C and provides quick recovery.
7 | Interior fans shut down when door is opened.
8 | Automatic condensate evaporation system.
9 | Temperature:
10 | Independent, adjustable microprocessor-based temperature control system set at 4oC.
11 | Monitor that will display chamber temperature to the tenth degree Celsius.
12 | A seven‑day temperature chart recorder.
13 | Alarms:
14 | Factory-set high/low alarms that can be reprogrammed. High and Low limits will be displayed.
15 | Audible and visual alarm to alert for power loss and/or temperature deviation beyond set limits.
16 | Audible alarms will not be defeatable or will have a corresponding visual indicator during a temporary alarm silence period.
17 | Power failure alarm, and Remote Alarm contacts.
18 | A method to manually test alarms.
19 | Cabinet Construction:
20 | Heavy gauge stainless steel interior.
21 | Pass-Thru dimensions that fit through standard doorways.
22 | Locking casters.
23 | At least six stainless steel rollout drawers with dividers. Durable, solid bottom and liquid tight stainless steel drawers for containment of spills (four inch depth).
24 | Tube-Guard drawer shields minimize tubing from catching in the drawer slides.
25 | Sure-Seal self-closing door system eliminates the problem of broken door latches (includes standard key lock).
26 | Heated dual-pane, insulated glass door removes fogging and condensation.
27 | Interior cabinet bottoms formed to contain spills for easy clean up.
28 | Interior fluorescent light.
29 | Security lock.
30 | Refrigerator should meet the standards established by the AABB, ANRC, FDA and BSI for whole blood storage.
31 | Operator safety and system performance should not be adversely affected by fluid spills.
32 | The unit should have no sharp edges. All external components should be securely mounted.
33 | The unit should be secure and provide adequate protection against moving and electrically energized parts.
34 | The unit should be well constructed with durable materials to withstand typical abuse and cleaning.
35 | Switches, knobs, and other controls should be designed for conditions of heavy use.
36 | Connections should be secure to resist accidental disconnection and should maintain sterility, when appropriate.
37 | The controls (e.g., switches, knobs) should be visible and clearly identified, and their functions should be self-evident.
38 | Device design should prevent misinterpretation of displays and control settings.
39 | Switches and controls should be protected against accidental setting changes (i.e., due to someone brushing against the panel).
40 | Controls should be sealed against penetration of liquids.
41 | The unit should be simple to use, operate, and maintain.
42 | The unit should be easy to clean, disinfect, and/or sterilize, as appropriate
43 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
44 | Compliance with standards & legislation:
45 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
46 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
47 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
48 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
49 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | MLB100022 | 0 | BLOOD BAG ROTATING MIXER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | USE FOR GENTLE AND UNIFORM MIXING OF BLOOD SAMPLES, CELL SUSPENSIONS, DYE FIXATIONS AND SOLUTIONS IN TEST TUBES, FLASKS, BOTTLES OR IN BLOOD BAGS. | YES
2 | 3 DIMENSIONAL ROCKING MOTION, 3D ROTATION MOVEMENTS | YES
3 | LOAD CAPACITY : ~ 2-4 KG | YES
4 | OPERATING TEMPERATURE RANGE : ~ +4 TO +45 °C | YES
5 | CAN MIX BLOOD SAMPLES IN BLOOD BAGS | YES
6 | CONTROLS SHALL HAVE : | YES
7 | FIXED-SPEED | YES
8 | FIXED-SHAKING ANGLE | YES
9 | ADJUSTABLE -SPEED ANGLE | YES
10 | RUBBER MAT SHALL BE SUITABLE FOR DIFFERENT TUBES SIMULTANEOUSLY AND SUITABLE FOR BLOOD BAGS | YES
11 | DIGITAL TIMER UP TO 99 MINUTES DISPLAY AND ADJUSTABLE | YES
12 | SPEED : 25 TO 30 RPM | YES
13 | PLATFORM SIZE SHALL BE AROUND W: ~23-30 CM , D: ~17-25 CM | YES
14 | STRAPS FOR BLOOD BAGS | YES
15 | TILT ANGLE : 15° - 20° | YES
16 | SHAKING ANGLE: HORIZONTAL TO 20◦ | YES
17 | COMPLETE STANDARD ACCESSORIES | YES
18 | STACKING PLATFORM WITH FLAT MAT | YES
19 | AUTOCLAVABLE DIMPLED MAT | YES
20 | 220 - 240VAC 60HZ WITH HOSPITAL STANDARD PLUG | YES
21 | 5 YEARS WARRANTY ON PARTS AND LABOR. AGENT MUST PROVIDE WARRANTY DOCUMENT UPON THE DELIVERY OF THE UNIT. | YES
22 | OPERATION, SERVICE, SCHEMATIC & SPARE PARTS LIST MANUALS AND CD COPIES.(BOTH SOFT AND HARD COPY) | YES
23 | ON SITE TRAINING FOR THE END-USER AND CLINICAL ENGINEERING STAFF | YES
24 | SFDA APPROVAL | YES
25 | CALIBRATION CERTIFICATE | YES
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | MLB10019 | 0 | BINOCULAR MICROSCOPE DUAL OBSERVATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Objective Lenses:
2 | Reversed facing sextuple objective turret revolving nosepiece holds Lead-Free
3 | Color Corrected Infinity System (CCIS) EC-H Plan Achromat objectives 4x, 10x, 40xr, 100xr
4 | Optional 2x, 20x and 60x
5 | Eyepieces:
6 | Siedentopf wide-field 10x eyepieces
7 | FN 22mm
8 | 30° inclined binocular tubes with interpupillary distance adjustments from 55 mm to 75 mm
9 | Independent diopter adjustments are on each eyepiece
10 | rubber eyeshields
11 | Trinocular port includes a 20/80 beam splitter
12 | Viewing Heads:
13 | Optical bridge with side-by-side two head dual observation
14 | Total Magnification
15 | 40x, 100x, 400x, 1000x
16 | Specimen Stage:
17 | Rectangular mechanical stage
18 | 174mm x 145mm with a travel range of 76mm x 50mm
19 | Low positioned coaxial control knobs
20 | The stage is protected with a special hard coating to protect against abrasion and fungus growth
21 | Stand
22 | Effective ergonomics and maximum expandability
23 | Focusing
24 | Fine and coarse coaxial focusing
25 | total vertical movement of 25mm at 2 µm
26 | Condenser
27 | The substage condenser mount is height adjustable
28 | Homogeneous illumination for observation and photomicrography at all magnifications from 2x to 100x
29 | Illumination:
30 | LED light
31 | Unit shall have a camera adapter
32 | Shall have Quadruple nosepiece configuration
33 | Shall specify the exact dimension and weight for the offered equipment
34 | Compliance with standards & legislation:
35 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
36 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
37 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
38 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
39 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | MLB10039 | 0 | CENTRIFUGE FLOOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Shall be a floor model with temperature range from -10 to +40 ºC and CFC free refrigeration
2 | Shall be configured to support the maximum speed of 15,000 rpm & 25000 RCF
3 | Shall have minimum of 5 direct recall program keys and capability of up to 99 programs
4 | Microprocessor control system with the features of direct, brushless induction low profile motor drive system
5 | Shall have Smart spin imbalance detection system with Auto-lock rotor locking system &pre-cooling function
6 | Shall have a dual time mode: time at start and time at speed to ensure reproducible results
7 | Shall have a run time set from 1 second to 99 minute 59 second
8 | Shall be a provision of run without brakes to allow friction less deceleration
9 | Shall have capability of password protection for the programs and lid opening
10 | Shall be able to display both set and actual parameter. Also, able to display both air/sample temperature as well as temperature in the sample, ensuring biological activity is maintained
11 | The rotors shall be either metallic or made of carbon fibre material and should be autoclavable at 121o C for sterilization whenever needed
12 | The installation and removal of rotors shall be user friendly without the requirement of any tool
13 | Shall be an option for automatic lid opening at the end of run
14 | Shall should have manual lid opening function in case of power failure
15 | Rotor Specifications:
16 | Rotor 1:
17 | Rotor Maximum Capacity: 4 liters (1000ml x 4)
18 | Rotor Maximum Speed: 9000 rpm
19 | Rotor Maximum Force: 16,000 or more
20 | Tubes/Bottles: 1000 ml PP bottles with caps, adaptors to be included wherever necessary (40 bottles)
21 | Rotor Material: Aluminum (Fixed angle)
22 | Rotor 2
23 | Rotor Maximum Capacity: 3 liters (6 x 500ml)
24 | Rotor Maximum Speed: 10,000 rpm
25 | Rotor Maximum Force: 17000 or more
26 | Tubes/Bottles: 500 ml tubes and 250 ml tubes to be provided. Adaptors to be included wherever necessary (80 bottles each)
27 | Rotor 3:
28 | Rotor Maximum Capacity: 1 liter (4 x 250 ml)
29 | Rotor Maximum Speed: 7,500 rpm
30 | Rotor Maximum Force: 10000 or more
31 | Tubes/Bottles: Necessary 250 ml bottles/adaptors to be provided
32 | Rotor Material: Aluminum (Swinging bucket)
33 | Rotor 4:
34 | Rotor Maximum Capacity: 300 ml ( 6 x 50 ml)
35 | Rotor Maximum Speed: 13000 rpm
36 | Rotor Maximum Force: 26000 or more
37 | Tubes/Bottles: Necessary 50 ml tubes/adaptors to be provided (100 tubes)
38 | Rotor 5:
39 | Rotor Maximum Capacity: 3 liters ( 4 x 750 ml)
40 | Rotor Maximum Speed: 4300 rpm
41 | Rotor Maximum Force: 4200 or more
42 | Tubes/Bottles: Necessary 750 ml bottles/adaptors to be provided (80 bottles)
43 | Shall specify the exact dimension and weight for the offered equipment
44 | Compliance with standards & legislation:
45 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
46 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
47 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
48 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
49 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | MLB10044 | 0 | CENTRIFUGE REFRIGERATED TABLE TOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | To incorporate digital display for speed (set and actual) and time, as well as other operating indicators such as lid locked, door closed/open, status, temperature.
2 | Manually adjustable parameters as well as user programming capability. State number of user defined programs and programmable parameters such as speed, time, acceleration, rotor type, brake, etc..
3 | Variable centrifugation speed up to ~ 5000 rpm
4 | Variable acceleration with acceleration and brake ramping. Specify range.
5 | Variable Relative Centrifugation Force. Specify range.
6 | Variable timer up to 60 min or better as well as continuous operation
7 | Acceleration/deceleration specs should be provided.
8 | Variable temperature from -10°C to 30°C in 10 increments.
9 | State accuracy
10 | Hermetically sealed, vibration-free compressor
11 | CFC free
12 | Swing out rotor with four buckets (and corresponding inserts) that can be used as open or closed. Sealing lid shall be included for each bucket.
13 | Rotator adaptors included:
14 | 4 x (15 x 7 ml) tubes
15 | 4 x (8 x 15 ml) tubes
16 | 4 x (13 x 10 ml) tubes
17 | The unit shall possess high end technical features and characteristics, especially those relating to sample and user safety. These features shall include but not be limited to:
18 | Brushless induction drive
19 | SS bowl and lid interior
20 | Easily removable rotor for cleaning the bowl
21 | Internal safety chamber (steel or similar) between the bowl and outer case
22 | Insulated interior for noiseless and vibration-free operation. Specify noise level.
23 | Smooth, low noise motor (state motor suspension method)
24 | Rubber positioning feet to prevent the unit from slipping
25 | Lid opening prevention mechanism when the rotor is turning
26 | Emergency stop capability
27 | Rotor imbalance detection system
28 | Audiovisual alarm with identification of problem or error code for easy troubleshooting. List all alarms.
29 | Electromagnetic lid locking system with power fail manual override
30 | Each of the following additional features will be considered as an asset during evaluation of the offered unit. State availability as well as characteristics where applicable:
31 | Pulse button
32 | Automatic rotor identification and verification with auto program
33 | Over speed detection system
34 | Compliance with standards & legislation:
35 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
36 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
37 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
38 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
39 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | MLB10045 | 0 | CHAIR BLOOD DRAWING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Wide with padded arm support that rotates 360 degrees.
2 | Suitable for right or left hand use.
3 | Arm rest adjusts to any size/height patient
4 | Stain resistant upholstery
5 | The armrest should lock into place at the desired location.
6 | Capability to carry up to 200Kg
7 | Storage drawer
8 | Non slip rubber feet
9 | Color to be coordinated with interior design / Engineer.
10 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | MLB40003 | 0 | ANALYZER AUTOMATED ELISA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled, fully automated, temperature controlled micro-plate reader and washer for routine and blood bank enzyme linked immuno-absorbance assays (ELISA)
2 | Digital LCD alphanumeric display of test parameters and status as well as other relevant information
3 | Printer for result documentation and Report customization shall be included.
4 | Accommodate all standard 96-well microplates (U, V, round and flat bottom; others shall be stated)
5 | Auto-positioning and centering
6 | State accuracy
7 | Multi-channel optical system for reading minimum of 8 channels (and 1 reference) simultaneously, with auto self-calibration capability
8 | Photometric accuracy +/- 1 %; +/- 0.01 Abs or better
9 | Reproducibility and linearity specifications should be specified
10 | Pre-defined stored protocols with user modification and saving capability for mostly performed test procedures.
11 | Specify number and description of stored protocols
12 | Specify overall on-board storage capacity for tests and results (in MB)
13 | Measurement modes shall include but not be limited to:
14 | Mono and bi-chromatic end point measurement
15 | Kinetic measurement (single and dual wavelength). Specify number of reading cycles and minimum interval time between cycles
16 | Wavelength range from approx. 400 to 700 nm
17 | Measurement range: 0 to 3.5 OD or better; (0.001 resolution)
18 | Specify light source type and life
19 | Specify time from power on to ready state (including warm up and any necessary calibration)
20 | Specify photodetector type and characteristics:
21 | Accuracy
22 | Reproducibility
23 | Noise level
24 | Linearity
25 | Specify filter type and characteristics.
26 | Specify reading speed for all modes (seconds/plate)
27 | Integrated incubator / shaker :Specify temperature control type, range and accuracy.
28 | Specify all software capabilities in terms of data acquisition, calculation and result presentation:
29 | Qualitative evaluations
30 | Quantitative evaluations
31 | Kinetic data reduction
32 | Mathematical formula transformation calculations
33 | Computer interfacing / connectivity for operational control, results and report formatting and generation, etc. and HIS including archiving capabilities .
34 | The option shall include all software and hardware requirements (PC, monitor, printer, other accessories and peripherals), communication hardware and software, etc. Each shall be specified in detail
35 | Barcode reader for plate identification shall be offered if available. Specify characteristics
36 | The offer shall include a fully automated microplate washer specifically designed for use with the microplate reader stated above. The washer shall incorporate the following specifications (if integrated, state so clearly):
37 | Microprocessor controlled with alphanumeric display
38 | To accommodate all standard microplate types (96 well, flat, U, V and round. Others shall be specified)
39 | Pre-programmed protocols for all standard applications, as well as user programming capability. State details including default and total number of programs
40 | Programmable parameters shall include the following. List corresponding specifications for each parameter as applicable (volume, height, time, range, limit, accuracy, resolution, precision, etc.)
41 | Dispense
42 | Aspirate (specify minimum residual volume)
43 | Wash (full plate, strip or multi-strips and number of wash heads)
44 | Shake
45 | Soak
46 | Flow
47 | Number of cycles
48 | Interval between steps
49 | The unit shall incorporate wash, rinse and waste fluid reservoirs
50 | Specify each reservoir capacity (minimum ~ 2 L)
51 | Method and ease of filling/emptying
52 | Incorporated safeties (overfilling protection, liquid level detection, alert messages, etc.)
53 | The offer shall include as option, a fully automated microplate reagent dispenser. The reagent dispenser shall incorporate the following specifications (if integrated, state so clearly):
54 | To accommodate all standard 96-well microplates (U, V, round and flat bottom; others shall be stated)
55 | Simultaneous 8-channel dispensing
56 | Sterilizable dispensing cassette
57 | Auto-positioning and centering. State accuracy
58 | Variable dispensing volume from ~ 1 μl to 100 μl in 1 μl increments or better
59 | Specify accuracy
60 | Specify precision
61 | Specify dispensing speed
62 | Capability to dispense multiple reagents in any number of columns.
63 | Specify number of liquids
64 | Specify method of delivery and line flushing / cleaning
65 | Shall specify the exact dimension and weight for the offered equipment
66 | Compliance with standards & legislation:
67 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
68 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
69 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
70 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
71 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | MLB40023 | 0 | INCUBATOR CO2 WITH O2 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled, direct heat air jacketed CO2 incubator with two consecutive doors, stainless steel body, compact design and clear back-lit display facilitating easy visualization of adjusted parameters
2 | Up 170 liter capacity
3 | Temperature control from 4°C above ambient to 50°C, with control accuracy ±0.1 ° C
4 | Direct heating feature (air jacketed) for better uniformity and rapid recovery
5 | Shall have CO2 control range from 0.2 to 20 % with control accuracy and uniformity of ± 0.1 % High quality CO2 sensors (infrared sensors).
6 | Shall have active control of O2 range from 0.1-19% or more
7 | Shall have with minimum 4 or 5 adjustable height perforated shelves & humidity reservoir (removable) to achieve at least 95 % RH
8 | Data storage facility (at least 72 Hr.) to record parameters such as CO2/O2 concentration, Temperature changes, door opening
9 | Shall have HEPA filter for incoming gases
10 | Shall have high heat decontamination in system temp up to 120°C or more during sterilization the system Shall operate without removing HEPA filter and CO2 sensor
11 | Compliance with standards & legislation:
12 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
13 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
14 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | MPH50004 | 0 | SHELVING PHARMACY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To store pharmaceutical products
2 | TYPE | Modular design
3 | NO. OF STORAGE LEVELS (SHELVES) | 8 approx.
4 | DEPTH OF SHELF | Modular design
5 | LENGTH OF SHELF | 120 cm approx.
6 | LAMINATED END PANELS | Yes
7 | BASE | Fixed on floor
8 | OTHER SPECIFICATION | Vendor is responsible for complete installation at site
9 | SHELVING MATERIAL | eboxy cooted metal
10 | DIVIDER KET | YES
11 | LABEL HOLDER | YES
12 | LEVEL ADJUSTABLE | YES
Company Stamp |  |  |  | Signature