## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | VENTILATOR TRANSPORT ADULT & PEDIATRIC
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | CART CRASH
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | CART DRESSING
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | CYLINDER O2 SIZE E WITH REGULATOR
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | REGULATOR SUCTION HIGH WALL MOUNT
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | STOOL FOOT 2 STEP
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | DEFIBRILLATOR UNIT
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | ECG 12 LEAD WITH CART
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | GLUCOMETER
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | LIGHT EXAM MOBILE
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | OXIMETER PULSE HAND HELD
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | PUMP SYRINGE
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | SPHYGMOMANOMETER ANEROID MOBILE
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | STETHOSCOPE ADULT
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | STETHOSCOPE PEDIATRIC
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | THERMOMETER DIGITAL HANDHELD
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | SUCTION UNIT PORTABLE
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | OTO OPTHALMOSCOPE DESKTOP
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | SCALE CHAIR
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | SCALE PATIENT WITH HEIGHT
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | BAG AMBO ADULT & PEDIATRIC
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | WHEELCHAIR STANDARD SIZE 24
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | CHAIR COMMODE ADULT
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | INTUBATION DIFFICULT VIDEO ADULT
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | ANALYZER MARKER CARDIAC
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | MONITOR CAPNOGRAPHY
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | THROMBOSIS PREVENTION VIEN DEEP
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | ULTRASOUND POC GENERAL PORTABLE
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | MONITOR BEDSIDE CENTRAL 10 BEDS ADULT & PEDIATRIC
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | CENTRIFUGE 48 TUBES TABLE TOP
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | MIXER BLOOD ROLLER TYPE
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | PIPETTOR AUTOMATIC
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | ANALYZER CHEMISTRY POC
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | X-RAY DIGITAL MOBILE
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | PROTECTION X-RAY APRON
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | PROTECTION X-RAY APRON VEST & SQUIRT
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | OPTHALMOSCOPE WALL MOUNT
ITEM 42 | This sheet contains the details of the bid with each technical specification for | ITEM 42 | HEADLIGHT FIBEROPTIC SURGICAL
ITEM 43 | This sheet contains the details of the bid with each technical specification for | ITEM 43 | SUCTION UNIT HIGH VACUUM
ITEM 44 | This sheet contains the details of the bid with each technical specification for | ITEM 44 | DETECTOR BLOOD FLOW
ITEM 45 | This sheet contains the details of the bid with each technical specification for | ITEM 45 | BLANKET WARMING SYSTEM AIR
ITEM 46 | This sheet contains the details of the bid with each technical specification for | ITEM 46 | INTUBATION SET COMPLETE CASE
ITEM 47 | This sheet contains the details of the bid with each technical specification for | ITEM 47 | LARYNGOSCOPE SET COMPLETE
ITEM 48 | This sheet contains the details of the bid with each technical specification for | ITEM 48 | WARMER BLOOD
ITEM 49 | This sheet contains the details of the bid with each technical specification for | ITEM 49 | REFRIGERATOR MEDICATION 700L
ITEM 50 | This sheet contains the details of the bid with each technical specification for | ITEM 50 | CABINET MEDICATION
ITEM 51 | This sheet contains the details of the bid with each technical specification for | ITEM 51 | CABINET NARCOTIC
ITEM 52 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | STRETCHER EMERGENCY
ITEM 53 | This sheet contains the details of the bid with each technical specification for | ITEM 49 | COUNTER BLOOD CELL 18 PARAMETERS
ITEM 54 | This sheet contains the details of the bid with each technical specification for | ITEM 50 | PROTECTION X-RAY THYROID COLLAR
ITEM 55 | This sheet contains the details of the bid with each technical specification for | ITEM 51 | TABLE EXAM GENERAL
ITEM 56 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | WARMER BLOOD RAPID INFUSION
ITEM 57 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | REFRIGERATOR MEDICATION 350L

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4217000002900 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | Medical Eqpt | 20 | [10 empty] | 0
2 | 4218190000900 | CART CRASH |  | Medical Eqpt | 35 | [10 empty] | 0
3 | 4214000000800 | CART DRESSING |  | Medical Eqpt | 20 | [10 empty] | 0
4 | 4214000002300 | CYLINDER O2 SIZE E WITH REGULATOR |  | Medical Eqpt | 100 | [10 empty] | 0
5 | 4214000003200 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | Medical Eqpt | 50 | [10 empty] | 0
6 | 4214000002800 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | Medical Eqpt | 550 | [10 empty] | 0
7 | 4214000005600 | REGULATOR SUCTION HIGH WALL MOUNT |  | Medical Eqpt | 50 | [10 empty] | 0
8 | 4227170104100 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER |  | Medical Eqpt | 400 | [10 empty] | 0
9 | 4214000007100 | STOOL FOOT 2 STEP |  | Medical Eqpt | 20 | [10 empty] | 0
10 | 4218190011900 | DEFIBRILLATOR UNIT |  | Medical Eqpt | 35 | [10 empty] | 0
11 | 4218190001800 | ECG 12 LEAD WITH CART |  | Medical Eqpt | 20 | [10 empty] | 0
12 | 4217000001200 | GLUCOMETER |  | Medical Eqpt | 20 | [10 empty] | 0
13 | 4214000004200 | LIGHT EXAM MOBILE |  | Medical Eqpt | 35 | [10 empty] | 0
14 | 4214000004800 | OXIMETER PULSE HAND HELD |  | Medical Eqpt | 50 | [10 empty] | 0
15 | 4218190007000 | PUMP SYRINGE |  | Medical Eqpt | 250 | [10 empty] | 0
16 | 4214000006300 | SPHYGMOMANOMETER ANEROID MOBILE |  | Medical Eqpt | 10 | [10 empty] | 0
17 | 4214000006700 | STETHOSCOPE ADULT |  | Medical Eqpt | 35 | [10 empty] | 0
18 | 4214000006800 | STETHOSCOPE PEDIATRIC |  | Medical Eqpt | 10 | [10 empty] | 0
19 | 4214000007800 | THERMOMETER DIGITAL HANDHELD |  | Medical Eqpt | 100 | [10 empty] | 0
20 | 4229352200100 | SUCTION UNIT PORTABLE |  | Medical Eqpt | 150 | [10 empty] | 0
21 | 4218200501300 | OTO OPTHALMOSCOPE DESKTOP |  | Medical Eqpt | 10 | [10 empty] | 0
22 | 4214000006100 | SCALE CHAIR |  | Medical Eqpt | 10 | [10 empty] | 0
23 | 4214000006000 | SCALE PATIENT WITH HEIGHT |  | Medical Eqpt | 10 | [10 empty] | 0
24 | 4214000012200 | BAG AMBO ADULT & PEDIATRIC |  | Medical Eqpt | 120 | [10 empty] | 0
25 | 4219260200800 | WHEELCHAIR STANDARD SIZE 24 |  | Medical Eqpt | 150 | [10 empty] | 0
26 | 4221161000400 | CHAIR COMMODE ADULT |  | Medical Eqpt | 20 | [10 empty] | 0
27 | 4218190012300 | INTUBATION DIFFICULT VIDEO ADULT |  | Medical Eqpt | 12 | [10 empty] | 0
28 | 4218190003400 | ANALYZER MARKER CARDIAC |  | Medical Eqpt | 10 | [10 empty] | 0
29 | 4218190004300 | MONITOR CAPNOGRAPHY |  | Medical Eqpt | 35 | [10 empty] | 0
30 | 4218190001500 | THROMBOSIS PREVENTION VIEN DEEP |  | Medical Eqpt | 50 | [10 empty] | 0
31 | 4220171200700 | ULTRASOUND POC GENERAL PORTABLE |  | Medical Eqpt | 5 | [10 empty] | 0
32 | 4218190300100 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE |  | Medical Eqpt | 60 | [10 empty] | 0
33 | 4218190014800 | MONITOR BEDSIDE CENTRAL 10 BEDS ADULT & PEDIATRIC |  | Medical Eqpt | 5 | [10 empty] | 0
34 | 4115000007300 | CENTRIFUGE 48 TUBES TABLE TOP |  | Medical Eqpt | 5 | [10 empty] | 0
35 | 4115000021900 | MIXER BLOOD ROLLER TYPE |  | Medical Eqpt | 5 | [10 empty] | 0
36 | 4115000024500 | PIPETTOR AUTOMATIC |  | Medical Eqpt | 5 | [10 empty] | 0
37 | 4115000002000 | ANALYZER CHEMISTRY POC |  | Medical Eqpt | 10 | [10 empty] | 0
38 | 4220180001300 | X-RAY DIGITAL MOBILE |  | Medical Eqpt | 5 | [10 empty] | 0
39 | 4220370000400 | PROTECTION X-RAY APRON |  | Medical Eqpt | 5 | [10 empty] | 0
40 | 4220370000500 | PROTECTION X-RAY APRON VEST & SQUIRT |  | Medical Eqpt | 5 | [10 empty] | 0
41 | 4218200500400 | OPTHALMOSCOPE WALL MOUNT |  | Medical Eqpt | 10 | [10 empty] | 0
42 | 4229510002600 | HEADLIGHT FIBEROPTIC SURGICAL |  | Medical Eqpt | 10 | [10 empty] | 0
43 | 4229510007000 | SUCTION UNIT HIGH VACUUM |  | Medical Eqpt | 50 | [10 empty] | 0
44 | 4218190001700 | DETECTOR BLOOD FLOW |  | Medical Eqpt | 5 | [10 empty] | 0
45 | 4229350000800 | BLANKET WARMING SYSTEM AIR |  | Medical Eqpt | 10 | [10 empty] | 0
46 | 4229350001600 | INTUBATION SET COMPLETE CASE |  | Medical Eqpt | 15 | [10 empty] | 0
47 | 4229350001800 | LARYNGOSCOPE SET COMPLETE |  | Medical Eqpt | 45 | [10 empty] | 0
48 | 4229350002700 | WARMER BLOOD |  | Medical Eqpt | 10 | [10 empty] | 0
49 | 4110301100700 | REFRIGERATOR MEDICATION 700L |  | Medical Eqpt | 10 | [10 empty] | 0
50 | 4214000010300 | CABINET MEDICATION |  | Medical Eqpt | 35 | [10 empty] | 0
51 | 4219190400000 | CABINET NARCOTIC |  | Medical Eqpt | 5 | [10 empty] | 0
52 | 4618181001100 | STRETCHER EMERGENCY |  | Medical Eqpt | 20 | [10 empty] | 0
53 | 4115000010000 | COUNTER BLOOD CELL 18 PARAMETERS |  | Medical Eqpt | 10 | [10 empty] | 0
54 | 4220370001200 | PROTECTION X-RAY THYROID COLLAR |  | Medical Eqpt | 5 | [10 empty] | 0
55 | 4214000014400 | TABLE EXAM GENERAL |  | Medical Eqpt | 50 | [10 empty] | 0
56 | 4229350002800 | WARMER BLOOD RAPID INFUSION |  | Medical Eqpt | 10 | [10 empty] | 0
57 | 4110301102600 | REFRIGERATOR MEDICATION 350L |  | Medical Eqpt | 5 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4217000002900 | 0 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To Transport Critical Care Adult, Pediatric and Neonatal Patients
2 | CONTROLS | Microprocessor
3 | - Tidal volume, mL | 219179
4 | - Inspiratory pressure, cmH20 | 0 - 100
5 | - Ventilation rate, Breaths / min | 0 - 150
6 | - Inspiratory flow, lpm | 1 - 180
7 | - Insp time. Sec | 0.1 - 10 with increments of 0.1
8 | - Exp time. Sec | 0.1 - 30
9 | - Insp Hold / plateau | Yes
10 | - Expiratory Hold | Yes
11 | - IE ratio | Specify
12 | - Fi02, % | 21 - 100 with Increments of 1
13 | - Manual Oxygen Breath | For 2 min
14 | - Manual breath | Yes
15 | - PEEP / CPAP. cmH20 | 0 - 35
16 | - Sensitivity range, cmH20 / L / min | 0 - 35
17 | - Trigger mechanism | flow / pressure
18 | - Capnography (adult, pediatric, neonate) | Yes
19 | - Leak Test | Yes
20 | - Leak Compensation | 100% invasive & 200% non-invasive
21 | - Manual Calibration | Yes
22 | - Apnea time, secs | 5-60 with of Increments of 5
23 | Neonatal Option:
24 | - Flow Sensor | Yes with 20 Pcs Each Machine
25 | - Flow Sensor Cable | Yes 2 Each Machine
26 | - Nasal CPAP | Yes
27 | Patient Transport Capability | Yes
28 | OPERATING MODES
29 | - Pressure - controlled | Yes
30 | - Assist / control ( Volume / Pressure) | Yes / Specify
31 | - IMV / SIMV ( Volume / Pressure) | Yes / Specify
32 | - Volume -controlled | Yes
33 | - Pressure support, cmH2O | 0 - 60
34 | - Pressure Regulated | Yes
35 | - Combination modes (Assist + Spontaneous) | Specify
36 | - CPAP | Yes
37 | - APRV /Bilevel mode | Yes
38 | - NIV | Yes
39 | - Apnea-backup vent | Yes
40 | DISPLAY
41 | Type | Specify
42 | Data Displayed | Digital Number , Graphics waveform, Loops and Trends
43 | INTERFACING
44 | - Output ports | Yes RS 232/USB
45 | - Remote alarm/ display | Yes
46 | - Report generation | Yes / SIMV
47 | - Display | Yes or Via RS232/USB
48 | - Hard Copy | Yes or Via RS232/USB
49 | - Archival disk | Yes or Via RS232/USB
50 | - Network | Yes or Via RS232/USB
51 | ALARMS
52 | Volume
53 | - Low & High VT | Yes
54 | - Low & High MV | Yes
55 | Pressure
56 | - High PIP | Yes
57 | - Low & High PEEP | Yes
58 | - Pressure Limited | Yes
59 | Respiratory Rate
60 | - High RR, bpm | 15 - 100
61 | - Apnea | Yes
62 | FIO2
63 | - Low FIO2 | 5% below the set
64 | - High FIO2 | 5 % above the set
65 | - Gas supply failure | Yes
66 | - Power failure | Yes
67 | - Vent inoperative | Yes
68 | - Low battery | Yes
69 | Self Diagnostics | Yes
70 | Other Alarms | PEEP / Patient trig. / I:E ratio / Inv.I:E ratio / Spontaneous
71 | Expiratory cassette Reusable (2 backup for each machine) | Yes
72 | PHYSICAL SPECIFICATIONS
73 | Dimensions | Specify
74 | Weight | Specify
75 | Disposables | Disposable Transport circuit with full accessories (Exhalation valves)
76 | Durable Trolley | Yes
77 | APPLICATION | To Transport Critical Care Adult, Pediatric and Neonatal Patients
78 | CONTROLS | Microprocessor
79 | - Tidal volume, mL | 219179
80 | - Inspiratory pressure, cmH20 | 0 - 100
81 | - Ventilation rate, Breaths / min | 0 - 150
82 | - Inspiratory flow, lpm | 1 - 180
83 | - Insp time. Sec | 0.1 - 10 with increments of 0.1
84 | - Exp time. Sec | 0.1 - 30
85 | - Insp Hold / plateau | Yes
86 | - Expiratory Hold | Yes
87 | - IE ratio | Specify
88 | - Fi02, % | 21 - 100 with Increments of 1
89 | - Manual Oxygen Breath | For 2 min
90 | - Manual breath | Yes
91 | - PEEP / CPAP. cmH20 | 0 - 35
92 | - Sensitivity range, cmH20 / L / min | 0 - 35
93 | - Trigger mechanism | flow / pressure
94 | - Capnography (adult, pediatric, neonate) | Yes
95 | - Leak Test | Yes
96 | - Leak Compensation | 100% invasive & 200% non-invasive
97 | - Manual Calibration | Yes
98 | - Apnea time, secs | 5-60 with of Increments of 5
99 | Neonatal Option:
100 | - Flow Sensor | Yes with 20 Pcs Each Machine
101 | - Flow Sensor Cable | Yes 2 Each Machine
102 | - Nasal CPAP | Yes
103 | Patient Transport Capability | Yes
104 | OPERATING MODES
105 | - Pressure - controlled | Yes
106 | - Assist / control ( Volume / Pressure) | Yes / Specify
107 | - IMV / SIMV ( Volume / Pressure) | Yes / Specify
108 | - Volume -controlled | Yes
109 | - Pressure support, cmH2O | 0 - 60
110 | - Pressure Regulated | Yes
111 | - Combination modes (Assist + Spontaneous) | Specify
112 | - CPAP | Yes
113 | - APRV /Bilevel mode | Yes
114 | - NIV | Yes
115 | - Apnea-backup vent | Yes
116 | DISPLAY
117 | Type | Specify
118 | Data Displayed | Digital Number , Graphics waveform, Loops and Trends
119 | INTERFACING
120 | - Output ports | Yes RS 232/USB
121 | - Remote alarm/ display | Yes
122 | - Report generation | Yes / SIMV
123 | - Display | Yes or Via RS232/USB
124 | - Hard Copy | Yes or Via RS232/USB
125 | - Archival disk | Yes or Via RS232/USB
126 | - Network | Yes or Via RS232/USB
127 | ALARMS
128 | Volume
129 | - Low & High VT | Yes
130 | - Low & High MV | Yes
131 | Pressure
132 | - High PIP | Yes
133 | - Low & High PEEP | Yes
134 | - Pressure Limited | Yes
135 | Respiratory Rate
136 | - High RR, bpm | 15 - 100
137 | - Apnea | Yes
138 | FIO2
139 | - Low FIO2 | 5% below the set
140 | - High FIO2 | 5 % above the set
141 | - Gas supply failure | Yes
142 | - Power failure | Yes
143 | - Vent inoperative | Yes
144 | - Low battery | Yes
145 | Self Diagnostics | Yes
146 | Other Alarms | PEEP / Patient trig. / I:E ratio / Inv.I:E ratio / Spontaneous
147 | Expiratory cassette Reusable (2 backup for each machine) | Yes
148 | PHYSICAL SPECIFICATIONS
149 | Dimensions | Specify
150 | Weight | Specify
151 | Disposables | Disposable Transport circuit with full accessories (Exhalation valves)
152 | Durable Trolley | Yes
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 4218190000900 | 0 | CART CRASH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Cardio-respiratory resuscitation system | (1) Top Cavity Tray, (3) Locking Side Bins, (1) Adjustable Defibrillator Tray, (1) Suction Pump Shelf, (1) Backboard with Front Assembly Kit, (1) Oxygen Tank Holder, (1) I.V. Pole with Cart Mount, (1) Storage/Gel Bin, (1) Pack Plastic Security Seal
2 | Accessories for Cart
3 | Oxygen supply system | Yes
4 | No.of cylinder | 2 with press. Reducing valve
5 | Oxygen flow meter | 0 - 15 L/min
6 | Suction Unit | Prefered Laerdal / Specify
7 | Built-in AC/DC charger | Yes
8 | Large knob controls vacuum regulation & unit power | Yes
9 | Flows, LPM | 30 LPM
10 | Vacuum power | 500+ mmHg
11 | Disposable canister and patient tubing | Yes
12 | Medical grade power extention | Built-in, at the back
13 | Laryngoscope (for pedia and adult) with set Blades ( Miller sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
14 | Laryngoscope (for pedia and adult) with set Blades ( Mickintosh sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
15 | Guedel airway set (size 0-4) | Yes
16 | Oxygen therapy mask set | Yes
17 | Resuscitation bag (for adult + pediatric) | Yes
18 | Mask | Yes
19 | Suction handle | Yes
20 | Tubing | Yes
21 | Tourniquet | Yes
22 | Fasteners (velcro) | Yes
23 | Forceps - Magil / Artery | Yes
24 | Scissors | Yes
25 | Dressing | Yes
26 | Suture | Yes
27 | Penlight | Yes
28 | Bite stick | Yes
29 | Nasal canula set | Yes
30 | Stop watch | Yes
31 | Aneroid Sphygmomanometer with all sizes cuffs (Adult & Pediatric) | Yes, aneroid
32 | Infusion pole | Yes
33 | ECG patient lead set | Yes
34 | ECG Electrodes | Yes
35 | ECG Chart recorder paper | 5 roll /each unit
36 | Additional , 1. NIBP hose 2.Pulse oxymetry cable 3.Pacing electrodes+cable | Yes
37 | Disposables | Yes
38 | Defibrillator tray | Yes
39 | Side tray for holding a portable suction machine | Yes
40 | Adult and pediatric Defibrillation pads. | Yes
41 | Arterial clamp | Yes
42 | Stethoscope | Yes
43 | Medication tray-3" (76 mm) Drawer tray approx.20" x 15" ( 508 x 381 mm) with 4 dividers, 2ea. Front to rear rails, sheet of blank labels, 5 security bags. | Yes
44 | PHYSICAL SPECIFICATIONS
45 | Dimensions | Specify
46 | Weight | Specify
47 | Cardio-respiratory resuscitation system | (1) Top Cavity Tray, (3) Locking Side Bins, (1) Adjustable Defibrillator Tray, (1) Suction Pump Shelf, (1) Backboard with Front Assembly Kit, (1) Oxygen Tank Holder, (1) I.V. Pole with Cart Mount, (1) Storage/Gel Bin, (1) Pack Plastic Security Seal
48 | Accessories for Cart
49 | Oxygen supply system | Yes
50 | No.of cylinder | 2 with press. Reducing valve
51 | Oxygen flow meter | 0 - 15 L/min
52 | Suction Unit | Prefered Laerdal / Specify
53 | Built-in AC/DC charger | Yes
54 | Large knob controls vacuum regulation & unit power | Yes
55 | Flows, LPM | 30 LPM
56 | Vacuum power | 500+ mmHg
57 | Disposable canister and patient tubing | Yes
58 | Medical grade power extention | Built-in, at the back
59 | Laryngoscope (for pedia and adult) with set Blades ( Miller sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
60 | Laryngoscope (for pedia and adult) with set Blades ( Mickintosh sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
61 | Guedel airway set (size 0-4) | Yes
62 | Oxygen therapy mask set | Yes
63 | Resuscitation bag (for adult + pediatric) | Yes
64 | Mask | Yes
65 | Suction handle | Yes
66 | Tubing | Yes
67 | Tourniquet | Yes
68 | Fasteners (velcro) | Yes
69 | Forceps - Magil / Artery | Yes
70 | Scissors | Yes
71 | Dressing | Yes
72 | Suture | Yes
73 | Penlight | Yes
74 | Bite stick | Yes
75 | Nasal canula set | Yes
76 | Stop watch | Yes
77 | Aneroid Sphygmomanometer with all sizes cuffs (Adult & Pediatric) | Yes, aneroid
78 | Infusion pole | Yes
79 | ECG patient lead set | Yes
80 | ECG Electrodes | Yes
81 | ECG Chart recorder paper | 5 roll /each unit
82 | Additional , 1. NIBP hose 2.Pulse oxymetry cable 3.Pacing electrodes+cable | Yes
83 | Disposables | Yes
84 | Defibrillator tray | Yes
85 | Side tray for holding a portable suction machine | Yes
86 | Adult and pediatric Defibrillation pads. | Yes
87 | Arterial clamp | Yes
88 | Stethoscope | Yes
89 | Medication tray-3" (76 mm) Drawer tray approx.20" x 15" ( 508 x 381 mm) with 4 dividers, 2ea. Front to rear rails, sheet of blank labels, 5 security bags. | Yes
90 | PHYSICAL SPECIFICATIONS
91 | Dimensions | Specify
92 | Weight | Specify
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 4214000000800 | 0 | CART DRESSING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety standard | FDA approval or CE Marking
2 | Design | Heavy duty, compact design & high quality finishing
3 | CONSTRUCTION | Rectangular tube frame, Made from S/S or Equiv. that is corrosion resistance & easy to clean. Rugged construction to with stand mechanical shocks caused by movement over uneven surfaces
4 | COMPARTMENT | Please Specify .
5 | WORKING TRAY | pull-out & made from S/S
6 | DRAWERS | 5 utility drawers with partitions to store dressing instruments, gauze etc?.
7 | CUPBOARD | One included
8 | ADJUSTABLE SHELF | Please Specify
9 | WASTE RECEPTACLE | Removable with locking mechanism
10 | TRAY | Swivel mounted & Made from S/S or Equiv
11 | BUCKET WITH LID | Available within fixture & made from S/S or equiv.
12 | AMPOULE OPENER | Included
13 | CASTORS | 4-antistatic castors with 2 efficient braking mechanism
14 | Table Top surface | Not less than 900 mm x 600 mm approx. Made from S/S or Equiv.
15 | Table top height | Not less than 1500 mm approx.
16 | BUMPERS | Revolving Bumpers
17 | Push Handles | Full-width hand rail on both sides
18 | Guard rail | 3-sided made from S/S or equiv.
19 | Accessories & Options | Any available options shall be quoted separately
20 | Catalogues | Original/colored detailed catalogue stating all the above mentioned specifications should be submitted with offer
21 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 4214000002300 | 0 | CYLINDER O2 SIZE E WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | OXYGEN CYLINDER E-SIZE | YES
2 | ALUMINUM CYLINDER | YES
3 | Nominal contents (litres) | 680L
4 | PIN INDEX | YES
5 | CAPACITY | SPECIFY
6 | ALUMINUM TROLLEY | INCLUDED
7 | OXYGEN REGULATOR | INCLUDED WITH FLOW METER 15 L/Min
8 | OTHER SPECIFICATION | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 4214000003200 | 0 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Flow Meter | Wall Mounted
2 | Flow Range | 0 - 15 L / Min
3 | Color Coded | specify
4 | Pressure Compensated | Yes
5 | Glass Envelop | Included with indicator ball
6 | Nipple | Yes
7 | Probe | As Per Site
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 4214000002800 | 0 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Pressure compensated
2 | Flow tube | Impact resistant
3 | Flow indication tube | Single
4 | Lower scale | Expanded
5 | Flow range | Up to 15 L/min, approx
6 | Humidifier cup volume | Reusable
7 | Flowmeter compatibility to all hospital wall outlets | Yes
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 4214000005600 | 0 | REGULATOR SUCTION HIGH WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MOUNTING | Wall / Pneumatic
2 | STANDARD FEATURES AND TECHNICAL SPECIFICATIONS
3 | Autoclaveable Bottles | Yes
4 | Suction Regulator Modes | Continuous & Intermittent
5 | Adjustable Timing Cycles | At least from 3 to 30 sec
6 | Fully Adjustable Negative Pressure | From 0 to -200 mmHg or more
7 | Flow Rate at Continuous Mode | At least 90 L/min
8 | Overflow shut off valve/trap | Yes
9 | British Standard Probe | Yes
10 | Compatible with KFMC Standards | Yes
11 | Suction Bottles and Canisters | Disposable / reusable
12 | PHYSICAL SPECIFICATIONS
13 | Dimensions | Specify
14 | Weight | Specify
15 | MOUNTING | Wall / Pneumatic
16 | STANDARD FEATURES AND TECHNICAL SPECIFICATIONS
17 | Autoclaveable Bottles | Yes
18 | Suction Regulator Modes | Continuous & Intermittent
19 | Adjustable Timing Cycles | At least from 3 to 30 sec
20 | Fully Adjustable Negative Pressure | From 0 to -200 mmHg or more
21 | Flow Rate at Continuous Mode | At least 90 L/min
22 | Overflow shut off valve/trap | Yes
23 | British Standard Probe | Yes
24 | Compatible with KFMC Standards | Yes
25 | Suction Bottles and Canisters | Disposable / reusable
26 | PHYSICAL SPECIFICATIONS
27 | Dimensions | Specify
28 | Weight | Specify
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 4227170104100 | 0 | CYLINDER OXYGEN SIZE K/H WITH FLOWMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CYLINDER SIZE H | FOR MEDICAL OXYGEN
2 | CONSTRUCTION | STEEL
3 | CAPACITY | 7,700 LITERS APPROX
4 | REGULATORS / VALVE | CGA 540
5 | FLOW METER 15L/MIN | INCLUDED WITH CYLINDER
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | 4214000007100 | 0 | STOOL FOOT 2 STEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FOOT STOOL - two step
2 | MATERIAL | Specify the material type
3 | DIMENSION | ALMOST ( 14x19x6 ) WxLxH
4 | TOP SURFACE | SLIP RESISTANCE
5 | WEIGHT - ALONE | Minimum 2.5 Kg
6 | WEIGHT - STACKED | Maximum 200 Kg
7 | OTHERS | EASY CLEAN - INFECTION CONTROL REGULATION
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | 4218190011900 | 0 | DEFIBRILLATOR UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Automatic/Semi Automatic
2 | DEFIBRILLATOR FEATURES
3 | Manual override | Yes
4 | Voice prompting | Yes
5 | Number of Wave forms | More than 8
6 | Unit with adult & pediatric paddles | Yes
7 | Energy sequence, AED mode, J | as per AHA guidelines
8 | Energy settings, manual mode, J | Step up 2-200 (or higher)
9 | Protocol configured | Specified
10 | Output waveshape | Truncated exponential or rectilinear
11 | Automated External Defibrillation Option | Yes
12 | Shock advisory. | Yes
13 | Synchronized cardioversion. | Yes
14 | Non-Invasive blood pressure monitoring | Yes
15 | Non-invasive pacing | Yes
16 | Side stream Capnography. | Yes
17 | Temperature monitor. | Yes
18 | 12-lead ECG monitoring. | Yes
19 | SpO2 monitoring. | Yes
20 | Using CPR metronome Built-in | Yes
21 | Audible tone guide without distracting vocal critique | Yes
22 | Sunview button | Yes
23 | Data display shall be arranged in such way that ease the reading and quick identification of abnormal reading. | Yes
24 | Trending and storage features for all parameters in both Defib/Monitormode. | Yes
25 | The unit shall comply with the most recent AHA guidelines. | Yes
26 | Discharge waveform shall provide the most effective energy delivery with the lowest possible damage; vendor shall submit all documents supporting their claim. | Yes
27 | Unit shall be provided with versatile arrhythmia detection algorithm.Vendor shall specify how many leads are being analyzed simultaneously. | Yes
28 | System shall be equipped with advisory messages and/or a feedback technique that provide the caregiver with information about the number and depth of compressions during CPR. | Yes
29 | The unit shall have full –energy Biphasic Manual Defibrillators. | Yes
30 | The equipment shall be easily switched from AED to manual mode. | Yes
31 | The unit shall be light weighted and have ergonomic design. | Yes
32 | The unit shall be easy to operate with clearly labeled big buttons for Analyze, Charge and Discharge. | Yes
33 | The unit shall be capable to recharge to the maximum energy in less than 5 seconds. | Yes
34 | The unit should incorporate hard adults paddles with integrated pediatrics paddles. | Yes
35 | Parameters Monitored | ECG, Pacer
36 | Pacemaker | Yes
37 | MONITOR
38 | The unit shall be provided with the following hardware features:
39 | Large ≥ 8” high-contrast TFT color display with resolution ≥ 800 x 600 | Yes
40 | Thermal array printer | Yes
41 | AC adaptor | Yes
42 | Patient cables and probes/sensors and all necessary accessories. | Yes
43 | Status indicator that should visually warn the users about the readiness of the system even if the battery is fully depleted and AC power is off. | Yes
44 | The unit shall be provided at least with the following alerting/alarming features
45 | The system shall provide an auditable signal for QRS | Yes
46 | Audiovisual alarms for any parameter(s) out of range | Yes
47 | Accurate Battery gauge and near battery depletion | Yes
48 | Indication of AD/DC power | Yes
49 | Arrhythmia detection | Yes
50 | System shall perform a daily automatic self test or on-demand and shall report (print) the readiness of the system, and log the result of testing in an internal memory. | Yes
51 | All hardware and software problems shall be logged with date and time stamp; the system shall be capable of recording at least 3000 events. Accessing and deleting events shall be authenticated. | Yes
52 | Unit shall recognize the type of the pads and limit the energy accordingly. (Pediatrics shall have a limited energy level that is different from that of adult. | Yes
53 | The unit should provide defibrillation energy selection, and demand/fixed rate pacing, and selection of ECG leads monitoring. | Yes
54 | The unit paddles should provide Charge/Discharge controls. | Yes
55 | The unit should provide AC/DC operation, with Li-Ion/Li-Pol batteries. | Yes
56 | The batteries should be fully charged 100% within < 3 hours. When fully charged, the unit should provide > 200 discharges or 6 hours of continuous monitoring with full parameters. | Yes
57 | The unit should be suitable for crash cart purposes. | Yes
58 | The printer should provide summary report and code markers. | Yes
59 | All external components should be securely mounted. | Yes
60 | The unit should be secure and provide adequate protection against moving and electrically energized parts. | Yes
61 | Switches, knobs, and other controls should be designed for conditions of heavy use. | Yes
62 | The controls (e.g., switches, knobs) should be visible and clearly identified, and their functions should be self-evident. Device design should prevent misinterpretation of displays and control settings. | Yes
63 | Controls should be sealed against penetration of liquids. | Yes
64 | The unit should be easy to clean, disinfect, and/or sterilize. | Yes
65 | Vendor shall provide itemized price for all option, disposable come with machine. | Yes
66 | Monitor ECG acquisition | Dispo. Defeb. Elect/ multifunctunal/quick patch
67 | ECG display Type / Message display Type | LCD/LCD
68 | Heart rate display Gain | Yes /autogain/manual
69 | Spo2 display with probe | Yes, Nellcor, finger, neonatal
70 | Disposable /Reusable | Yes
71 | Electrodes | Solid gel
72 | Conductive area | 82 cm2 approx.
73 | Shelf life | Not less than 1 8 months
74 | DOCUMENTATION /DATA STORAGE/ANALYSIS
75 | Chart recorder > 100mm | Yes
76 | ECG/voice recording | Standard/Optional
77 | Capacity, min | 60/90
78 | Playback | Yes
79 | Solid-state memory | Yes
80 | Information stored | Analysis/events
81 | Capacity, min | Specify
82 | Playback | Yes
83 | Event record database storage | Yes
84 | Data management software | Yes
85 | Analysis
86 | - Auto or manual | Auto/Manual/Both
87 | - Segment analyzed, sec | <3
88 | - Analysis time, sec | 43994
89 | - V tach rate threshold, bpm | >120
90 | - VF amplitude threshold, mV | >0.08
91 | BATTERY | QTY : 2 sets
92 | Integrated battery charger | Yes
93 | External battery charger | Yes
94 | Type | rechargable
95 | Power supply | Yes
96 | Integral/replaceble | Integral, with built in charger
97 | Charging method | Line cord/external/plugin
98 | Charge time, hr | < 3 Hrs
99 | PHYSICAL SPECIFICATIONS
100 | Dimensions | Specify
101 | Weight | Specify
102 | TYPE | Automatic/Semi Automatic
103 | DEFIBRILLATOR FEATURES
104 | Manual override | Yes
105 | Voice prompting | Yes
106 | Number of Wave forms | More than 8
107 | Unit with adult & pediatric paddles | Yes
108 | Energy sequence, AED mode, J | as per AHA guidelines
109 | Energy settings, manual mode, J | Step up 2-200 (or higher)
110 | Protocol configured | Specified
111 | Output waveshape | Truncated exponential or rectilinear
112 | Automated External Defibrillation Option | Yes
113 | Shock advisory. | Yes
114 | Synchronized cardioversion. | Yes
115 | Non-Invasive blood pressure monitoring | Yes
116 | Non-invasive pacing | Yes
117 | Side stream Capnography. | Yes
118 | Temperature monitor. | Yes
119 | 12-lead ECG monitoring. | Yes
120 | SpO2 monitoring. | Yes
121 | Using CPR metronome Built-in | Yes
122 | Audible tone guide without distracting vocal critique | Yes
123 | Sunview button | Yes
124 | Data display shall be arranged in such way that ease the reading and quick identification of abnormal reading. | Yes
125 | Trending and storage features for all parameters in both Defib/Monitormode. | Yes
126 | The unit shall comply with the most recent AHA guidelines. | Yes
127 | Discharge waveform shall provide the most effective energy delivery with the lowest possible damage; vendor shall submit all documents supporting their claim. | Yes
128 | Unit shall be provided with versatile arrhythmia detection algorithm.Vendor shall specify how many leads are being analyzed simultaneously. | Yes
129 | System shall be equipped with advisory messages and/or a feedback technique that provide the caregiver with information about the number and depth of compressions during CPR. | Yes
130 | The unit shall have full –energy Biphasic Manual Defibrillators. | Yes
131 | The equipment shall be easily switched from AED to manual mode. | Yes
132 | The unit shall be light weighted and have ergonomic design. | Yes
133 | The unit shall be easy to operate with clearly labeled big buttons for Analyze, Charge and Discharge. | Yes
134 | The unit shall be capable to recharge to the maximum energy in less than 5 seconds. | Yes
135 | The unit should incorporate hard adults paddles with integrated pediatrics paddles. | Yes
136 | Parameters Monitored | ECG, Pacer
137 | Pacemaker | Yes
138 | MONITOR
139 | The unit shall be provided with the following hardware features:
140 | Large ≥ 8” high-contrast TFT color display with resolution ≥ 800 x 600 | Yes
141 | Thermal array printer | Yes
142 | AC adaptor | Yes
143 | Patient cables and probes/sensors and all necessary accessories. | Yes
144 | Status indicator that should visually warn the users about the readiness of the system even if the battery is fully depleted and AC power is off. | Yes
145 | The unit shall be provided at least with the following alerting/alarming features
146 | The system shall provide an auditable signal for QRS | Yes
147 | Audiovisual alarms for any parameter(s) out of range | Yes
148 | Accurate Battery gauge and near battery depletion | Yes
149 | Indication of AD/DC power | Yes
150 | Arrhythmia detection | Yes
151 | System shall perform a daily automatic self test or on-demand and shall report (print) the readiness of the system, and log the result of testing in an internal memory. | Yes
152 | All hardware and software problems shall be logged with date and time stamp; the system shall be capable of recording at least 3000 events. Accessing and deleting events shall be authenticated. | Yes
153 | Unit shall recognize the type of the pads and limit the energy accordingly. (Pediatrics shall have a limited energy level that is different from that of adult. | Yes
154 | The unit should provide defibrillation energy selection, and demand/fixed rate pacing, and selection of ECG leads monitoring. | Yes
155 | The unit paddles should provide Charge/Discharge controls. | Yes
156 | The unit should provide AC/DC operation, with Li-Ion/Li-Pol batteries. | Yes
157 | The batteries should be fully charged 100% within < 3 hours. When fully charged, the unit should provide > 200 discharges or 6 hours of continuous monitoring with full parameters. | Yes
158 | The unit should be suitable for crash cart purposes. | Yes
159 | The printer should provide summary report and code markers. | Yes
160 | All external components should be securely mounted. | Yes
161 | The unit should be secure and provide adequate protection against moving and electrically energized parts. | Yes
162 | Switches, knobs, and other controls should be designed for conditions of heavy use. | Yes
163 | The controls (e.g., switches, knobs) should be visible and clearly identified, and their functions should be self-evident. Device design should prevent misinterpretation of displays and control settings. | Yes
164 | Controls should be sealed against penetration of liquids. | Yes
165 | The unit should be easy to clean, disinfect, and/or sterilize. | Yes
166 | Vendor shall provide itemized price for all option, disposable come with machine. | Yes
167 | Monitor ECG acquisition | Dispo. Defeb. Elect/ multifunctunal/quick patch
168 | ECG display Type / Message display Type | LCD/LCD
169 | Heart rate display Gain | Yes /autogain/manual
170 | Spo2 display with probe | Yes, Nellcor, finger, neonatal
171 | Disposable /Reusable | Yes
172 | Electrodes | Solid gel
173 | Conductive area | 82 cm2 approx.
174 | Shelf life | Not less than 1 8 months
175 | DOCUMENTATION /DATA STORAGE/ANALYSIS
176 | Chart recorder > 100mm | Yes
177 | ECG/voice recording | Standard/Optional
178 | Capacity, min | 60/90
179 | Playback | Yes
180 | Solid-state memory | Yes
181 | Information stored | Analysis/events
182 | Capacity, min | Specify
183 | Playback | Yes
184 | Event record database storage | Yes
185 | Data management software | Yes
186 | Analysis
187 | - Auto or manual | Auto/Manual/Both
188 | - Segment analyzed, sec | <3
189 | - Analysis time, sec | 43994
190 | - V tach rate threshold, bpm | >120
191 | - VF amplitude threshold, mV | >0.08
192 | BATTERY | QTY : 2 sets
193 | Integrated battery charger | Yes
194 | External battery charger | Yes
195 | Type | rechargable
196 | Power supply | Yes
197 | Integral/replaceble | Integral, with built in charger
198 | Charging method | Line cord/external/plugin
199 | Charge time, hr | < 3 Hrs
200 | PHYSICAL SPECIFICATIONS
201 | Dimensions | Specify
202 | Weight | Specify
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | 4218190001800 | 0 | ECG 12 LEAD WITH CART |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile
2 | Cart | Yes
3 | LEADS
4 | Lead switching | Automatic/Manual
5 | Sensitivity, mm/mV | 5, 10, 20., other specify.
6 | Calibration signal | Automatic
7 | FREQUENCY RANGE, HZ
8 | Diagnostic | 0.05-150
9 | Filtered | Notch, EMG, Baseline, Wander, filters
10 | INPUT IMPEDANCE,
11 | Mega ohms | one hundred
12 | CMRR @ 60 Hz, dB | 100-140 approx
13 | Leads-off indicator | Yes
14 | RECORDER
15 | Channels | twelve
16 | Recording | YES A4 PAPER
17 | Lead marker | Automatic
18 | Timing marker | Yes
19 | Event marker | Yes
20 | Chart speed, mm/sec | 5, 10, 25, 50.
21 | LCD MONITOR 10 OR MORE | Yes ADJSTABLE
22 | No. of traces | 3, 6, 12
23 | PREVIEW SCREEN | Yes
24 | No. Waveforms stored | Min. 30
25 | ECG transmission | Yes
26 | Interpretation | Yes
27 | ECG measurements | Yes
28 | Auxiliary output | Yes
29 | Auxiliary input | Yes
30 | DEFIBRILLATOR SYRORNAIZ | YES
31 | SINGLE AVREGING | YSE
32 | Overload protection | Yes
33 | BATTERY OPERATION | Yes
34 | Battery type | Built in rechargeable
35 | Operating time, hr | Specify
36 | ACCESSORIES | All standard accessories
37 | PRINTER | Built-in
38 | A4 paper | Yes
39 | SAFTY AND PERFORMANCE | IEC 60601-2-51
40 | POWER REQUIREMENTS | 220V, 60Hz
41 | OTHER SPECIFICATIONS | Spare patient cable, chest electrodes, clamp electrodes for limb leads, with ICU Analysis. SOFT WARE Patient cable holder.FDA, CE, ISO, Approved. HL7
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | 4217000001200 | 0 | GLUCOMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor-controlled, hand held glucose meter to measure blood glucose levels.
2 | Should be IVD marked (In Vitro Diagnostic Medical Device).
3 | Measuring range: 20-500mg/dl Approx.
4 | Must give accurate results within Approx. 15 seconds.
5 | Must store at least the 20 results.
6 | Must be supplied with 200 strips.
7 | Must be supplied with a virtually pain-free blood sampling tool (lancing device).
8 | Specify the shelf life of strips (vial closed, after opening the vial)
9 | Must be supplied with long-life batteries.
10 | Unit must have an automatic shut off to protect battery life
11 | Unit must have low battery warning
12 | Must have a case for storage.
13 | Compliance with standards & legislation:
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | 4214000004200 | 0 | LIGHT EXAM MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Laryngoscope Set : | Neonatal & Infant
2 | BLADES
3 | Pediatric (Different Sizes) | 2 Different Sizes (0,0 / 0,1 / 1,1 / 1,2 )
4 | Pediatric (Different Sizes) | 2 Different Sizes 0,1,2,3
5 | ILLUMINATION | With Incorporated Fiber Optic Light Carrier inside the Blade (Xenon Light)
6 | Spare lamps | 6
7 | Rechargeable Battery Shell | With High Power LED Technology, more than 50,000 lux, Lithium-ion batteries
8 | Reduction Sleeve | For rechargeable batteries
9 | Handle Sleeve | Yes
10 | Charging Unit | For 2 Rechargeable batterie (110-240 VAC, 60 Hz)
11 | Bag for All Laryngoscopes | Splash - Protected
12 | Other Specifications | FDA, ISO Approved.
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | 4214000004800 | 0 | OXIMETER PULSE HAND HELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mobile examination light on sturdy and stable base
2 | On/Off switch, easily accessible on the light stand or light head
3 | Stability and precise positioning in every desired position.
4 | Weight balancing system in the rotating joint
5 | Smooth surfaces for perfect hygiene
6 | Light intensity: approx. 40,000 lux at 80 cm
7 | Quick lamp replacement
8 | Color temperature ≥ 3500K
9 | Temperature increase in the head area ≤ 2° C
10 | Angle of rotation: unlimited
11 | Minimal heat dissipation
12 | 3 white LED lights
13 | Sterilizable hand grip
14 | Main and battery-operate
15 | Shadowless light
16 | Swiveling castors with locking feature
17 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | 4218190007000 | 0 | PUMP SYRINGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Portable / Stand-alone/transport/HAND HELD FOR ADULT/PEDIATRIC/INFANT
2 | DISPLAYS | SpO2, pulse rate, pulse strength and/or signal, low battery, Volumes, alarm limits.
3 | Type | LCD/ with backlit LED SHOWING PARAMETERS & READINGS
4 | Sp02 RANGE, % | 0-100
5 | Accuracy, ,% | < ± 2 %
6 | PULSE RATE, bpm | 30-240
7 | PERFUSION INDEX AND/OR SIGNAL STRENGTH INDICATOR | YES
8 | RESPONSE TIME, sec | Pulse to Pulse or time ( Specify)
9 | SETTLING TIME | Pulses or time ( Specify)
10 | START-UP TIME, sec or pulses
11 | ALARMS:
12 | Audiovisual | Yes
13 | Visual | Yes, Low battery
14 | ALARM OVERRIDE | Yes
15 | Reactivation method | Yes, after 2 mins.
16 | Volume control | Yes
17 | SELF-TEST MODE | Yes
18 | Memory: up to 72 hours of data storage | YES
19 | PROBE TYPES | YES REUSABLE (1 ADULT, 1 PEDIATRIC, AND 1 NEONATE)
20 | Patient range | Adult / Neonate/ Pediatric
21 | Cable length, m | Not less than1.5 m
22 | BATTERY BACKUP :
23 | Rechargeable battery | Yes, Specify the type
24 | Low battery notice | Yes
25 | Rechargeable time, hr | Specify
26 | Battery life, hrs | Specify
27 | The unit should operate on batteries as well as AC electricity, with automatic battery charging while in use. | YES
28 | The unit shall operate on batteries | YES
29 | battery type | Specify
30 | VAC | 220 V
31 | CURRENT | 13 A
32 | FREQUENCY | 60 Hz.
33 | PLUG TYPE | 3 Pin British
34 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
35 | IMPORTANT NOTE: | Please specify whether accessories are standard or optional, otherwise they will be considered standard.
36 | MDMA | Yes
Company Stamp |  |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | 4214000006300 | 0 | SPHYGMOMANOMETER ANEROID MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled with digital LCD alphanumeric display of parameters and alarms
2 | Variable rate ranging from 0.1 to 500 mL/hr or better, with 0.1 mL/hr increments
3 | 3 % accuracy or better
4 | Variable volume-to-be-infused from 1 to 1,000 mL or similar
5 | Digitally displayed parameters to include:
6 | Infusion rate
7 | Battery / AC operation
8 | Running indicator
9 | Alarming condition when active, with indication of alarm type or code
10 | Back pressure monitor / indicator
11 | Event history
12 | Capability to accept different syringe types and sizes with automatic syringe detection and identification
13 | Syringe compatibility and auto detection shall include but not be limited to all sizes of the following (3 to 60 mL):
14 | BD
15 | Terumo
16 | Monoject
17 | Braun
18 | Fresenius
19 | Variable bolus rate
20 | Up to maximum flow rate
21 | Bolus infused volume indicator during bolus activation
22 | Protected access
23 | Audiovisual alarms shall include but not be limited to the following:
24 | Syringe installation and integrity (detection)
25 | Line disconnection (sudden drop in back pressure)
26 | Occlusion pressure pre-alarm
27 | Occlusion pressure
28 | Near end of infusion alarm
29 | End of infusion
30 | Volume limit pre-alarm
31 | Volume limit
32 | KVO (1 ml/hr; if other, specify)
33 | Low battery pre-alarm
34 | Discharged battery
35 | Internal malfunction
36 | Syringe unlocked
37 | Plunger disengaged
38 | Empty syringe
39 | Data log capability and data port for data transmission, display and printing. Any required software for such function shall be included.
40 | Logged data to include ( ≥1 year) :
41 | Settings
42 | Alarms
43 | Errors
44 | Key presses
45 | Error codes
46 | Amount infused
47 | Safety features shall include but not be limited to:
48 | Self test at start-up
49 | Nurse call interfacing capability
50 | Splash proof design
51 | Auto priming
52 | Adjustable alarm volume. No permanent silencing shall be possible.
53 | Keypad lock
54 | Impossibility to improperly install infusion set
55 | Free flow prevention system
56 | Last parameter setting retention
57 | IV stand mounting accessory shall be included
58 | Battery autonomy of 3 hrs or more when fully charged. specify:
59 | Battery type and characteristics (voltage and current capacity)
60 | Autonomy at 10 mL/hr
61 | Recharging time from depleted to 90%
62 | Compliance with standards & legislation:
63 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
64 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
65 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
66 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
67 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | 4214000006700 | 0 | STETHOSCOPE ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | METHOD | ANEROID
2 | INFLATION PRESSURE, mm Hg | 300 mm Hg
3 | CUFF SIZES | Adult, Pediatric, Velcro cuff
4 | SYSTOLIC, mm Hg | 0-240 approx.
5 | DIASTOLIC, mm Hg | 0-130 approx.
6 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | 4214000006800 | 0 | STETHOSCOPE PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Double head
2 | MATERIAL
3 | EAR TUBE | Stainless steel
4 | CHEST PIECE | Stainless steel
5 | BINAURAL
6 | ADJUSTMENT | To angle of ear
7 | CHEST PIECE | Adult and pediateric & Neonatal
8 | ADDITIONAL DIAPHARGM | 1 EACH
9 | EAR TIPS, TYPE | COMFORT SEALING
10 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | 4214000007800 | 0 | THERMOMETER DIGITAL HANDHELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Double head
2 | MATERIAL
3 | EAR TUBE | Stainless steel
4 | CHEST PIECE | Stainless steel
5 | BINAURAL
6 | ADJUSTMENT | To angle of ear
7 | CHEST PIECE | pediateric & Neonatal
8 | ADDITIONAL DIAPHARGM | 1 EACH
9 | EAR TIPS, TYPE | COMFORT SEALING
10 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | 4229352200100 | 0 | SUCTION UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Digital Thermometer safe for all ages and shall be for oral, underarm, or rectal use.
2 | Should use disposable probes for measurement
3 | Fast, accurate readings
4 | Temperature accuracy: 0.1°C
5 | Automatic shut-off feature
6 | Large LCD display for reading
7 | Should signal after temperature taken
8 | Safe to use, with no glass and mercury-free
9 | Should have an auto memory to show last 10 temperatures taken
10 | Should include rechargeable batteries with charging desk unit
11 | Unit should operate >1,000 uses on full battery
12 | Unit should have a self test capability
13 | Unit shall have visual indicators
14 | Final temperature stops flashing
15 | low-battery
16 | low and high temperatures
17 | Should have a suitable case
18 | Compliance with standards & legislation:
19 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
20 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | 4218200501300 | 0 | OTO OPTHALMOSCOPE DESKTOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Portable, Capable of connecting to bed
2 | TRACHEAL USE
3 | Pre-Hospital (EMS) | Yes
4 | Intra-Hospital (crash carts) | Yes
5 | PUMP TYPE | Piston
6 | COLLECTION CANISTER | Disposable
7 | Limit line, mL | 800-1000 mL
8 | Capacity measured, mL | 800-1000 mL
9 | Graduations | Yes
10 | Overflow protection | Yes
11 | Connector labels | Yes
12 | Filter Type | Bacterial, HEPA
13 | TUBING
14 | Internal diameter, mm (in) | 8 (0.31)
15 | Length, m (ft) | 1.8 (6)
16 | VACUUM
17 | Range, mm Hg | 50 to 600 mm Hg
18 | Time to 300 mm Hg, sec | <3
19 | Flow rate at max vacuum, L/min | 27
20 | Vacuum Level Control | Regulator, variable
21 | Vacuum Gauge Type | LED indicators
22 | Vacuum Gauge Scale, mmHg | 50 to 600 mm Hg
23 | PHYSICAL SPECIFICATIONS
24 | Dimensions | Specify
25 | Weight | Specify
26 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
27 | IMPORTANT NOTE | Please specify whether accessories are standard or optional, otherwise they will be considered standard
28 | External Charger | Yes/ Specify
29 | Wire Canister bracket | Yes/ Specify
30 | Rechargeable 12 Volts DC battery | Yes/ Specify
31 | AC and DC Power cords | Yes/ Specify
32 | AC/DC Adapter Charger | Yes/ Specify
33 | Disposable Canisters | Yes/ Specify
34 | Soft Pouch carrying case | Specify
35 | TYPE | Portable, Capable of connecting to bed
36 | TRACHEAL USE
37 | Pre-Hospital (EMS) | Yes
38 | Intra-Hospital (crash carts) | Yes
39 | PUMP TYPE | Piston
40 | COLLECTION CANISTER | Disposable
41 | Limit line, mL | 800-1000 mL
42 | Capacity measured, mL | 800-1000 mL
43 | Graduations | Yes
44 | Overflow protection | Yes
45 | Connector labels | Yes
46 | Filter Type | Bacterial, HEPA
47 | TUBING
48 | Internal diameter, mm (in) | 8 (0.31)
49 | Length, m (ft) | 1.8 (6)
50 | VACUUM
51 | Range, mm Hg | 50 to 600 mm Hg
52 | Time to 300 mm Hg, sec | <3
53 | Flow rate at max vacuum, L/min | 27
54 | Vacuum Level Control | Regulator, variable
55 | Vacuum Gauge Type | LED indicators
56 | Vacuum Gauge Scale, mmHg | 50 to 600 mm Hg
57 | PHYSICAL SPECIFICATIONS
58 | Dimensions | Specify
59 | Weight | Specify
60 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
61 | IMPORTANT NOTE | Please specify whether accessories are standard or optional, otherwise they will be considered standard
62 | External Charger | Yes/ Specify
63 | Wire Canister bracket | Yes/ Specify
64 | Rechargeable 12 Volts DC battery | Yes/ Specify
65 | AC and DC Power cords | Yes/ Specify
66 | AC/DC Adapter Charger | Yes/ Specify
67 | Disposable Canisters | Yes/ Specify
68 | Soft Pouch carrying case | Specify
Company Stamp |  |  |  | Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | 4214000006100 | 0 | SCALE CHAIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Tabletop
2 | TECHNICAL SPECIFICATIONS
3 | Lightweight | Yes
4 | Line Powered with transformer | Yes
5 | Handle Power Line Cord Length | More than 2.5 m/ Specify
6 | Otoscope Halogen or Xenon Or LED Lamp (LED Preferred) | 3.5 V
7 | Ophthalmoscope same as Otoscope | Yes
8 | Curved End Probes | Yes
9 | Place to store in units | Yes
10 | Lamp Spare as chosen | 1 For each
11 | PHYSICAL SPECIFICATIONS
12 | Dimensions | Specify
13 | Weight | Specify
14 | CONFIGURATION | Tabletop
15 | TECHNICAL SPECIFICATIONS
16 | Lightweight | Yes
17 | Line Powered with transformer | Yes
18 | Handle Power Line Cord Length | More than 2.5 m/ Specify
19 | Otoscope Halogen or Xenon Or LED Lamp (LED Preferred) | 3.5 V
20 | Ophthalmoscope same as Otoscope | Yes
21 | Curved End Probes | Yes
22 | Place to store in units | Yes
23 | Lamp Spare as chosen | 1 For each
24 | PHYSICAL SPECIFICATIONS
25 | Dimensions | Specify
26 | Weight | Specify
Company Stamp |  |  |  | Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | 4214000006000 | 0 | SCALE PATIENT WITH HEIGHT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Clinical, chair
2 | TYPE | Electronic
3 | MEASUREMENT UNITS
4 | Weight | Metric
5 | Height | Metric
6 | CAPACITY | 250 kg. Approx.
7 | ACCURACY | 0.1 g
8 | CALIBRATION ADJUST | Yes
9 | Power | 220V 60Hz
10 | TARE WEIGHT ADJUST | Yes
11 | OTHER SPECIFICATIONS. | Chargeable Battery with charger
12 | Weight: 14 Kg
Company Stamp |  |  |  | Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | 4214000012200 | 0 | BAG AMBO ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Measure Weight and Body Mass
2 | DISPLAY | Digital, LCD
3 | UNIT DISPLAYED | Kg & lbs
4 | TRANSDUCER TYPE | Load cell / Strain gauge
5 | BODY MASS INDEX FUNCTION | Yes
6 | MAX WEIGHT | Not less than 270 Kg
7 | WEIGHT CHANGE | 0.05 Kg
8 | ACCURACY | ± 2%
9 | ZERO ADJUSTMENT | Automatic
10 | CALIBRATION ADJUSTMENT | Yes
11 | HEIGHT SCALE OPTION | Specify
12 | - Max height | Not less than 200 cm
13 | BATTERY | Yes
14 | - Type | Rechargeable
15 | - Charger | Internal
16 | - Operating time /Charging time | Specify/Specify
17 | The scale should be highly accurate with no need of adjustment or periodic calibration. Accuracy shall be better than or equal to 1 %. Provide detailed characteristics | Yes
18 | Unit shall include a measuring column made from corrosion resistant hard aluminum and shall have a measuring range of approximately 750 mm to 2,000 mm in divisions of approximately 5 mm. | Yes
19 | Measuring rod shall have both inches/cm increments. | Yes
20 | Calibration weight set to be included | Yes
21 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
22 | APPLICATION | Measure Weight and Body Mass
23 | DISPLAY | Digital, LCD
24 | UNIT DISPLAYED | Kg & lbs
25 | TRANSDUCER TYPE | Load cell / Strain gauge
26 | BODY MASS INDEX FUNCTION | Yes
27 | MAX WEIGHT | Not less than 270 Kg
28 | WEIGHT CHANGE | 0.05 Kg
29 | ACCURACY | ± 2%
30 | ZERO ADJUSTMENT | Automatic
31 | CALIBRATION ADJUSTMENT | Yes
32 | HEIGHT SCALE OPTION | Specify
33 | - Max height | Not less than 200 cm
34 | BATTERY | Yes
35 | - Type | Rechargeable
36 | - Charger | Internal
37 | - Operating time /Charging time | Specify/Specify
38 | The scale should be highly accurate with no need of adjustment or periodic calibration. Accuracy shall be better than or equal to 1 %. Provide detailed characteristics | Yes
39 | Unit shall include a measuring column made from corrosion resistant hard aluminum and shall have a measuring range of approximately 750 mm to 2,000 mm in divisions of approximately 5 mm. | Yes
40 | Measuring rod shall have both inches/cm increments. | Yes
41 | Calibration weight set to be included | Yes
42 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
Company Stamp |  |  |  | Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | 4219260200800 | 0 | WHEELCHAIR STANDARD SIZE 24 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Aluminium wheelchair | Yes
2 | Powder coating finish black | Yes
3 | Foldable backrest, aluminium folding mechanism | Yes
4 | Flip back armrest | Yes
5 | Detachable footrest, heel loop | Yes
6 | Nylon black upholstery | Yes
7 | Front castor 8"x1" solid tire | Yes
8 | Steel brake | Yes
9 | Rear wheel 24" x 1 3/8" PU tire , , | Yes
10 | Quick release rear axle | Yes
11 | aluminium hand rim | Yes
12 | brakes | Yes
13 | Anti tipper | Yes
14 | 5 Cm seat cushion | Yes
15 | Capacity 150 Kg | Yes
16 | Wieght 14 Kg | Yes
Company Stamp |  |  |  | Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | 4221161000400 | 0 | CHAIR COMMODE ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Aluminium wheelchair | Yes
2 | Powder coating finish black | Yes
3 | Foldable backrest, aluminium folding mechanism | Yes
4 | Flip back armrest | Yes
5 | Detachable footrest, heel loop | Yes
6 | Nylon black upholstery | Yes
7 | Front castor 8"x1" solid tire | Yes
8 | Steel brake | Yes
9 | Rear wheel 24" x 1 3/8" PU tire , , | Yes
10 | Quick release rear axle | Yes
11 | aluminium hand rim | Yes
12 | brakes | Yes
13 | Anti tipper | Yes
14 | 5 Cm seat cushion | Yes
15 | Capacity 120 Kg | Yes
16 | Wieght 14 Kg | Yes
Company Stamp |  |  |  | Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | 4218190012300 | 0 | INTUBATION DIFFICULT VIDEO ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Easy-to-release arm mechanism allows for safe lateral patient transfers to and from commode
2 | Durable, heavy-duty, gray powder-coated steel tubing
3 | Large, durable snap-on seat
4 | Adjustable height: approx. From 45 to 55 cm
5 | Seat dimensions: approx. 60w x 48 d cm
6 | Commode seat shall be compatible to be placed over most toielts
7 | The offered equipment shall have an approved international certificate( ce,fda,tuv, etc..)
Company Stamp |  |  |  | Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | 4218190003400 | 0 | ANALYZER MARKER CARDIAC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Battery/ electricity operated with complete accessories for use
2 | Capacity to run ROPONIN T / TROPONIN I, MYOGLOBIN, D- DIMMER, NT proBNP / BNP, CK-MB. Specify Range
3 | Sample type: whole blood (capillary, syringe or vacutainer)
4 | Sample volume: 100µL or less
5 | Reportable results: Quantitative measurements
6 | Test time: 2-10minutes
7 | Shall have a Bar code reader system
8 | Automated Internal quality control system
9 | Weight up to 1 kg
10 | Data storage capacity: Equal or greater than 1000 patient results
11 | Ability to interface with wireless technology and LIMS for data storage
12 | Reagent for three months supply for (25 -30) samples per day should be supplied at the time of commissioning
13 | Compliance with standards & legislation:
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | 4218190004300 | 0 | MONITOR CAPNOGRAPHY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA approval or CE marked | yes
2 | Type | rapid diagnostic test system comprised of a meter and various test devices that improve a physician’s ability to aid in the diagnosis of critical diseases and health conditions including heart failure and myocardial infarction as well as aid in
3 | Applications | POC test to improve a physician’s ability to aid in the diagnosis of critical diseases and health conditions including heart failure and myocardial infarction as well as aid in assessing patients for pulmonary embolismMeasure
4 | Test Card for Quantitative measurements of Myoglobin, CK-MB , Troponin I, BNB and D-Dimer as an aid in the diagnosis | Yes
5 | Test Card for Quantitative measurements of Myoglobin, CK-MB and Troponin I as an aid in the diagnosis of myocardial infarction (injury) all in one test. | One Panel/box
6 | BNP in one panel | Yes
7 | Test Card for quantitative immunoassay to aid in assessing and evaluating patients with suspected disseminated intravascular coagulation (DIC) and thromboembolic events including pulmonary embolism (PE) and deep vein thrombosis (DVT). | Yes
8 | Test Card for Quantitative determination of Neutrophil Gelantinase-Associated Lipocalin (NGAL). to provide early information regarding the risk of AKI “Acute Kidney injury” | Yes, specify
9 | Test Card for Cardiorenal assessment | Yes
10 | 10-15 min test for first run then 2 min per test | Yes
11 | Bar-coded test panels / cards | Yes
12 | Compact size | Yes, approx “30X15X6” cm
Company Stamp |  |  |  | Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | 4218190001500 | 0 | THROMBOSIS PREVENTION VIEN DEEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | application | monitoring CO2 measurements for neonatal, pediatric and adult patients
2 | measurements | EtCO2,FiCO2
3 | respiration rate | 0 bpm 150 bpm
4 | spo2 | yes
5 | printer | yes
6 | disply | color screen
7 | wave form trends | yes
8 | trends storage | yes 24hr.,specify
9 | accuracy | specify
10 | alarms | yes, visual and audible,specify
11 | Battery | yes
12 | Battery operating time | hr? specify
13 | power Supply | 220 VAC / 60 HZ
14 | All standard accessories required. | Yes ,Provide Itemized Price
Company Stamp |  |  |  | Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | 4220171200700 | 0 | ULTRASOUND POC GENERAL PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Regulation Standard :
2 | FDA 510 (K) Clearance / CE Mark & SFDA Registration | Yes
3 | Validated Up to Date Certificates Must be Submitted with the Offer | Yes, Failure of Submission and Comply Will Disqualify the Offer
4 | Product Description :
5 | Range | Laptop Ultrasound
6 | Purpose | Used for point of care Critical Care
7 | Console Design :
8 | Laptop / Hand Carried on height adjusted mobility Cart | Cart Should Include the Following :Integrated Power Supply , Probe Ports , wheels with locking braking system
9 | Total Weight - (incl. battery) | ≤ ( 5 ) kg .
10 | Easy Acess mode keys | Yes, (Specify)
11 | Screen Monitor | ≥ (10) inch. LCD with High Resolution ( maximum active screen size ratio )
12 | Security Lock | Yes
13 | Built-In Battery | Yes
14 | Battery scanning time | 1 hour
15 | Power Requirement | 220-240 V / 60 Hz
16 | Computing Capabilities :
17 | Operation System | Windows or equivalent
18 | Integrated Hard Desk type | Flash memory≥8 GB for better shock absorbence
19 | Fast System boot up | Yes, <=30 sec
20 | DVD /CD burner/USB port | Yes
21 | Clinical Applications : | Analysis Packages Standard
22 | Vascular Access and Nerve Block | Yes
23 | Other Available Applications | Yes
24 | Multi Frequency Probes : | One Each , With dedicated Operational Software
25 | Linear array | 5-10 MHz ,
26 | Convex array | 2-5 MHz,
27 | Phased array | 1-5 Mhz,
28 | Active Probe Ports | 2 Active Ports
29 | Operating Modes :
30 | Brightness Mode (B-Mode) (2D) | Yes
31 | Motion Mode (M-Mode) | Yes
32 | Color Doppler mode | yes
33 | Pulsed wave Doppler (PW) with HPRF | Yes
34 | High PRF Doppler Mode | Yes
35 | Color Flow Doppler Mode (CFM) | Yes
36 | Power Doppler Mode | Yes
37 | Continuous Wave Doppler | Yes
38 | Scanning Parameters :
39 | Imaging Depth, cm | Not less than 30 cm
40 | Dynamic Range | Up to 150 dB
41 | TGC Control | Yes
42 | Frame Rate 2D | Specify
43 | Frame rate Color Doppler | Specify
44 | GrayScale Levels | Specify
45 | Scanning Automatic Optimization | Capability for Auto optimization in 2D & Doppler
46 | Digital Calipers | Yes
47 | Preprocessing | Yes
48 | Postprocessing | Yes, (Measurement, adjust the Dynamic Range, TGC, color map & Gain) on stored images
49 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
50 | Doppler :
51 | Type | PW in all probes
52 | Frequency Display | Yes
53 | Doppler baseline shift after freeze | Yes
54 | Velocity Display | Yes
55 | User – defined annotation | Yes
56 | GrayScale Levels | Yes
57 | Transmit focal zones | Specify
58 | Digital reconstructed zoom | Specify
59 | Gain Control after freeze | Yes
60 | Adaptive Gain Compensation | Yes
61 | Directional color power Doppler | Yes
62 | Cine Features :
63 | Cine Storage | Yes, (Still images / Clips / Volume)
64 | Length / Capacity | 140 MB (up to 3000 images)
65 | Scanning Features :
66 | Speckle Reduction Imaging | Yes, Organ Specific
67 | Trapeziodal Imaging / Virtual Convex | Yes
68 | B-Mode Steering | Yes
69 | Compound Imaging | Yes
70 | Tissue Harmonic Imaging | Yes
71 | TGC (Time Gain Compensation) | Yes
72 | Color flow mapping | Standard
73 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
74 | Automatic Doppler Calculations | Yes
75 | Adjustable Transmit Focus | Yes
76 | Dynamic Receive Focus | Yes
77 | Selectable Dynamic Range | Yes
78 | Zoom / Pan :
79 | Real- time image | Yes
80 | Frozen image | Yes
81 | Connectivity :
82 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicable
83 | Wireless Connection | Yes
84 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
85 | Storage and Playback of M-mode, PW Spectral Doppler. | yes
86 | Printing of Patient Reports Via a USB Connection. | Yes
87 | Recalculation for stored images | Yes
88 | Security / Privacy :
89 | Accessories :
90 | Secondary Battery | Yes
91 | Carrying Case | Yes
92 | MOH Requirement :
93 | All the vendors should supply all available options of their machine not stated in the above specifications as an optional items with separate price. | Yes
94 | Supplier is an authorized representative of the manufacturer by SFDA | Yes (Attach official documents)
95 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | 4218190300100 | 0 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | NONINVASIVE HEMODYNAMIC MONITORING SYSTEM | YES
2 | THE SYSTEM MUST BE ABLE TO CONTINUOUSLY MEASURE PARAMETERS | YES
3 | SV/SI (STROKE VOLUME / INDEX) | YES
4 | CO/CI (CARDIAC OUTPUT / INDEX) | YES
5 | HR (HEART RATE) | YES
6 | CARDIAC OUTPUT TECHNOLOGY & ALGORITHM | YES
7 | SHOULD HAVE MULTIPLE SCREENS DISPLAYING VARIOUS PARAMETERS AND THE PHYSIORELATION SCREEN | YES
8 | SHOULD DISPLAY THE PERIOPERATIVE GOAL DIRECTED THERAPY SCREEN, | YES
9 | CONTINUOUSLY MEASUREMENT OF HAEMDYNAMIC PARAMETERS | YES
10 | SYSTEM SHOULD AUTO-CALIBRATE USING PATIENT DEMOGRAPHICS | YES
11 | SVV (STROKE VOLUME VARIATION) – FLUID RESPONSIVENESS INDICATOR | YES
12 | CALCULATE SVR (SYSTEMIC VASCULAR RESISTANCE / INDEX) | YES
13 | CALCULATE DO2/DO2I (DELIVERED OXYGEN) | YES
14 | THE SYSTEM SHOULD BE ABLE TO DISPLAY AND PERFORM FLUID RESPONSIVENESS TEST (SV AND PLR) WHEN OPERATED | YES
15 | THE SYSTEM SHOULD BE USABLE FOR ALL DEPARTMENTS - ICU, OR , ER AND WARDS | YES
16 | THE TECHNOLOGY USED SHOULD HAVE OPTION OF TOTALLY NON-INVASIVE TO FACILITATE THE USE IN EMERGENCY SITUATIONS | YES
17 | EACH PARAMETER SHOULD BE DISPLAYED WITH IT'S NORMAL RANGE AND THE RANGE SHOULD BE ADJUSTABLE IF REQUIRED | YES
18 | THE MONITOR SHOULD BE COMPACT WITH EASY SETUP. | YES
19 | THE MONITOR SHOULD BE CONNECTABLE TO A SOFTWARE PROVIDES PARAMETER TRENDING, PASSIVE LEG RAISING TEST, PRINTING, ...ETC | YES
20 | THE MONITOR SHOULD BE TRANSPORTABLE AND EASY TO USE ON STAND OR TROLLEY | YES
21 | THE MONITOR SHOULD HAVE A BUILT-IN BATTERY AND STORAGE MEMORY FOR 72 HOURS PER PATIENT | YES
22 | BIDDER SHALL LIST ALL THE AVAILABLE ACCESSORIES | YES
23 | COMPLIANCE WITH STANDARDS & LEGISLATION: | YES
24 | SHOULD HAVE A FDA APPROVAL AND/OR CE MARK & SFDA REGISTRATION, WHERE APPLICABLE. LIST ANY OTHER INTERNATIONAL STANDARDS (CE, UL, TUV, CSA), IF ANY. | YES
25 | ALL ELECTRICAL CONNECTIONS AND PLUGS SHOULD BE HOSPITAL GRADE AND FOLLOW INTERNATIONAL, LOCAL AND HOSPITAL REQUIREMENTS. | YES
26 | PROVIDE HARD/SOFT COPIES OF THE OPERATION AND MAINTENANCE MANUALS AS PER THE TENDER TERMS AND CONDITIONS | YES
27 | ALL OTHER BASIC ACCESSORIES DEEMED NECESSARY THAT ARE NOT MENTIONED IN THIS SPECIFICATION BUT ARE REQUIRED FOR FULL FUNCTION AND HIGHEST CLINICAL OUTCOMES AND OUTPUT OF THE EQUIPMENT MUST BE INCLUDED. | YES
28 | SHOULD BE ACCURATE AND USER INDEPENDEDNT TO GET ALL PARAMETERS AND REDUCE THE NEED OF CLOSE CONTACT WITH PATIENT | YES
29 | SCREEN SOULD BE 10.4 INCHES APPROX COLOUR LCD TOUCH SCREEN MONITOR, WINDOWS OS | YES
30 | AUDIBLE ALARM WHEN PARAMETERS GO OUT OF RANGE AND THE ALARM SHOULD BE ABLE TO BE MANUALLY SILENCED | YES
31 | VENDOR MUST PROVIDE FREE DEVICE WITH SERTAIN NUMBERS OF SENSORS | YES
32 | 220 VOLTS / 60 HZ | YES
33 | CART MUST BE PROVIDED WITH BASKET | YES
34 | PRINTER | OPTIONAL
Company Stamp |  |  |  | Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | 4218190014800 | 0 | MONITOR BEDSIDE CENTRAL 10 BEDS ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | ADULT , PEDIATRIC , NEONATE
2 | CENTRAL STATION | YES
3 | NO. OF PATIENT | LISCENCE FOR 10 BEDS
4 | DISPLAY MONITOR | HIGH RESOLUTION COLOR NOT LESS THAN 19 INCH MEDICAL GRADE, FLAT SCREEN TFT. TYPE
5 | NUMBER OF MONITORS | 1 SCREEN WITH CONFIGURABLE NUMBER OF PATIENT , SPECIFY
6 | TRENDING FULL DISCLOSURE | NOT LESS THAN 72 HRS.
7 | OPERATION | FIXED KEY AND MENU DRIVEN/TOUCH SCREEN, OR OTHER - SPECIFY
8 | COMMUNICATION | NETWORK, CENTRAL STATION, REMOTE BEDSIDE MONITOR
9 | BED TO BED COMMUNICATION | YES
10 | CONFIGURABLE REPORTING | YES
11 | ECG INTERPRETATION | YES
12 | COMPUTER MEDICAL GRADE | YES
13 | KEYBOARD AND MOUSE. | WIRELESS OR WIRED SPECIFY
14 | RECORDER | 4 CHANNEL, DIGITAL
15 | RECORDING | YES, ONE FOR THE SYSTEM
16 | LASER PRINTER | YES, ONE FOR THE SYSTEM
17 | DATA | ALL, PATIENTS
18 | WAVEFORMS | YES
19 | STRIP CHART | YES
20 | REMOTE OPERATION | FROM BEDSIDE MONITOR
21 | ALARM - AUDIO AND VISIBLE | YES
22 | ALARM LAMP AND FLASHING | YES
23 | THE SYSTEM SHOULD BE BASED ON THE LATEST TECHNOLOGY IN MONITORS THAT COVER ALL APPLICATIONS NEEDED IN ICU. | YES
24 | THE MONITOR SHOULD HAVE INTENSIVE CARE SOFTWARE FOR ADULT, PEDIATRIC. | YES
25 | THE MONITOR SHOULD PROVIDE A HIGH LEVEL OF ACCURACY OF MONITORED PARAMETERS. | YES
26 | THE MONITOR SHOULD BE MOBILE STAND WITH LOCKABLE WHEELS (MINIMUM OF 5), WITH BASKET, | YES
27 | THE MONITOR SHOULD PROVIDE A HIGH LEVEL OF ACCURACY (E.G. 2% APPROXIMATE FOR ECG) OF ALL MONITORED PARAMETERS. VENDORS SHOULD PROVIDE PARAMETERS ACCURACY DETAILS. | YES
28 | THE MONITOR SHOULD BE MODULAR WITH SINGLE OR MULTI-PARAMETERS PER MODULE. | YES
29 | BEDSIDE MONITOR: | 10 MONITORS
30 | THE DISPLAY SHOULD BE COLOR LCD AND SHOULD HAVE AN APPROXIMATE SIZE OF 19 INCHES WITH A HIGH LEVEL OF CONTRAST, BRIGHTNESS, RESOLUTION, AND VISIBILITY FROM A DISTANCE AND AT A WIDE ANGLE. MONITORED PARAMETERS VALUES SHOULD BE LARGE ENOUGH TO BE | YES
31 | READ FROM A DISTANCE.
32 | THE MONITOR SHOULD BE A TOUCHSCREEN. | YES
33 | THE MONITOR SHOULD BE CAPABLE OF DISPLAYING SIX (6) PHYSIOLOGICAL WAVEFORMS SIMULTANEOUSLY AS A MINIMUM. | YES
34 | THE MONITOR SHOULD BE CAPABLE OF INTERFACING WITH THE HOSPITAL AND BEDSIDE CLINICAL INFORMATION SYSTEM (HCIS). IT SHOULD BE ABLE TO DOWNLOAD ANY RECORDED PARAMETER, WAVEFORMS AND NUMERICAL DATA, TO THE HCIS. THE VENDOR SHOULD SUBMIT A LIST OF | YES
35 | COMPATIBLE HCIS. (PLEASE PROVIDE IT AS ITEMIZED PRICE)
36 | THE VENDOR SHOULD PROVIDE STATE OF THE ART TECHNOLOGY FOR INFORMATION TECHNOLOGY USED IN THE MONITOR APPLICATIONS. | YES
37 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL CLINICAL SOFTWARE THAT IS USED ON THE MONITOR (VENDOR SHOULD SUPPORT DOCUMENTS WITH EVIDENCE THAT PROVE THE BENEFIT OF THE CLINICAL PROGRAMS) | YES
38 | THE UNIT SHOULD BE WELL CONSTRUCTED WITH DURABLE MATERIALS TO WITHSTAND TYPICAL ABUSE, HEAVY USED ENVIRONMENT. THE VENDOR SHOULD PROVIDE SUPPORT DOCUMENT THAT PROVES THE DURABILITY OF HIS PRODUCT (INTERNATIONAL REFERENCE SITE SHOULD BE MENTIONED). | YES
39 | THE MONITOR SHOULD BE CAPABLE OF PROVIDING AN AUXILIARY ECG OUTPUT SIGNAL FOR DEFIBRILLATOR/IABP SYNCHRONIZATION, AND AN AUXILIARY BP OUTPUT SIGNAL FOR IABP SYNCHRONIZATION (PREFERABLE). | YES
40 | THE MONITOR SHOULD PROVIDE USER-ADJUSTABLE VISUAL AND AUDIBLE ALARMS TAILORED FOR INDIVIDUAL PARAMETERS. | YES
41 | THE MONITOR SHOULD BE CAPABLE OF PROVIDING GRAPHICAL AND NUMERICAL 24-HOURS TRENDS OF ALL MONITORED PARAMETERS. | YES
42 | THE MONITOR SHOULD BE SOFTWARE-UPGRADEABLE TO INCLUDE FUTURE UPDATES AND/OR ADDITIONS OF SOFTWARE APPLICATION PACKAGES. | YES
43 | THE VENDOR SHOULD PROVIDE DETAILED TECHNICAL INFORMATION FOR THE UNIQUE FEATURE IN THE MONITOR. | YES
44 | THE VENDOR SHOULD PROVIDE DETAILED TECHNICAL INFORMATION ABOUT THE FOLLOWING: | YES
45 | HR ACCURACY | YES
46 | ECG HR RANGE, BPM | YES
47 | INTERPRETATION | YES
48 | ARRHYTHMIA DETECT TECHNIQUE AND NUMBER OF LEAD ANALYZED | YES
49 | ST TECHNIQUE AND NO. OF LEADS ANALYZED | YES
50 | RESPIRATION METHOD AND WAVEFORM DISPLAYED FEATURE | YES
51 | DISPLAY SIZE, (CM) IN | YES
52 | DISPLAY USER INTERFACE FEATURE | YES
53 | REMOTE DISPLAY AVAILABILITY | YES
54 | TRENDING PARAMETERS | YES
55 | NETWORKING AND CENTRALIZATION. | YES
56 | HARDWIRED/WIRELESS | YES
57 | COMMUNICATION PROTOCOLS | YES
58 | ALARM TYPES AND FEATURE FOR THE MACHINE | YES
59 | THE MONITOR SHOULD BE CAPABLE OF RECORDING AND DISPLAY IN WAVEFORMS AND NUMERICAL VALUES THE FOLLOWING PARAMETERS (PLEASE PROVIDE IT AS ITEMIZED PRICE): | YES
60 | ECG, WITH 5 LEADS DISPLAY, DETECTION OF PACEMAKER SPIKES, AND ARRHYTHMIA. | YES
61 | RESPIRATION, WITH WAVEFORM DISPLAY AND ALARMS. | YES
62 | INVASIVE PRESSURE (IP). | YES
63 | PULSE OXIMETRY (SPO2), WITH NELLCOR-OXYMAX OR MASSIMO RD SET CABLES AND FINGER PROBES. | YES
64 | NON-INVASIVE BLOOD PRESSURE (NIBP), AND QUICK-CONNECT HOSES. | YES
65 | TEMPERATURE, IN TWO CHANNELS FOR TWO SIMULTANEOUS PROBES, SKIN AND RECTAL. | YES
66 | END-TIDAL CO2 (ETCO2), WITH WAVEFORM DISPLAY, RAPID WARM-UP AND CALIBRATION TIMES, AND LOW DEAD-SPACE CONNECTORS. | YES
67 | CARDIAC OUTPUT. | YES
68 | LEVEL OF CONSCIOUSNESS. | YES
69 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR DIFFERENT MODULES. | YES
70 | THE MONITOR SHOULD HAVE A MODULE THAT CAN MEASURE AT THE SAME TIME: | YES
71 | ECG, RESP. | YES, QTY: 1
72 | PULSE OXIMETRY (SPO2), NELLCOR-OXYMAX OR MASSIMO. | YES, QTY: 1
73 | NON-INVASIVE BLOOD PRESSURE (NIBP). | YES, QTY: 1
74 | INVASIVE BLOOD PRESSURE (IBP). | YES, QTY: 2
75 | TEMPERATURE. | YES, QTY: 1
76 | ALL PATIENT INPUT SHOULD BE DEFIBRILLATOR-PROTECTED. | YES
77 | ALL CABLES, PROBES, AND TRANSDUCERS (EXCEPT FOR DISPOSABLES) FOR ALL REQUIRED PARAMETERS SHOULD BE PROVIDED WITH THE MONITOR, FOR ADULTS AND PEDIATRIC. | YES
78 | ALL HARDWARE AND ACCESSORIES FOR EACH MODULE SHOULD BE DESIGNED FOR ALL TYPE OF PATIENT USE (PROVIDE ITEMIZED PRICE). | YES
79 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL OPTION, ACCESSORIES ON THE MACHINE. | YES
80 | THE MANUFACTURER SHOULD GUARANTY THAT SPARE PARTS AND TECHNICAL SUPPORT WILL BE PROVIDED FOR AT LEAST TEN YEARS. | YES
81 | HARDWARE AND SOFTWARE SHOULD BE INDUSTRY STANDARD WITH THE LATEST SPECIFICATION; THE VENDOR SHOULD PROVIDE A SPECIFICATION OF BOTH HARDWARE AND SOFTWARE USED IN THE SYSTEM. | YES
82 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL OPTIONS IN THE SYSTEM. | YES
83 | THE UNIT SHOULD BE PROVIDED WITH A LINE (POWER) CORD OF ACCEPTABLE DURABILITY, QUALITY, LENGTH, AND SHOULD BE SECURED WITH ADEQUATE STRAIN RELIEFS. | YES
84 | THE CHASSIS SHOULD BE GROUNDED AND GROUNDING RESISTANCE SHOULD NOT EXCEED 0.15 OHM. | YES
85 | IF THE UNIT IS DOUBLE INSULATED, IT SHOULD BE SO LABELED. | YES
86 | SWITCHES, KNOBS, AND OTHER CONTROLS SHOULD BE VISIBLE AND CLEARLY IDENTIFIED, AND THEIR FUNCTIONS SHOULD BE SELF-EVIDENT. | YES
87 | SWITCHES, KNOBS, AND OTHER CONTROLS SHOULD BE PROTECTED AGAINST ACCIDENTAL SETTING CHANGES AND SHOULD BE SEALED AGAINST PENETRATION OF LIQUIDS. | YES
88 | ALL AVAILABLE OPTIONS, CONSUMABLES, OR DISPOSABLES DEEMED NECESSARY FOR THE INTENDED FUNCTION OF THE EQUIPMENT AND NOT LISTED IN THE TECHNICAL SPECIFICATIONS SHOULD BE INCLUDED AND SPECIFIED. | PROVIDE ITEMIZED PRICE
89 | WALL MOUNT BRACKETS AS ITEMIZED PRICE | PROVIDE ITEMIZED PRICE
90 | ETCO2 ITEMIZED | PROVIDE ITEMIZED PRICE
91 | CENTRAL STATION ITEMIZED (LICENSE OF 20 BEDS) | PROVIDE ITEMIZED PRICE
92 | ITEMIZED: THE MONITOR SHOULD HAVE A MODULE THAT CAN MEASURE 8 WAVES AT THE SAME TIME AS FOLLOWS : | PROVIDE ITEMIZED PRICE
93 | QTY: 1 ECG, RESP. | PROVIDE ITEMIZED PRICE
94 | QTY: 1 PULSE OXIMETRY (SPO2), NELLCOR-OXYMAX OR MASSIMO. | PROVIDE ITEMIZED PRICE
95 | QTY: 1 NON-INVASIVE BLOOD PRESSURE (NIBP). | PROVIDE ITEMIZED PRICE
96 | QTY: 3 INVASIVE BLOOD PRESSURE (IBP). | PROVIDE ITEMIZED PRICE
97 | QTY: 2 TEMPERATURE. | PROVIDE ITEMIZED PRICE
98 | ELECTRICAL LEAKAGE CURRENT FROM THE CHASSIS OF THE SYSTEM SHOULD NOT EXCEED [0.5 MA PER IEC 601-1 OR 0.3 MA IN THE U.S. PER NFPA 99-1993].OPERATOR SAFETY AND SYSTEM PERFORMANCE SHOULD NOT BE ADVERSELY AFFECTED BY FLUID SPILLS. | YES
99 | THE SYSTEM PERFORMANCE SHOULD NOT BE AFFECTED BY EMI RADIATED OR CONDUCTED THROUGH THE POWER LINES FROM ANOTHER DEVICE. | YES
100 | THE POWER CORD SHOULD BE AUTOMATIC RETRACTABLE. | YES
101 | GENERAL SPECIFICATIONS AND REQUIREMENTS: | YES
102 | POWER SUPPLY: 230 VAC 60 HZ. | YES
103 | HOSPITAL GRADE BRITISH TYPE POWER PLUGS, LINE CORDS, AND STRAIN RELIEF. | YES
104 | THE UNIT MUST MEET THE APPLICABLE REQUIREMENTS AND STANDARDS OF THE UNDERWRITERS LABORATORIES (UL), CE MARK, IEC, ISO, AND/OR FDA APPROVAL. | YES
105 | THE UNIT SHOULD BE MANUFACTURED ACCORDING TO GMP GUIDELINES. | YES
106 | LABELED WITH ALL APPROPRIATE OPERATION AND SAFETY TAGS AND SYMBOLS. | YES
107 | THE UNIT SHOULD BE SIMPLE TO OPERATE AND EASY TO CLEAN AND DISINFECT WITH ISOPROPANOL ALCOHOL 70%. | YES
108 | THE EXTERIOR SHOULD BE WELL CONSTRUCTED WITH DURABLE FLUID/SHOCK RESISTANT MATERIALS TO WITHSTAND TYPICAL ABUSE AND CLEANING. | YES
109 | RUGGED TO HANDLE IN THE ROUGH TRANSPORT ENVIRONMENT. | YES
110 | THE UNIT SHOULD HAVE NO SHARP EDGES. ALL EXTERNAL COMPONENTS SHOULD BE SECURELY MOUNTED. | YES
Company Stamp |  |  |  | Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | 4115000007300 | 0 | CENTRIFUGE 48 TUBES TABLE TOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | To incorporate digital display for speed (actual and set), RCF and time, as well as other operating indicators such as lid locked, door closed/open, status, etc. (list details)
2 | Manually adjustable parameters as well as user programming capability. State number of user defined programs and programmable parameters such as speed, time, acceleration, rotor type, brake, etc..
3 | Variable centrifugation speed up to ~ 6000 rpm
4 | Variable acceleration with acceleration and brake ramping. Specify range.
5 | Variable Relative Centrifugation Force. Specify range.
6 | Variable timer up to 60 min or better as well as continuous operation
7 | Acceleration/deceleration specs should be provided.
8 | Swing out rotor with four buckets (and corresponding inserts) that can be used as open or closed. Sealing lid shall be included for each bucket.
9 | Rotator tube capacity up to 48 tubes
10 | To incorporate air circulation system to prevent excessive temperature increase within the centrifugation chamber
11 | The unit shall possess advanced features and characteristics, especially those relating to sample and user safety. These features shall include but not be limited to:
12 | Brushless induction drive
13 | SS bowl and lid interior
14 | Easily removable rotor for cleaning the bowl
15 | Internal safety chamber (steel or similar) between the bowl and outer case
16 | Insulated interior for noiseless and vibration-free operation with NOISE LEVEL < 55 dB
17 | Smooth, low noise motor (state motor suspension method)
18 | Rubber positioning feet to prevent the unit from slipping
19 | Lid opening prevention mechanism when the rotor is turning
20 | Emergency stop capability
21 | Rotor imbalance detection system
22 | Audiovisual alarm with identification of problem or error code for easy troubleshooting. List all alarms.
23 | Electromagnetic lid locking system with power fail manual override
24 | Each of the following additional features will be considered as an asset during evaluation of the offered unit. State availability as well as characteristics where applicable:
25 | Pulse button
26 | Automatic rotor identification and verification with auto program
27 | Over speed detection system
28 | Shall accept tube Sizes 12, 15 and16 ML
29 | Compliance with standards & legislation:
30 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
31 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
32 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
33 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
34 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | 4115000021900 | 0 | MIXER BLOOD ROLLER TYPE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Rolling and tilting tube mixer for homogeneous, gentle mixing of hematology blood sample tubes
2 | Rolling only and rolling and tilting modes of operation should be allowed
3 | Fixed speed at ~ 40 rpm
4 | Roller length ~ 35 cm
5 | Vertical movement ~ 16 mm
6 | To incorporate ~ 5 rollers for a total capacity of ~ 15 tubes
7 | Individual rollers should be removable to accommodate larger bottles
8 | Compliance with standards & legislation:
9 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
10 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
11 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
12 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
13 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | 4115000024500 | 0 | PIPETTOR AUTOMATIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Light weight and easy to use
2 | Automatic tip recognition
3 | LCD display
4 | Adjustable dispensing volumes
5 | Battery-operated
6 | Shall have three operational modes:
7 | Pipette
8 | Dispense
9 | Auto-dispense
10 | Shall specify the range volume, increment, and precision
11 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
12 | Shall list all the necessary accessories for proper installation and functioning
Company Stamp |  |  |  | Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | 4115000002000 | 0 | ANALYZER CHEMISTRY POC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | single-test/ pannel tests clinical chemistry system which allows the measurement of clinical parameters from whole blood, plasma or serum – including liver and pancreas enzymes, metabolites, blood lipids, hemoglobin and potassium.
2 | To incorporate a LCD screen/ touch screen and alphanumeric keyboard for user interfacing. Provide detailed specs (touch control, size, color or mono, resolution, etc.)
3 | To incorporate a built-in thermal printer/ external printer. Provide detailed specs.
4 | Throughput shall be max. 60 samples/hr
5 | Aanalysis time < = 60 sec
6 | Sample volume: 30 µL
7 | TEST PANEL
8 | a) Kidney diseases
9 | b) Liver diseases
10 | c) Lipid disorders
11 | d) Diabetes
12 | e) Anaemia
13 | Measuring ranges:
14 | Glucose 0.56 – 33.3 mmol/l
15 | Choesterol 2.59 – 12.9 mmol/l
16 | HDL Cholesterol 0.26 – 2.59 mmol/l
17 | Triglycerides 0.80 – 6.86 mmol/l
18 | Creatinine 44.5 – 884 μmol/l
19 | Urea 3.33 – 50mmol/l
20 | Uric aicd 120 – 1190 μmol/l
21 | Haemoglobin 3.1 – 12.4 mmol/l
22 | GPT (ALT) 5.0 – 2000 U/l (37oC)
23 | GOT (AST) 5.0 – 500 U/l (37oC)
24 | Light sources | Light-emitting diodes (LEDs)
25 | Wavelengths | 567 nm, 642 nm, 951 nm
26 | Light detectors | Photodiodes (2)
27 | Accuracy | ± 0.5 % reflectance
28 | Precision | ≤ 0.2 % reflectance
29 | QUALITY CONTROL KITS
30 | Specify test throughput (analysis time) from “READY” state as well as start up time from off to ready
31 | Automatic calibration. State levels, schedule (or time interval) and duration. Calibration interruption for STAT samples
32 | Internal troubleshooting (self diagnostic) software capability is an asset. Specify details.
33 | RS232 interface
34 | Data management capabilities
35 | The system should incoporate the following features: Display, printer, bar code reader
36 | Shall specify the exact dimension and weight for the offered equipment
37 | Compliance with standards & legislation:
38 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
39 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
40 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
41 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
42 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | 4220180001300 | 0 | X-RAY DIGITAL MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | GENERAL STANDARD
2 | (FDA 510 (K) Clearance OR CE Mark) & SFDA Registration | Yes
3 | MOBILE X-RAY UNIT
4 | Microprocessor Controlled | Yes
5 | Forward & Reverse Motor Driven | Yes, Self Propelled Handle Drive Control
6 | Max incline, degree | ( 5-7 ) Degree is Preferable
7 | Speed | Adjustable from ( 0.5 - 1.3 ) m/s
8 | AC line | Yes
9 | Line-voltage Compensator | Yes / Automatic
10 | Type of Power Sources | Both AC & Battery, The Generator Should Work on AC in Case of Battery Failure (Preferable)
11 | Power Requirements | 220 Volt / 60 Hz.
12 | Single or Dual Battery System | Two Independent Systems Provide Power for Driving Motors & Imaging Exposure (Preferable)
13 | Battery Type | Lead acid sealed lead / NiCad
14 | Battery Charging Time - ( From Empty to Full Charged ) | Not more than ( 10 ) hrs
15 | Low-Battery Indicator | Yes
16 | Battery Power Storage | Specify number of maximum oprating hours
17 | Power Cord Length, m | Not less than ( 2.5-5 ) m or More
18 | Wired Handheld Switch & Cordless Remote control Switch | Yes, Exposure By Using Remote Control for Optimal Radiation Protection
19 | Manual Movement in Case of Battery or Motor Failure | Yes
20 | Remote Control Operating range | Not Less than ( 7 ) m
21 | Anti-Collision Protection System or Equivalent | Yes
22 | System Weight | Specify
23 | Mobile X-ray can be Moved Forward & Backward by using the Bedside Drive Controls Located on the Collimator or on the Articulated Arm | Preferred
24 | X-RAY TUBE
25 | Rotating Anode | Yes
26 | Anode Heat Storage Capacity | Not Less Than ( 120 ) kHU
27 | Maximum output Voltage, kVp | Approx. ( 140 ) KV , ( ±10 )
28 | Nominal Focal Spot Value, mm | ( 0.8 / IEC ) or Dual Focal Spots
29 | Nominal Anode Input power | Not Less Than ( 30 ) KW
30 | Tube Movement & Angulation | Yes
31 | Clear view while moving (telescopic movement of the column or dropdown) | Yes
32 | Shortest Exposure Time | 2 ms or less
33 | Working Column Movement Range | Rotating ( 180 ) Degree
34 | X-RAY GENERATOR
35 | High Frequency X-Ray Generator | Yes
36 | Nominal power Rating | Not Less Than ( 30 ) KW
37 | Kv Range | Approx. ( 40 - 140 ) Kv ( ±10 ) step of 5 (as minimum)
38 | mA range | Not Less Than ( 300 ) mA
39 | mAs range | Min.: ≤ 0.4 , Max.: ≥ 320
40 | Increments | Step Up / Step Down / Continuous
41 | Digital Indicator meters | Yes, (If applicable) For KVp, mAs, mA, Dose
42 | X-RAY COLLIMATOR
43 | Type | Manual Collimator
44 | Rotation | Approx. ( ±90 ) Degree
45 | Aluminum filter, mm | ≤ (3.5) mm
46 | SID Range, cm | Min.: ≤ 68 , Max.: ≥ 200
47 | Centering indicator (light /laser) | Yes, for Quick & Easy Indication of SID
48 | LARGE DETECTOR ASSEMBLY
49 | Waterproof | Yes
50 | Shock Sensor | Yes
51 | Portable Digital Detector | Wireless
52 | Semiconductor material | Amorphous silicon ( A-Si )
53 | Scintillator | Cesium iodide ( Csl )
54 | Detector Active Area Size | ≥ ( 35 x 43 ) cm, Approx.
55 | Image Depth / A/D Conversion / Acquisition Depth | ( 14 ) Bit
56 | Pixel size | 150 Micron or less is Preferred
57 | Matrix Size / Resolution | Not Less Than ( 2022 X 2022 )
58 | Detective Quantum Efficiency ( DQE ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
59 | Modulation Transfer Function ( MTF ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
60 | Detector Weight Including Battery | Less Than or equal ( 4 ) kg
61 | Max. Load Capacity On Lying Position | More Than ( 100 ) kg
62 | Rechargeable Battery (Built in or Removable) | Lithium-ion , in Case of Removable Second Battery is Required or equivalent
63 | Charging Time | Not More Than ( 4 ) hours for 100% Charge
64 | Battery Operation Time | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
65 | Charging Station | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
66 | Rigidity of Detector Housing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
67 | Detector Handle | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
68 | Detector Sharing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at (90) images per hour
69 | PROCESING & VIEWING WORKSTATION (Operator Console)
70 | Acquisition Workstation | Built in
71 | LCD Touch Screen Monitor | Not Less Than ( 15 ) inch
72 | Radiation Dose Monitoring ( DAP ) | indicate
73 | Had Disk Storage | Not Less Than ( 3000 ) images
74 | Automatic Programming ( APR ) | Yes
75 | IMAGE POST PROCESSING FUNCTIONS
76 | Window Width & Leveling | Yes
77 | Gray Scale Invert / Annotation | Yes
78 | Image Rotate / Free Rotation | Yes
79 | Electronic L/R Marker | Yes
80 | Patient Edit / Emergency Exam | Yes
81 | NETWORKING & SECURITY
82 | DICOM Compatible | Yes
83 | DICOM Store and DICOM structured dose report | Yes
84 | Image auto transfer | Yes
85 | Query, retrieve, Send, Receive | Yes
86 | Modality worklist | Yes
87 | Storage commitment | Yes
88 | DICOM print | Yes
89 | Patient edit | Yes
90 | DICOM MPPS | Yes
91 | DICOM viewer on CD "Burn exam on CD with DICOM viewer" | Yes
92 | HIPAA Patient Data Security | Yes
93 | DATA PROTECTION AGAINST DELETION & export "Different privilege levels " | Yes
94 | Security Package | Yes
95 | Virus Protection | Yes
96 | The ability to interact with the processing workstation from remote PC to view and process studies within the same hospital | Yes
97 | ACCESSORIES
98 | Two light weight Aprons (0.5) Equivalent of ( Non Leaded Material ) | ( Two ) - (medium & large)
99 | One Moveable Detector holder | ( One ) with Vertical, Horizontal & Swivel Movement
100 | Add-on specifications on top of base specs | Vendor input
101 | Extra detector (small)
102 | Extra detector (large)
103 | Above specs with telescopic column movement for X-Ray tube (if available)
104 | Extra detector (small) with exchange price
105 | Extra detector (large) with exchange price
106 | One Mobile Apron Hanger, and must hold 5 aprons
107 | Please specify Add-on specifications on top of base specs
Company Stamp |  |  |  | Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | 4220370000400 | 0 | PROTECTION X-RAY APRON |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | X-ray protective apron made of flexible lightweight material. Specify weight and length for each size
2 | Wrap around model to provide 0.50 mm protective lead equivalency in the front and 0.3 mm in the back
3 | Wide shoulder cut to provide wide weight distribution and reduce inflicted fatigue
4 | Fastening shall be easy and reliable (Velcro type fasteners for example). Specify.
5 | Set of three different sizes
6 | Specify all available colors.
7 | Features ;
8 | High Quality alloy.
9 | Easily cleaned front
10 | Non-slip backing
11 | Stain-resistant binding
12 | Permit hanging
13 | Protective multi-layered
14 | Provides protection secondary (scatter) radiation.
15 | radiological lab inspected and certified.
16 | Souple, drapes comfortably
17 | Lightweight, easy handling
18 | Life cycle and store time :7-10 years specify
Company Stamp |  |  |  | Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | 4220370000500 | 0 | PROTECTION X-RAY APRON VEST & SQUIRT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Lumbar Vest & Skirt apron
2 | Back Flap for better weight balance
3 | Built-in Black Bel tin Vest and Skirt for better lumbar support
4 | Partial overlap to increase protection
5 | Professionally designed for an exceptional fit and comfort
6 | VEST Specification:
7 | Each front panel: .5 mm Pb (1.0 mm Pb at overlap)
8 | Back panel: .25 mm Pb
9 | SKIRT Specifications
10 | Each front panel: .25 mm Pb (.5 mm Pb at overlap)
11 | Specify all available colors
12 | Medium size
13 | Features ;
14 | High Quality alloy.
15 | Easily cleaned front
16 | Non-slip backing
17 | Stain-resistant binding
18 | Permit hanging
19 | Protective multi-layered
20 | Provides protection secondary (scatter) radiation.
21 | radiological lab inspected and certified.
22 | Souple, drapes comfortably
23 | Lightweight, easy handling
24 | Life cycle and store time :7-10 years specify
Company Stamp |  |  |  | Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | 4218200500400 | 0 | OPTHALMOSCOPE WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Advanced, high performance color imaging system for abdominal, vascular, obstetrics, gynecology, neonatal, and small parts applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing, Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R/W Drive and USB
22 | The system shall be suitable for the following applications:
23 | Abdominal
24 | Obstetrical
25 | Gynecological
26 | Musculoskeletal
27 | Vascular
28 | Urological
29 | Small Parts and Superficial
30 | Pediatric and Neonatal
31 | Transcranial Doppler
32 | The system shall possess capabilities for the following scanning methods:
33 | Electronic Sector
34 | Electronic Convex
35 | Electronic Linear
36 | The system shall be capable of accommodating the following transducer types:
37 | Phased array sector
38 | Convex
39 | Micro-convex
40 | Linear
41 | Operating modes shall include but not be limited to:
42 | B-Mode
43 | M-Mode
44 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
45 | Color Flow Mode
46 | Power Doppler Imaging
47 | PW Doppler
48 | M-Color Flow Mode
49 | Steerable CW Doppler
50 | Trapezoidal Imaging
51 | Optional pan view
52 | shear-wave elastography, and contrast-enhanced imaging , Specify.
53 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
54 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
55 | Tilt/rotate adjustment
56 | Digital Brightness/Contrast Adjustment
57 | Integrated Task Light
58 | Slave monitor :
59 | A second slave monitor Wi-Fi connected with the scanner shall also be included in the offer, for patient viewing. All necessary mounting accessories shall be supplied.
60 | System must be offered with an above 19-inch high resolution,
61 | flat panel, medical grade
62 | monitor with wide viewing angles & good colour resolution.
63 | Resolution : 1024x768pixels or better
64 | The user interface panel shall possess ergonomic and friendly functional design, including:
65 | Alphanumeric keyboard for data entry with ergonomic key operations
66 | Interactive backlighting
67 | Indicator lights to identify activated keys
68 | The system cart shall possess features of high stability and maneuverability, including:
69 | On-board Storage for color-printer
70 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
71 | Multiple probe holders
72 | Gel dispenser holder
73 | Rear handle for transport maneuverability
74 | Video connection shall be integrated for film printing by laser camera.
75 | A color small format (A4), high resolution printer shall be included with the offer
76 | The system shall possess extensive display annotations including:
77 | Institution name
78 | Date / Time
79 | Patient Demographics
80 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
81 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
82 | Focal zone markers
83 | TGC Curve (LGC or Lateral Gain Compensation)
84 | Body patterns
85 | System messages display
86 | Trackball functionality status
87 | Biopsy guide line and zone
88 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
89 | h as IM
90 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
91 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
92 | Endo-vaginal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint. (Biopsy guide shall be quoted)
93 | Linear Probe for vascular, small parts, breast, neonatal and pediatrics: specify the frequency range, the number of elements, the convex radius and the FOV
94 | Compliance with standards & legislation:
95 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
96 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
97 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
98 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
99 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 42
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
42 | 4229510002600 | 0 | HEADLIGHT FIBEROPTIC SURGICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Wall mounted | Yes
2 | OPHTHALMOSCOPE | Yes
3 | Lamp type | Specify
4 | Diopter corrections, D | -25 to + 40 approx.
5 | Filters | Red free
6 | Handle 2.5 volts | Yes
7 | Diagnostic beams | Two
8 | 4 Spare bulb | Yes
9 | Power Supply 220 VAC / 60 HZ | Yes.
10 | FDA or CE approved | Yes.
Company Stamp |  |  |  | Signature

## ITEM 43
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
43 | 4229510007000 | 0 | SUCTION UNIT HIGH VACUUM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Light source 300 watt | Yes
2 | Headlight with ball joint and vertical height adjustment | Yes
3 | front selection movable in all directions | Yes
4 | holder with ball joint | Yes
5 | fiber optic light cable, diameter 3.5 mm Length 230cm | Yes
6 | fixation clip | Yes
7 | Power | 220v 60 hz
8 | OTHER SPECIFICATION | All access. needed
Company Stamp |  |  |  | Signature

## ITEM 44
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
44 | 4218190001700 | 0 | DETECTOR BLOOD FLOW |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Stable mobile stand made of lightweight, heavy-duty, antirust metal with 4 antistatic swiveling castors and wide base to prevent accidental tripping.
2 | The pump shall possess (or exceed) the following technical specifications:
3 | Double piston pump continuous rating type.
4 | Pressure guage up to 800 mm Hg
5 | Pressure regulator .
6 | All Castors with brakes .
7 | Over flow protection.
8 | Silicon Tubings.
9 | Suction Bottle Capacity - 2 x 2500 ml minimum (with safety valve)
10 | Bottles graduation scale 100ml - material polysulphone . sterelisable at 134°C .
11 | Vacuum control : ≥ 50 LPM specify
12 | Variable (regulating valve) vacuum pressure from 0 to approximately – 750 mm Hg, with negative pressure gauge indicating actual pressure.
13 | Low noise operation (≤ 45 dB @ 1 meter)
14 | Incorporated filter (specify type)
15 | Foot switch operation .
16 | Hand control .
17 | The offer shall include all the accessories and parts necessary for the full and efficient operation of the system. A detailed list of such standard accessories (with part numbers and quantities) shall be included.
Company Stamp |  |  |  | Signature

## ITEM 45
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
45 | 4229350000800 | 0 | BLANKET WARMING SYSTEM AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DOPPLER WITH 5 MHz PROBE | 1
2 | DOPPLER WITH 8 MHz PROBE | 1
3 | TYPE | Portable / Handheld
4 | APPLICATION | High level vascular assessment, ideal for the detection of peripheral arterial disease (probe sizes neonatal, infant, child, adult)
5 | PROBE FREQUENCY (INTERCHANGEABLE) | 5, 8, 10 MHz ( Probe 10 Mhz compatible with the device)
6 | PROBE CONSTRUCTION | Durable Long Lasting Sealed Probes
7 | PROBE SHAPE / LENGTH | Pencil / 109 mm (approx)
8 | BUILT-IN SPEAKER (200 mW) | High fidelity speaker
9 | POWER CONSUMPTION : | 500 mW or less
10 | BI-DIRECTIONAL LCD DISPLAY | Yes
11 | WAVEFORM CALIBRATION FUNCTION | Yes
12 | GAIN CONTROL | Yes
13 | OBSTETRIC CAPABILITY | Yes/ Neonatal
14 | BATTERY
15 | - Type of battery | Specify
16 | - Operating time, hrs | 250 one min tests
17 | If, rechargeable battery
18 | - Battery life | Specify
19 | - Charger | Yes
20 | AUTO SHUT-OFF : | After 3 min. if left ON
21 | BATTERY CHECK INDICATOR | Yes, (LED).
22 | EXTERNAL OUTPUTS: | Phone-220 mV Impedance 270Ω, Record 500 mV/kHz (3.5 mm jack) Impedance 20 kΩ
23 | HEADSET / HEADPHONE | Yes
24 | PROBE HOLDER | Yes
25 | USAGE | Detecting peripheral pulses
26 | PRINTER | Specify
27 | Computer Interface Or PC-based | Yes
28 | ACCESSORIES :
29 | - Large Carry Bag | with complete Assessment kit
30 | - Probes | Yes
31 | The Device Should Integrate with the current and future HIS. | Yes / Specify
32 | DOPPLER WITH 5 MHz PROBE | 1
33 | DOPPLER WITH 8 MHz PROBE | 1
34 | TYPE | Portable / Handheld
35 | APPLICATION | High level vascular assessment, ideal for the detection of peripheral arterial disease (probe sizes neonatal, infant, child, adult)
36 | PROBE FREQUENCY (INTERCHANGEABLE) | 5, 8, 10 MHz ( Probe 10 Mhz compatible with the device)
37 | PROBE CONSTRUCTION | Durable Long Lasting Sealed Probes
38 | PROBE SHAPE / LENGTH | Pencil / 109 mm (approx)
39 | BUILT-IN SPEAKER (200 mW) | High fidelity speaker
40 | POWER CONSUMPTION : | 500 mW or less
41 | BI-DIRECTIONAL LCD DISPLAY | Yes
42 | WAVEFORM CALIBRATION FUNCTION | Yes
43 | GAIN CONTROL | Yes
44 | OBSTETRIC CAPABILITY | Yes/ Neonatal
45 | BATTERY
46 | - Type of battery | Specify
47 | - Operating time, hrs | 250 one min tests
48 | If, rechargeable battery
49 | - Battery life | Specify
50 | - Charger | Yes
51 | AUTO SHUT-OFF : | After 3 min. if left ON
52 | BATTERY CHECK INDICATOR | Yes, (LED).
53 | EXTERNAL OUTPUTS: | Phone-220 mV Impedance 270Ω, Record 500 mV/kHz (3.5 mm jack) Impedance 20 kΩ
54 | HEADSET / HEADPHONE | Yes
55 | PROBE HOLDER | Yes
56 | USAGE | Detecting peripheral pulses
57 | PRINTER | Specify
58 | Computer Interface Or PC-based | Yes
59 | ACCESSORIES :
60 | - Large Carry Bag | with complete Assessment kit
61 | - Probes | Yes
62 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Signature

## ITEM 46
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
46 | 4229350001600 | 0 | INTUBATION SET COMPLETE CASE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Inflation time | 90 sec
2 | Warming up time | 4 minutes
3 | Length of Air Hoses | 1500 mm
4 | Temperature | 20°C - 42°C, step of 1°C
5 | Pressure | 0 – 60 mb. Step of 5 mb
6 | Alarm | Audio and visual
7 | Hand Controller | Specify
8 | Warm-Air Mattress, Wide | Length 200 cm, Width 60 cm.
9 | Warm-Air Mattress, Small | Length 100 cm, Width 50 cm
10 | Strap | Yes
11 | Hose seal | 2 pcs
12 | Hose set complete | 2 pcs
13 | Itemized price for all accessories | Yes
Company Stamp |  |  |  | Signature

## ITEM 47
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
47 | 4229350001800 | 0 | LARYNGOSCOPE SET COMPLETE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Video Laryngoscope for difficult airway intubation
2 | Design | light weight , portable
3 | Video laryngoscope with LCD screen | yes
4 | Intubation of cervical spine immobilization | yes
5 | Difficult Airway management and routine intubation | yes
6 | Disposable Blades | yes
7 | Range in blade sizes allow coverage patient weights | YES
8 | Monitor unit
9 | Screen Size | LCD or TFT
10 | Cormack-Lehane grades I - IV laryngeal views | yes
11 | Video output capabilities for viewing on external monitors and recording | yes
12 | BLADES
13 | Different sizes and disposable | yes
14 | Made of Polycarbonate, TPE (Thermoplastic Polyurethane Elastomer) | yes
15 | High resolution camera at good Position on blade to protect from blood and secretion | yes
16 | Built in LED light for illumination | yes
17 | Degree of blade angulation | Must be Clinically sutible for Difficult Intubation
18 | Rechargable Battery | yes
19 | Original trolley | yes
20 | Accessories | Any accessories, options and consumable items necessary to operate the offered system(s) must be clearly identified and priced separately
21 | OTHER SPECIFICATION | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 48
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
48 | 4229350002700 | 0 | WARMER BLOOD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile, IV pole mounted
2 | MAXIMUM TEMPERATURE SETTING, °C | Up to 42°C
3 | TECHNICAL SPECIFICATIONS
4 | Heat Exchanger
5 | - Aluminum heat exchanger | Yes
6 | - Technology | Specify
7 | - Maximum Temperature Setting, °C | Up to 42°C
8 | - Increments, °C | 1
9 | - Warm-up time, mins | < 2
10 | Maximum Flow, mL/min, 10°c input to ≥ 37°c output | ≥ 150
11 | Priming Volume, mL | 20 - 150
12 | Display | LCD
13 | - Temperature range, °C | 33 - 42
14 | - Increments, °C | 1
15 | Air Vent or Trap | Manual bubble trap on standard flow sets; Autoventing bubble trap on all high-flow sets.
16 | APPLICATIONS
17 | - IV | Yes
18 | - Heater power, W | 800-850 W (Approx)
19 | - High Temperature Cutoff, °C | 43 °C
20 | - Automatic air elimination on all high flow sets | Yes
21 | - Electronic temperature control | Yes
22 | - Audible, visual & alarm test High Temperature alarms | Yes
23 | PHYSICAL SPECIFICATIONS
24 | Dimensions | Specify
25 | Weight | Specify
26 | CONFIGURATION | Mobile, IV pole mounted
27 | MAXIMUM TEMPERATURE SETTING, °C | Up to 42°C
28 | TECHNICAL SPECIFICATIONS
29 | Heat Exchanger
30 | - Aluminum heat exchanger | Yes
31 | - Technology | Specify
32 | - Maximum Temperature Setting, °C | Up to 42°C
33 | - Increments, °C | 1
34 | - Warm-up time, mins | < 2
35 | Maximum Flow, mL/min, 10°c input to ≥ 37°c output | ≥ 150
36 | Priming Volume, mL | 20 - 150
37 | Display | LCD
38 | - Temperature range, °C | 33 - 42
39 | - Increments, °C | 1
40 | Air Vent or Trap | Manual bubble trap on standard flow sets; Autoventing bubble trap on all high-flow sets.
41 | APPLICATIONS
42 | - IV | Yes
43 | - Heater power, W | 800-850 W (Approx)
44 | - High Temperature Cutoff, °C | 43 °C
45 | - Automatic air elimination on all high flow sets | Yes
46 | - Electronic temperature control | Yes
47 | - Audible, visual & alarm test High Temperature alarms | Yes
48 | PHYSICAL SPECIFICATIONS
49 | Dimensions | Specify
50 | Weight | Specify
Company Stamp |  |  |  | Signature

## ITEM 49
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
49 | 4110301100700 | 0 | REFRIGERATOR MEDICATION 700L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TEMPERATURE
2 | Temperature Control | Microprocessor
3 | Temperature Range | 1 to 12 ⁰C
4 | Factory Preset | 4 ⁰C
5 | Accuracy | ± 1 ⁰C
6 | Average Temperature Stability | ± 3 ⁰C
7 | Temperature Adjustable | Yes
8 | Automatic Defrost | Yes
9 | Cooling & Air Circulation Method | Forced Cooling
10 | FEATURES
11 | Capacity | 25 Cu.Ft. Approx
12 | Exterior Dimension | Specify
13 | Interior Dimension | Specify
14 | Interior Finish | Stainless Steel
15 | Door Type | Single Glass Door
16 | Gasket | Magnetic
17 | Access Ports | Additonal Probes to Map Temperature
18 | Digital Display | Specify
19 | ALARM
20 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temperature, Power Failure, Door Ajar & Condensor
21 | Remote Alarm Connected to Central Alarm System | Yes
22 | Microprocessor with Backup Battery | Yes
23 | Visual / Audio Alarms | Yes
24 | Key-Operated Alarm & Setpoint Security | Lock in Temperature & Setpoint Security
25 | OTHER SPECIFICATION
26 | High Density CFC-Free Urethane Foam Insulation | Yes
27 | Number of Adjustable Shelves & Type | Specify
28 | Medical Grade Hermetically Sealed Compressors | Yes
29 | Upgradeble Memory | Yes
30 | Digital Calibration | Yes
31 | Serial Port | USB
32 | Noise Level | Specify
33 | Interior Led Lights | Yes
34 | ACCESSORIES INCLUDED
35 | Temperature Chart Recorder | Yes
36 | Casters | Heavy Duty for Easy Mobility with Leveling
37 | USB | Yes
38 | Mounting Kit | Yes
39 | Integrated Electromagnetic Access Control | Yes
40 | The Device Should Integrate with the current and future HIS. | Yes / Specify
41 | TEMPERATURE
42 | Temperature Control | Microprocessor
43 | Temperature Range | 1 to 12 ⁰C
44 | Factory Preset | 4 ⁰C
45 | Accuracy | ± 1 ⁰C
46 | Average Temperature Stability | ± 3 ⁰C
47 | Temperature Adjustable | Yes
48 | Automatic Defrost | Yes
49 | Cooling & Air Circulation Method | Forced Cooling
50 | FEATURES
51 | Capacity | 25 Cu.Ft. Approx
52 | Exterior Dimension | Specify
53 | Interior Dimension | Specify
54 | Interior Finish | Stainless Steel
55 | Door Type | Single Glass Door
56 | Gasket | Magnetic
57 | Access Ports | Additonal Probes to Map Temperature
58 | Digital Display | Specify
59 | ALARM
60 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temperature, Power Failure, Door Ajar & Condensor
61 | Remote Alarm Connected to Central Alarm System | Yes
62 | Microprocessor with Backup Battery | Yes
63 | Visual / Audio Alarms | Yes
64 | Key-Operated Alarm & Setpoint Security | Lock in Temperature & Setpoint Security
65 | OTHER SPECIFICATION
66 | High Density CFC-Free Urethane Foam Insulation | Yes
67 | Number of Adjustable Shelves & Type | Specify
68 | Medical Grade Hermetically Sealed Compressors | Yes
69 | Upgradeble Memory | Yes
70 | Digital Calibration | Yes
71 | Serial Port | USB
72 | Noise Level | Specify
73 | Interior Led Lights | Yes
74 | ACCESSORIES INCLUDED
75 | Temperature Chart Recorder | Yes
76 | Casters | Heavy Duty for Easy Mobility with Leveling
77 | USB | Yes
78 | Mounting Kit | Yes
79 | Integrated Electromagnetic Access Control | Yes
80 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Signature

## ITEM 50
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
50 | 4214000010300 | 0 | CABINET MEDICATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | يستخدم لحفظ الادوية. | نعم
2 | يحتوي على 4 أدراج بحد أدنى. | نعم
3 | مدهون بدهان حراري غير قابل للصدأ | نعم
4 | يحتوي على قفل. | نعم
5 | شكل عصري وعملي | نعم
6 | وجود فهرسة | نعم
7 | تقديم ضمان شامل لا يقل عن 5 سنوات | نعم
8 | الابواب زجاجية | نعم
9 | الابعاد | طول 100 سم, عرض 50 سم ,ارتفاع 180 سم
Company Stamp |  |  |  | Signature

## ITEM 51
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
51 | 4219190400000 | 0 | CABINET NARCOTIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Storage of Narcotics
2 | FINISH | Enamelled
3 | NUMBER OF DOORS
4 | door type | piano hinge
5 | locks | Yes separate for each door
6 | DIMENSION | 350mm x 250mm x 200mm H. approximately
7 | OTHER SPECIFICATION
8 | ALARMS
9 | visual | Yes
10 | audible | Yes
11 | POWER
12 | vac
13 | hz
Company Stamp |  |  |  | Signature

## ITEM 52
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
52 | 4618181001100 | 0 | STRETCHER EMERGENCY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Area of Use | Emergency
2 | Stretcher Type | Hydraulic
3 | Position Control | Dual sided foot pedal
4 | Trendelenburg | 12 degree APPROX.
5 | Reverse Trendelenburg | 12 degree APPROX.
6 | Emergency Trendelenburg Foot or Handle pedal | Yes, specify.
7 | Height Adjustment | Yes, specify range
8 | Back Rest | Adjustable Gas Spring Assisted
9 | Restraining straps | Yes
10 | Frame Material | Rigid tubular framework with epoxy painted steel.
11 | Patient platform | X-Ray translucent
12 | Safety Side rails | Integral Fold Down
13 | X-Ray cassette access | Yes
14 | Oxygen Cylinder Holder with Cylinder | Accept Size E
15 | Bumper | Yes
16 | Central Locking Bracking System | Four Corner Activation
17 | IV Pole | Yes
18 | Spare IV rod | Yes
19 | Sockets number
20 | Large Storage Area | Yes
21 | Maximum Patient Weight | 220kg OR more
22 | Four 20 cm Casters | All Lockable Antistatic
23 | Mattres | Yes
24 | Thickness cms | 6.cm or more
25 | Monitor Shelf | Yes
26 | 5th Wheel | Yes
Company Stamp |  |  |  | Signature

## ITEM 53
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
53 | 4115000010000 | 0 | COUNTER BLOOD CELL 18 PARAMETERS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Automated, blood cell counter that provide whole blood, and capillary blood analysis
2 | Digital LCD display of numeric counters
3 | Soft touch keys ergonomically placed for ease of use
4 | One step sample processing with touch screen user interface
5 | Full STAT capability
6 | Approximately 20µL sample size whole blood analysis from both venous and capillary collection
7 | Automated probe wide
8 | Minimum of 50 samples per hour
9 | up to 18 parameters
10 | Quality control (QC) programs
11 | Zero routine maintenance
12 | Compliance with standards & legislation:
13 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
16 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
17 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 54
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
54 | 4220370001200 | 0 | PROTECTION X-RAY THYROID COLLAR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | X-ray protective shield for throat (thyroid collar) made of flexible lightweight material.
2 | Wrap around model to provide 0.50 mm protective lead equivalency all around
3 | Fastening shall be easy and reliable (Velcro type fasteners for example). Specify.
4 | Specify all available colors .
5 | Set of three different sizes
6 | Features ;
7 | High Quality alloy.
8 | Easily cleaned front
9 | Non-slip backing
10 | Stain-resistant binding
11 | Permit hanging
12 | Protective multi-layered
13 | Provides protection secondary (scatter) radiation.
14 | radiological lab inspected and certified.
15 | Souple, drapes comfortably
16 | Lightweight, easy handling
17 | Life cycle and store time :7-10 years specify
Company Stamp |  |  |  | Signature

## ITEM 55
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
55 | 4214000014400 | 0 | TABLE EXAM GENERAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Urodynamic Treatment Table/ Electric
2 | APPLICATION | • Urodynamics • Pelvic Muscle Rehabilitation • Anorectal Manometry • Pelvic Exam
3 | TECHNICAL SPECIFICATIONS
4 | • Shall be ergonomics with improved foam to enhance patient comfort and feel | Yes
5 | • Shall be One touch movement locking mechanism for precise positioning | Yes
6 | • Shall be three motors to control lift, reclining back and pelvic tilt | Yes
7 | • Shall be easy access for elderly and patients with physical difﬁculties | Yes
8 | Colour | Standard / Specify
9 | • Inﬁnite height adjustment from | 16” (40 cm) min. to 41” (105 cm) max.
10 | Single foot switch control | Yes
11 | Working Load | ≥ 205kg
12 | Seat tilt range | 0º - 36º (approx)
13 | Motorized Height & seat &back recling back adjustment | Yes
14 | Lift Mechanism: | Electric
15 | Retractable Wheels | 4
16 | PHYSICAL SPECIFICATIONS
17 | Dimensions | Specify
18 | Cushion Width, cm | Specify
19 | Head Section Length, cm | Specify
20 | Mid Section Length, cm | Specify
21 | Leg Section Length, cm | Specify
22 | Total Length, cm | Specify
23 | Weight | Specify
24 | Standard Accessories
25 | 2- Arm rests 2- Leg rests 2- Leg rest holders 2- Fold up footrests 1- Seat cushion without cut out 1- Cushion with cut out 1- Paper towel holder + maneuvering handle 1- Three motor foot control 1- Funnel & funnel holder 1- Neck pillow | Yes / Specify
26 | TYPE | Urodynamic Treatment Table/ Electric
27 | APPLICATION | • Urodynamics • Pelvic Muscle Rehabilitation • Anorectal Manometry • Pelvic Exam
28 | TECHNICAL SPECIFICATIONS
29 | • Shall be ergonomics with improved foam to enhance patient comfort and feel | Yes
30 | • Shall be One touch movement locking mechanism for precise positioning | Yes
31 | • Shall be three motors to control lift, reclining back and pelvic tilt | Yes
32 | • Shall be easy access for elderly and patients with physical difﬁculties | Yes
33 | Colour | Standard / Specify
34 | • Inﬁnite height adjustment from | 16” (40 cm) min. to 41” (105 cm) max.
35 | Single foot switch control | Yes
36 | Working Load | ≥ 205kg
37 | Seat tilt range | 0º - 36º (approx)
38 | Motorized Height & seat &back recling back adjustment | Yes
39 | Lift Mechanism: | Electric
40 | Retractable Wheels | 4
41 | PHYSICAL SPECIFICATIONS
42 | Dimensions | Specify
43 | Cushion Width, cm | Specify
44 | Head Section Length, cm | Specify
45 | Mid Section Length, cm | Specify
46 | Leg Section Length, cm | Specify
47 | Total Length, cm | Specify
48 | Weight | Specify
49 | Standard Accessories
50 | 2- Arm rests 2- Leg rests 2- Leg rest holders 2- Fold up footrests 1- Seat cushion without cut out 1- Cushion with cut out 1- Paper towel holder + maneuvering handle 1- Three motor foot control 1- Funnel & funnel holder 1- Neck pillow | Yes / Specify
Company Stamp |  |  |  | Signature

## ITEM 56
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
56 | 4229350002800 | 0 | WARMER BLOOD RAPID INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | HEAT EXCHANGER
2 | Technology. Dry heat, sealed water cirduit, water bath, warm forced air, direct radial conduction, and magnetic induction,. | Specify
3 | MAX. TEMP. SETTING degree C | 38 - 42 degree C, Specify
4 | FLOW RANGE, mL/min, With 10OC input to at least 35OC | > OR = 1000
5 | APPLICATIONS
6 | IV , Irrigation | Specify
7 | DISPLAY | LCD, LED, TFT, Specify
8 | Temp range, C. Degree | From 35 TO 41 C Degree
9 | Increments, OC | 0.5 C degree
10 | HEATER POWER, W | Specify
11 | PRIMARY TEMPERATURE
12 | CONTROL | Electronic thermostatic
13 | HIGH-TEMPERATURE CUTOFF, OC | 40 - 43, Others specify
14 | HIGH-TEMPERATURE ALARM Audible, Visual, Alarm test | Specify
15 | ON/OFF SWITCH | Yes
16 | MOUNTING
17 | Freestanding, IV pole attachment | Specify
18 | ACCESSORIES
19 | 500 disposable set, | Yes
20 | life for disposable sets | 6 M0NTH
21 | Power Supply | 220 Vac , 60 Hz
22 | Other Specifications | FDA , CE , Iso Approved
Company Stamp |  |  |  | Signature

## ITEM 57
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
57 | 4110301102600 | 0 | REFRIGERATOR MEDICATION 350L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Medication Storage
2 | Capacity | 17 Cu.Ft. Approx
3 | TEMPERATURE
4 | Temperature Control | Programmable Microprocessor
5 | Temperature Range | 2 to 8 ⁰C
6 | Factory Preset | 4 ⁰C
7 | Accuracy | ± 1 ⁰C
8 | Average Temperature Stability | ± 2 ⁰C
9 | Temperature Adjustable | Yes
10 | Hi-Tech Integrated Sensor | Yes
11 | Automatic Defrost | Yes
12 | Cooling & Air Circulation Method | Forced Cooling
13 | TECHNICAL SPECIFICATION
14 | Exterior MaterialType | Electrostatic Powder Coated Stainless Steel
15 | Interior Materical Type | Stainless Steel
16 | Double Door Sliding Type | Yes, Two Heated Glass Door, Lockable
17 | Safety Lock to Prevent Unauthorised Access | Yes
18 | Gasket | Magnetic
19 | Number of Shelves & Type | 5-Shelf Epoxy Coated St. Steel/Adjustable
20 | Interior Lights | LED
21 | INTEGRATED WIRELESS MODULE
22 | Integrated Wireless Data Logging System | Yes
23 | Digital Display | 7" TFT Display
24 | USB Dowload Data | Records Temperature, Humidity, Alarm & Door Ajar Information
25 | Front Ports | Two USB / Ethernet
26 | Integrate with the Current & Future KFMC HIS System | Yes
27 | ALARM
28 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temperature, Power Failure, Door Ajar & Condensor
29 | Remote Alarm Connected to Central Alarm System | Yes
30 | Microprocessor with Backup Battery | Yes, (Approx. 72 Hrs Power Fail Alarm)
31 | Visual / Audio Alarms | Yes, Flash & Buzzer
32 | Key-Operated Alarm & Setpoint Security | Lock in Temperature & Setpoint Security
33 | OTHER SPECIFICATION
34 | High Density CFC-Free Polyurethane Foam Insulation | Yes
35 | Medical Grade Hermetically Sealed Compressors | Yes
36 | Upgradeble Memory | Yes
37 | Digital Calibration | Yes
38 | Noise Level | ˂ 60 db
39 | ACCESSORIES INCLUDED
40 | Casters | 4-Casters & 2-with Brake.
41 | USB | Yes
42 | Mounting Kit | Yes
43 | Additional Distinguish Features if any | Specify
44 | PHYSICAL SPECIFICATIONS
45 | Dimensions | Specify
46 | Weight | Specify
47 | APPLICATION | Medication Storage
48 | Capacity | 17 Cu.Ft. Approx
49 | TEMPERATURE
50 | Temperature Control | Programmable Microprocessor
51 | Temperature Range | 2 to 8 ⁰C
52 | Factory Preset | 4 ⁰C
53 | Accuracy | ± 1 ⁰C
54 | Average Temperature Stability | ± 2 ⁰C
55 | Temperature Adjustable | Yes
56 | Hi-Tech Integrated Sensor | Yes
57 | Automatic Defrost | Yes
58 | Cooling & Air Circulation Method | Forced Cooling
59 | TECHNICAL SPECIFICATION
60 | Exterior MaterialType | Electrostatic Powder Coated Stainless Steel
61 | Interior Materical Type | Stainless Steel
62 | Double Door Sliding Type | Yes, Two Heated Glass Door, Lockable
63 | Safety Lock to Prevent Unauthorised Access | Yes
64 | Gasket | Magnetic
65 | Number of Shelves & Type | 5-Shelf Epoxy Coated St. Steel/Adjustable
66 | Interior Lights | LED
67 | INTEGRATED WIRELESS MODULE
68 | Integrated Wireless Data Logging System | Yes
69 | Digital Display | 7" TFT Display
70 | USB Dowload Data | Records Temperature, Humidity, Alarm & Door Ajar Information
71 | Front Ports | Two USB / Ethernet
72 | Integrate with the Current & Future KFMC HIS System | Yes
73 | ALARM
74 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temperature, Power Failure, Door Ajar & Condensor
75 | Remote Alarm Connected to Central Alarm System | Yes
76 | Microprocessor with Backup Battery | Yes, (Approx. 72 Hrs Power Fail Alarm)
77 | Visual / Audio Alarms | Yes, Flash & Buzzer
78 | Key-Operated Alarm & Setpoint Security | Lock in Temperature & Setpoint Security
79 | OTHER SPECIFICATION
80 | High Density CFC-Free Polyurethane Foam Insulation | Yes
81 | Medical Grade Hermetically Sealed Compressors | Yes
82 | Upgradeble Memory | Yes
83 | Digital Calibration | Yes
84 | Noise Level | ˂ 60 db
85 | ACCESSORIES INCLUDED
86 | Casters | 4-Casters & 2-with Brake.
87 | USB | Yes
88 | Mounting Kit | Yes
89 | Additional Distinguish Features if any | Specify
90 | PHYSICAL SPECIFICATIONS
91 | Dimensions | Specify
92 | Weight | Specify
Company Stamp |  |  |  | Signature