## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | AED SYSTEM
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | ANALYZER AUTOMATED CHEMISTRY SEMI LIQUID AGENTS TABLE TOP
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | ANALYZER AUTOMATED URINE
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | ANALYZER BLOOD GAS BENHCTOP
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | ANALYZER CHEMISTRY POC
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | ANALYZER ELECTROLYTE
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | ANALYZER MARKER CARDIAC
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | ANALYZER SEMI AUTOMATED COAGULATION
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | APHERESIS UNIT
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | BAG HOLDER BIOHAZARD
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | BED INTENSIVE CARE
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | BEDHEAD UNIT HORIZONTAL
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | BLANKET WARMING SYSTEM AIR
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | CABINET MEDICATION
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | CABINET NARCOTIC
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | CART CRASH
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | CART DRESSING
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | CENTRIFUGE 12 TUBES TABLE TOP
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | CHAIR BLOOD DONATION
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | COUNTER CELL BLOOD POC
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | CURTAIN SINGLE USE SLIDING
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | CYLINDER O2 SIZE D WITH REGULATOR
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | CYLINDER O2 SIZE E WITH REGULATOR
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | DECREASING SYSTEM PATHOGENS
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | DEFIBRILLATOR UNIT
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | DETECTOR BLOOD FLOW
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | ECG 12 LEAD WITH CART
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | FREEZER ULTRA LOW TEMPERATURE -20 TO -40 C
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | GLUCOMETER
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | HEADLIGHT FIBEROPTIC SURGICAL
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | KIT SPILL CHEMICAL
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | LARYNGOSCOPE ADULT & PEDIATRIC
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | LARYNGOSCOPE PEDIATRIC
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | LIGHT EXAM MOBILE
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | MIXER BLOOD ROLLER TYPE
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | MONITOR CAPNOGRAPHY
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | OPTHALMOSCOPE DIRECT
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | OTOSCOPE TABLETOP
ITEM 42 | This sheet contains the details of the bid with each technical specification for | ITEM 42 | OXIMETER PULSE HAND HELD
ITEM 43 | This sheet contains the details of the bid with each technical specification for | ITEM 43 | PIPETTOR ADJUSTABLE
ITEM 44 | This sheet contains the details of the bid with each technical specification for | ITEM 44 | PROTECTION X-RAY APRON VEST & SQUIRT
ITEM 45 | This sheet contains the details of the bid with each technical specification for | ITEM 45 | PUMP INFUSION
ITEM 46 | This sheet contains the details of the bid with each technical specification for | ITEM 46 | PUMP SYRINGE
ITEM 47 | This sheet contains the details of the bid with each technical specification for | ITEM 47 | RACK BLOOD SEDIMENTATION
ITEM 48 | This sheet contains the details of the bid with each technical specification for | ITEM 48 | REFRIGERATOR MEDICATION 700L
ITEM 49 | This sheet contains the details of the bid with each technical specification for | ITEM 49 | REGULATOR SUCTION HIGH WALL MOUNT
ITEM 50 | This sheet contains the details of the bid with each technical specification for | ITEM 50 | SCALE CHAIR
ITEM 51 | This sheet contains the details of the bid with each technical specification for | ITEM 51 | SCALE PATIENT WITH HEIGHT
ITEM 52 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | SCOPE SET RIGID ENT
ITEM 53 | This sheet contains the details of the bid with each technical specification for | ITEM 53 | SPHYGMOMANOMETER ANEROID MOBILE
ITEM 54 | This sheet contains the details of the bid with each technical specification for | ITEM 54 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN
ITEM 55 | This sheet contains the details of the bid with each technical specification for | ITEM 55 | STAND IV
ITEM 56 | This sheet contains the details of the bid with each technical specification for | ITEM 56 | STETHOSCOPE ADULT
ITEM 57 | This sheet contains the details of the bid with each technical specification for | ITEM 57 | STETHOSCOPE PEDIATRIC
ITEM 58 | This sheet contains the details of the bid with each technical specification for | ITEM 58 | STOOL FOOT 2 STEP
ITEM 59 | This sheet contains the details of the bid with each technical specification for | ITEM 59 | STRETCHER EMERGENCY
ITEM 60 | This sheet contains the details of the bid with each technical specification for | ITEM 60 | STRETCHER RECOVERY
ITEM 61 | This sheet contains the details of the bid with each technical specification for | ITEM 61 | SUCTION UNIT HIGH VACUUM
ITEM 62 | This sheet contains the details of the bid with each technical specification for | ITEM 62 | SUCTION UNIT PORTABLE
ITEM 63 | This sheet contains the details of the bid with each technical specification for | ITEM 63 | TABLE EXAM GENERAL
ITEM 64 | This sheet contains the details of the bid with each technical specification for | ITEM 64 | THERMOMETER DIGITAL HANDHELD
ITEM 65 | This sheet contains the details of the bid with each technical specification for | ITEM 65 | THROMBOSIS PREVENTION VIEN DEEP
ITEM 66 | This sheet contains the details of the bid with each technical specification for | ITEM 66 | TIMER INTERVAL
ITEM 67 | This sheet contains the details of the bid with each technical specification for | ITEM 67 | TROLLEY WITH BASKET
ITEM 68 | This sheet contains the details of the bid with each technical specification for | ITEM 68 | ULTRASOUND UNIT PORTABLE
ITEM 69 | This sheet contains the details of the bid with each technical specification for | ITEM 69 | VENTILATOR ICU ADULT & PEDIATRIC
ITEM 70 | This sheet contains the details of the bid with each technical specification for | ITEM 70 | VENTILATOR TRANSPORT ADULT & PEDIATRIC
ITEM 71 | This sheet contains the details of the bid with each technical specification for | ITEM 71 | WARMER BABY BOTTLE
ITEM 72 | This sheet contains the details of the bid with each technical specification for | ITEM 72 | WARMER BLOOD
ITEM 73 | This sheet contains the details of the bid with each technical specification for | ITEM 73 | WASTE BIN CHEMICAL
ITEM 74 | This sheet contains the details of the bid with each technical specification for | ITEM 74 | WHEELCHAIR STANDARD SIZE 20
ITEM 75 | This sheet contains the details of the bid with each technical specification for | ITEM 75 | X-RAY DIGITAL MOBILE

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4217 0000 00600 | AED SYSTEM |  | MEDICAL GROUP | 322 | [10 empty] | 0
2 | 4115 0000 01000 | ANALYZER AUTOMATED CHEMISTRY SEMI LIQUID AGENTS TABLE TOP |  | MEDICAL GROUP | 38 | [10 empty] | 0
3 | 4115 0000 01800 | ANALYZER AUTOMATED URINE |  | MEDICAL GROUP | 38 | [10 empty] | 0
4 | 4218 1900 00600 | ANALYZER BLOOD GAS BENHCTOP |  | MEDICAL GROUP | 45 | [10 empty] | 0
5 | 4115 0000 02000 | ANALYZER CHEMISTRY POC |  | MEDICAL GROUP | 7 | [10 empty] | 0
6 | 4115 0000 02500 | ANALYZER ELECTROLYTE |  | MEDICAL GROUP | 50 | [10 empty] | 0
7 | 4115 0000 02900 | ANALYZER MARKER CARDIAC |  | MEDICAL GROUP | 44 | [10 empty] | 0
8 | 4115 0000 03100 | ANALYZER SEMI AUTOMATED COAGULATION |  | MEDICAL GROUP | 38 | [10 empty] | 0
9 | 4115 0000 03500 | APHERESIS UNIT |  | MEDICAL GROUP | 24 | [10 empty] | 0
10 | 4115 0000 03900 | BAG HOLDER BIOHAZARD |  | MEDICAL GROUP | 2450 | [10 empty] | 0
11 | 4219 1800 00500 | BED INTENSIVE CARE |  | MEDICAL GROUP | 104 | [10 empty] | 0
12 | 4219 1800 01100 | BEDHEAD UNIT HORIZONTAL |  | MEDICAL GROUP | 2500 | [10 empty] | 0
13 | 4229 3500 00800 | BLANKET WARMING SYSTEM AIR |  | MEDICAL GROUP | 50 | [10 empty] | 0
14 | 4214 0000 10300 | CABINET MEDICATION |  | MEDICAL GROUP | 110 | [10 empty] | 0
15 | 4219 1904 00000 | CABINET NARCOTIC |  | MEDICAL GROUP | 19 | [10 empty] | 0
16 | 4218 1903 00100 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE |  | MEDICAL GROUP | 720 | [10 empty] | 0
17 | 4218 1900 00900 | CART CRASH |  | MEDICAL GROUP | 170 | [10 empty] | 0
18 | 4214 0000 00800 | CART DRESSING |  | MEDICAL GROUP | 105 | [10 empty] | 0
19 | 4115 0000 07100 | CENTRIFUGE 12 TUBES TABLE TOP |  | MEDICAL GROUP | 38 | [10 empty] | 0
20 | 4219 2101 00400 | CHAIR BLOOD DONATION |  | MEDICAL GROUP | 18 | [10 empty] | 0
21 | 4111 5815 00600 | COUNTER CELL BLOOD POC |  | MEDICAL GROUP | 45 | [10 empty] | 0
22 | 4219 1607 00700 | CURTAIN SINGLE USE SLIDING |  | NON MEDICAL | 1955 | [10 empty] | 0
23 | 4217 0000 00800 | CYLINDER O2 SIZE D WITH REGULATOR |  | MEDICAL GROUP | 760 | [10 empty] | 0
24 | 4214 0000 02300 | CYLINDER O2 SIZE E WITH REGULATOR |  | MEDICAL GROUP | 372 | [10 empty] | 0
25 | 4115 0000 10700 | DECREASING SYSTEM PATHOGENS |  | MEDICAL GROUP | 11 | [10 empty] | 0
26 | 4218 1900 11900 | DEFIBRILLATOR UNIT |  | MEDICAL GROUP | 322 | [10 empty] | 0
27 | 4218 1900 01700 | DETECTOR BLOOD FLOW |  | MEDICAL GROUP | 90 | [10 empty] | 0
28 | 4218 1900 01800 | ECG 12 LEAD WITH CART |  | MEDICAL GROUP | 404 | [10 empty] | 0
29 | 4214 0000 03200 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | MEDICAL GROUP | 660 | [10 empty] | 0
30 | 4214 0000 02800 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | MEDICAL GROUP | 2450 | [10 empty] | 0
31 | 4110 3005 00000 | FREEZER ULTRA LOW TEMPERATURE -20 TO -40 C |  | MEDICAL GROUP | 38 | [10 empty] | 0
32 | 4217 0000 01200 | GLUCOMETER |  | MEDICAL GROUP | 86 | [10 empty] | 0
33 | 4229 5100 02600 | HEADLIGHT FIBEROPTIC SURGICAL |  | MEDICAL GROUP | 100 | [10 empty] | 0
34 | 4115 0000 17200 | KIT SPILL CHEMICAL |  | MEDICAL GROUP | 19 | [10 empty] | 0
35 | 4218 1900 03100 | LARYNGOSCOPE ADULT & PEDIATRIC |  | MEDICAL GROUP | 180 | [10 empty] | 0
36 | 4218 1900 03300 | LARYNGOSCOPE PEDIATRIC |  | MEDICAL GROUP | 178 | [10 empty] | 0
37 | 4214 0000 04200 | LIGHT EXAM MOBILE |  | MEDICAL GROUP | 171 | [10 empty] | 0
38 | 4115 0000 21900 | MIXER BLOOD ROLLER TYPE |  | MEDICAL GROUP | 38 | [10 empty] | 0
39 | 4218 1900 04300 | MONITOR CAPNOGRAPHY |  | MEDICAL GROUP | 195 | [10 empty] | 0
40 | 4218 2005 00300 | OPTHALMOSCOPE DIRECT |  | MEDICAL GROUP | 117 | [10 empty] | 0
41 | 4214 0000 04500 | OTOSCOPE TABLETOP |  | MEDICAL GROUP | 86 | [10 empty] | 0
42 | 4214 0000 04800 | OXIMETER PULSE HAND HELD |  | MEDICAL GROUP | 504 | [10 empty] | 0
43 | 4115 0000 24400 | PIPETTOR ADJUSTABLE |  | MEDICAL GROUP | 19 | [10 empty] | 0
44 | 4220 3700 00500 | PROTECTION X-RAY APRON VEST & SQUIRT |  | MEDICAL GROUP | 46 | [10 empty] | 0
45 | 4218 1900 06500 | PUMP INFUSION |  | MEDICAL GROUP | 920 | [10 empty] | 0
46 | 4218 1900 07000 | PUMP SYRINGE |  | MEDICAL GROUP | 260 | [10 empty] | 0
47 | 4115 0000 26500 | RACK BLOOD SEDIMENTATION |  | MEDICAL GROUP | 38 | [10 empty] | 0
48 | 4110 3011 00700 | REFRIGERATOR MEDICATION 700L |  | MEDICAL GROUP | 100 | [10 empty] | 0
49 | 4214 0000 05600 | REGULATOR SUCTION HIGH WALL MOUNT |  | MEDICAL GROUP | 2450 | [10 empty] | 0
50 | 4214 0000 06100 | SCALE CHAIR |  | MEDICAL GROUP | 56 | [10 empty] | 0
51 | 4214 0000 06000 | SCALE PATIENT WITH HEIGHT |  | MEDICAL GROUP | 56 | [10 empty] | 0
52 | 4229 5100 05700 | SCOPE SET RIGID ENT |  | MEDICAL GROUP | 215 | [10 empty] | 0
53 | 4214 0000 06300 | SPHYGMOMANOMETER ANEROID MOBILE |  | MEDICAL GROUP | 758 | [10 empty] | 0
54 | 4214 0000 06500 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | MEDICAL GROUP | 500 | [10 empty] | 0
55 | 4214 0000 06600 | STAND IV |  | MEDICAL GROUP | 2500 | [10 empty] | 0
56 | 4214 0000 06700 | STETHOSCOPE ADULT |  | MEDICAL GROUP | 250 | [10 empty] | 0
57 | 4214 0000 06800 | STETHOSCOPE PEDIATRIC |  | MEDICAL GROUP | 50 | [10 empty] | 0
58 | 4214 0000 07100 | STOOL FOOT 2 STEP |  | MEDICAL GROUP | 140 | [10 empty] | 0
59 | 4618 1810 01100 | STRETCHER EMERGENCY |  | MEDICAL GROUP | 100 | [10 empty] | 0
60 | 4219 1800 03700 | STRETCHER RECOVERY |  | MEDICAL GROUP | 70 | [10 empty] | 0
61 | 4229 5100 07000 | SUCTION UNIT HIGH VACUUM |  | MEDICAL GROUP | 2450 | [10 empty] | 0
62 | 4229 3522 00100 | SUCTION UNIT PORTABLE |  | MEDICAL GROUP | 92 | [10 empty] | 0
63 | 4214 0000 14400 | TABLE EXAM GENERAL |  | MEDICAL GROUP | 250 | [10 empty] | 0
64 | 4214 0000 07800 | THERMOMETER DIGITAL HANDHELD |  | MEDICAL GROUP | 295 | [10 empty] | 0
65 | 4218 1900 01500 | THROMBOSIS PREVENTION VIEN DEEP |  | MEDICAL GROUP | 120 | [10 empty] | 0
66 | 4115 0000 32500 | TIMER INTERVAL |  | MEDICAL GROUP | 19 | [10 empty] | 0
67 | 4214 0000 08200 | TROLLEY WITH BASKET |  | MEDICAL GROUP | 50 | [10 empty] | 0
68 | 4220 1700 00900 | ULTRASOUND UNIT PORTABLE |  | MEDICAL GROUP | 50 | [10 empty] | 0
69 | 4218 1900 08800 | VENTILATOR ICU ADULT & PEDIATRIC |  | MEDICAL GROUP | 366 | [10 empty] | 0
70 | 4217 0000 02900 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | MEDICAL GROUP | 98 | [10 empty] | 0
71 | 4219 1800 03100 | WARMER BABY BOTTLE |  | MEDICAL GROUP | 30 | [10 empty] | 0
72 | 4229 3500 02700 | WARMER BLOOD |  | MEDICAL GROUP | 30 | [10 empty] | 0
73 | 4215 2200 15700 | WASTE BIN CHEMICAL |  | MEDICAL GROUP | 2450 | [10 empty] | 0
74 | 4219 2210 06500 | WHEELCHAIR STANDARD SIZE 20 |  | MEDICAL GROUP | 815 | [10 empty] | 0
75 | 4220 1800 01300 | X-RAY DIGITAL MOBILE |  | MEDICAL GROUP | 48 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | MEDICAL GROUP | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
|  | NON MEDICAL | 0 | 0 | [5 empty] | If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4217 0000 00600 | 0 | AED SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Automatic External Defibrillator (AED)
2 | Waveform | biphasic
3 | Defibrillator Charge Hold Time | not less than 20 sec
4 | Energy Selecion | 120, 150, 200 J Adult - 50, 75, 85J Pediatric
5 | Charge Time | less than 10 sec
6 | Electrode with CPR sensor | Prefereed
7 | Automatic self test | Yes
8 | Defibrillator advisory | Yes
9 | Languages | Arabic or English
10 | Patient Impedance | 0 - 300 Ohms
11 | Display
12 | Display Format | LCD with CPR Feedback
13 | Depth Display | Yes
14 | Size | less than 4''
15 | ECG Viewing Time | specify
16 | Display sweep speed | 20 or 25 mm/sec
17 | Battery
18 | Capacity | 300 shock or 13 hours ECG monitoring
19 | Low Battery indication | not less than 30 discharge
20 | shelf life | 5 years
21 | Electrodes
22 | Adult with CPR sensor | Yes
23 | Pediatric | Yes
24 | Shelf life | ≥ 18 month
25 | General
26 | Size (HxWxD) cm | specify
27 | Weight KG | less than 4KGs
28 | Ingress Protection (IP) | water and dust proof, specify
29 | Carry Case | soft carry case
30 | OTHER SPECIFICATIONS | FDA, CE, ISO approved
Company Stamp |  |  |  | Company Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 4115 0000 01000 | 0 | ANALYZER AUTOMATED CHEMISTRY SEMI LIQUID AGENTS TABLE TOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Type | Dry Chemistry
3 | Parameters | 28 colormety + 3 electrolyts + 6 CALCULATED TESTS enzymes: ALP,AMYL,CHE,CKMB,CPK,GGT,GOT/AST,GPT/ALT,LAP,LDH,LIP GENERAL CHEMISTRY : ALB,BUN,CA,CRE,DBIL,GLU,HDL,IP,MG, NH3,TBIL,TCHO,TCO2,TG,TP,UA ELECTROLYTS: NA,K,CL IMMUNOLOGICAL TEST: CRP CALCULATED TESTS: LDL,NON-HDL,GLOB,ALB/GLOB RATIO,BUN/CRE RATIO, ANION GAP
4 | Table Top | YES , <= 47 cm (W)x 36 cm(D) x 42(H) cm
5 | Blood Sample Volume | <=100 μl of whole blood, serum or plasma
6 | Throughput | >= 128 Test /Hour
7 | Power Supply | 220 V - 60 Hz
8 | Other Specifications | FDA & CE & MDMA
9 | Supplier is an authorized representative of the manufacturer by SFDA and/or an official agent by Ministry of Trade | Yes (Attach official documents)
10 | Supplier is an authorized representative of the manufacturer by SFDA | Yes (Attach official documents)
11 | Attach MDMA Certificate | Yes (Give the MDMA Number)
1 | Application | Chemistry Analyzer Table Top
Company Stamp |  |  |  | Company Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 4115 0000 01800 | 0 | ANALYZER AUTOMATED URINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Tests shall include:
3 | Billirubin
4 | Blood
5 | Glucose
6 | Ketone body
7 | leukocyte esterase
8 | Leukocytes
9 | Nitrite
10 | pH
11 | Protein
12 | Specific gravity
13 | Urobilinogen
14 | Ascorbic acid
15 | Reflectance method shall be employed.
16 | Automatic calibration using calibration strips. If other, specify.
17 | Specify method of QC and verification
18 | Test completion time ≤ 1 minute
19 | Data storage for at least 100 tests
20 | Specify method of result viewing (screen, preferred, or LEDs) and user interfacing
21 | High-resolution printer for report printing shall be incorporated. If optional, price separately
22 | Preferably zero user maintenance
23 | Audio or visual alarms for errors shall be incorporated. List with details
24 | Laboratory information system interface and Bar code reader capability
25 | Bidder shall list all the safety precaution and features
26 | Bidder shall specify the exact dimension and weight for the offered equipment
27 | Compliance with standards & legislation:
28 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
29 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
30 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
31 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
32 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Semi automated Urine strip reader.
Company Stamp |  |  |  | Company Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 4218 1900 00600 | 0 | ANALYZER BLOOD GAS BENHCTOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | To incorporate a LCD screen and alphanumeric keyboard for user interfacing. Provide detailed specs (touch control, size, color or mono, resolution, etc.) | YES
3 | To incorporate a built-in thermal printer. Provide detailed specs (paper size and type (roll, Z-fold, etc.), resolution, printing format and information (data, tables, patient demographics, normal values, abnormal flags, etc.) | YES
4 | Measured blood gas parameters to include (Specify measurement range, units, accuracy as well as any other relevant information): | YES
5 | SAMPLE VOLUME | 200 uL APPROX
6 | ANALYSIS TIME | LESS THAN 100 SEC. APPROX
7 | ELECTRODE | LONG LIFE ELECTRODE FOR AT LEAST 1 YEAR
8 | Sodium / Potassium / Calcium / Chloride | YES
9 | PH | YES
10 | PCO2 / PO2 (total oxygen concentration in blood ) | YES
11 | LACTATE / HCT
12 | CO-Oximetry panel shall be integrated. | YES
13 | SO2 (Oxygen saturation of hemoglobin in blood ) | YES
14 | HCO 3 (Concentration of hydrogen carbonate in plasma ) | YES
15 | Aa DpO2 (Alveolo-arterial oxygen tension difference for arterial blood) | YES
16 | SBE & ABE (standard and actual base excess) | YES
17 | SBC (standard bicarbonate) | YES
18 | Plasma tCO2 (Concentration of total CO2 in plasma) | YES
19 | PT PH (pH of plasma at patient temperature) | YES
20 | PT PCO2 (CO2 tension in blood at patient temperature) | YES
21 | PT H+ (Concentration of hydrogen ions in blood at patient temperature | YES
22 | PT PO2 (O2 tension in blood at patient temperature) | YES
23 | Measured electrolytes (Specify measurement range, units, accuracy as well as any other relevant information): | YES
24 | User entered data shall include (provide details): | YES
25 | Patient demographics (name, ID, age, sex, height, weight) | YES
26 | Technician ID | YES
27 | Physician name | YES
28 | Location (outpatient, in-hospital department or bed, etc.) | YES
29 | Sample type and time taken | YES
30 | Patient temperature CORRECTION | YES
31 | Hb | YES
32 | FiO2 | YES
33 | Automatic washing capability. Specify parameters. | YES
34 | Hands free sample administration (syringe or capillary tube): | YES
35 | Fully automated aspiration, transport, analysis and processing of sample | YES
36 | State sample volume in all applicable modes (normal, micro, etc.) | YES
37 | State type, stability and accuracy of temperature control | YES
38 | Specify test throughput (analysis time) from “READY” state as well as start up time from off to ready | YES
39 | Automatic calibration. State levels, schedule (or time interval) and duration. Calibration interruption for STAT samples - LIQUID OR GAS CALIBRATION | YES
40 | On-board QC management | YES
41 | Barcode reading capability for sample and reagent handling. Automatic reagent tracking (level detection, expiry date, etc.) | YES
42 | Built-in automatic maintenance program, preferably zero user maintenance | YES
43 | State type of required gas mixture and method of delivery | YES
44 | Internal troubleshooting (self diagnostic) software capability is an asset. Specify details. | YES
45 | Computer connectivity for Bi-directional data, results and report transmission from and to the Hospital Information System including archiving capabilities | YES
46 | The following features of the offered system are preferred. State any that apply to your system. If optional, Price separately: | YES
47 | Sealed waste container | YES
48 | On-board help function | YES
49 | Remote diagnostics | YES
50 | PROVIDE ITEMIZED PRICE FOR 1,000 TESTS | SPECIFY ITEMIZED PRICE EACH TEST DEVICE CAN DO
51 | LANGUAGE | ENGLISH
52 | PLUG | BRITISH
53 | CERTIFICATION | FDA , CE OR APPROVED BY REGULATORY BODY
54 | UPS | 1 HOUR BACKUP THAT MATCH DEVICE POWER REQUIREMENT
1 | Fully automated, microprocessor controlled, table-top blood gas and electrolytes analyzer for whole blood, plasma and serum samples , PLEURAL FLUID & DIALYSATE | YES
Company Stamp |  |  |  | Company Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 4115 0000 02000 | 0 | ANALYZER CHEMISTRY POC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The point of care Hemoglobin Analyzer shall be easy‐to‐use with intuitive, screen step by step instructions .
2 | Measurement time : ≤ 15 secs Specify.
3 | Hb Range of : 4.5-25.6 g / dl Specify.
4 | It shall be battery operated : ≥2000 test operating capacity.
5 | It shall be made of very durable material
6 | It shall be practical and portable
7 | The following should be specified :
8 | Sample size
9 | Linearity
10 | Device auto shuts down :≥ 5 mins of non use
11 | Time to result
12 | Compliance with standards & legislation:
13 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
16 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
17 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Company Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 4115 0000 02500 | 0 | ANALYZER ELECTROLYTE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | To incorporate a LCD screen and alphanumeric keyboard for user interfacing. Provide detailed specs (touch control, size, color or mono, resolution, etc.)
3 | To incorporate a built-in thermal printer. Provide detailed specs.
4 | Throughput shall be more than 60 samples/hr
5 | Aanalysis time < = 60 sec
6 | Measured electrolytes (Specify measurement range, units, accuracy as well as any other relevant information):
7 | Na
8 | Serum/Plasma:40 250 or better
9 | Urine: 10-300 or better
10 | K: 50-200 or better
11 | Serum/Plasma : 0.5-15 or better
12 | Urine: 2-150 or better
13 | Cl
14 | Serum/Plasma: 50-200
15 | Urine: 15-300 or better
16 | Calcium: 0.1-6
17 | Li: 0.1-5
18 | Automatic washing capability. Specify parameters.
19 | Specify test throughput (analysis time) from “READY” state as well as start up time from off to ready
20 | Automatic calibration. State levels, schedule (or time interval) and duration. Calibration interruption for STAT samples
21 | Internal troubleshooting (self diagnostic) software capability is an asset. Specify details.
22 | RS232 interface
23 | Data management capabilities
24 | The system should incoporate the following features: Display, printer, bar code reader
25 | Shall specify the exact dimension and weight for the offered equipment
26 | Compliance with standards & legislation:
27 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
28 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
29 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
30 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
31 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Fully automated, microprocessor controlled, table-top electrolytes analyzer for whole blood, plasma, serum and urine samples
Company Stamp |  |  |  | Company Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 4115 0000 02900 | 0 | ANALYZER MARKER CARDIAC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Capacity to run ROPONIN T / TROPONIN I, MYOGLOBIN, D- DIMMER, NT proBNP / BNP, CK-MB. Specify Range
3 | Sample type: whole blood (capillary, syringe or vacutainer)
4 | Sample volume: 100µL or less
5 | Reportable results: Quantitative measurements
6 | Test time: 2-10minutes
7 | Shall have a Bar code reader system
8 | Automated Internal quality control system
9 | Weight up to 1 kg
10 | Data storage capacity: Equal or greater than 1000 patient results
11 | Ability to interface with wireless technology and LIMS for data storage
12 | Reagent for three months supply for (25 -30) samples per day should be supplied at the time of commissioning
13 | Compliance with standards & legislation:
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Battery/ electricity operated with complete accessories for use
Company Stamp |  |  |  | Company Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 4115 0000 03100 | 0 | ANALYZER SEMI AUTOMATED COAGULATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Two (minimum) channel measurement system
3 | High resolution LCD display of test parameters, calibration data and results. Specify characteristics.
4 | Built-in or attached printer for patient results, QC and calibration curves, etc.
5 | QC processing for normal and pathologic levels (N and P)
6 | Capability to be configured to run tests in single or duplicate runs
7 | Automatic test start with reagent pipetting
8 | Specify the number of reagent wells and on-board cooling method and stability for each assay
9 | Pre-defined (factory preset) test protocols with possibility of user modification and addition of new tests
10 | Testing capability should include:
11 | PT
12 | APTT
13 | Fibrinogen
14 | Thrombin time
15 | Optional testing capability should be offered if available:
16 | Factor assays
17 | Reptilase
18 | Lupus Anticoagulant
19 | PS
20 | PC
21 | Heparin assays
22 | Reagent life (room temperature and refrigerated) after reconstitution and number of tests per reagent pack for each assay shall be clearly stated
23 | Storage reagent requirements (before reconstitution) and life shall be stated
24 | Must be user-friendly with minimal user maintenance required
25 | Specify whether reagent is readily available for immediate use. If reconstitution is required, please state method and required time
26 | System shall be equipped with state-of-the-art advanced technological features providing highest levels of accuracy, linearity and repeatability. Data shall be provided with the offer
27 | Provide detailed information regarding sample size (normal, reduced and micro modes)
28 | Shall specify the exact dimension and weight for the offered equipment
29 | Compliance with standards & legislation:
30 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
31 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
32 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
33 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
34 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | The unit should be microprocessor controlled for semi-automated hemostasis testing and calibration
Company Stamp |  |  |  | Company Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | 4115 0000 03500 | 0 | APHERESIS UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Fully automated apheresis unit for separating, collecting and re-infusion of blood components for donating or therapeutic purposes
2 | The unit should be microprocessor controlled with pre-programmed protocols. Specify.
3 | Automatic parameter optimization based on donor profile and customer needs
4 | Interface to the Laboratory Information System
5 | Recognize ISBT128 barcodes and the barcodes generated by the LIS
6 | Digital alphanumeric display and keyboard for:
7 | Entry of patient demographic information
8 | Specifying type of procedure to be performed
9 | Provide users with set-up information and apheresis set installation/removal guidance
10 | Monitoring of session progress and related parameters list.
11 | Troubleshooting and system performance monitoring and verification
12 | Display of alarming conditions with user instructions
13 | Variable anticoagulant control (~ 1:8 – 1:16)
14 | Microprocessor controlled pumps for inlet, collect, anticoagulant and plasma lines with fail safe mechanisms. Provide details.
15 | Single and dual needle operation. Specify collection parameters for both modes (flow type, inlet rate, extra-corporeal volume, etc.)
16 | Component separation by centrifugation. Specify centrifugal speed range.
17 | The unit should be designed for all apheresis applications:
18 | Collection
19 | Exchange
20 | Therapeutic removal
21 | Shall list all the safety precaution and features
22 | Compliance with standards & legislation:
23 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
24 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
25 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
26 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
27 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Company Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | 4115 0000 03900 | 0 | BAG HOLDER BIOHAZARD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Bright biohazard epoxy-coated steel wire
3 | Chemical and corrosion-resistant
4 | Bottom tray
5 | Shall include 20 biohazard bags
6 | Bags shall withstand a temperature range of -20°C to 120°C
7 | Autoclavable
8 | Approx. Dimension: 40 D x 75H cm
1 | Sturdy wire holder keeps bags open while a separate hooked ring that rests over the top keeps them securely in place
Company Stamp |  |  |  | Company Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | 4219 1800 00500 | 0 | BED INTENSIVE CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Full electric ICU Patient Bed
2 | PATIENT CONTROLS | Back lighted
3 | Location | Rail embedded
4 | Functions/Indicators/Alarms | Nurse call, head up/down, knee up/down, room light, read light, night light and enhanced entertainment control; Specify
5 | NURSE CONTROLS
6 | Patient control lockouts | Yes
7 | Full low bed indicator | Yes; Specify the height
8 | CPR control | Yes
9 | Trendelenburg head section indicator | Yes
10 | Optional Hands free Foot Control | Preferred
11 | Function Control Display/Screen | LCD Touch screen
12 | Displays | Automatic contour, etc/ Specify
13 | TECHNICAL SPECIFICATIONS
14 | CONTOUR | Automatic by touch of a button
15 | RETRACTABLE AND EXTENDABLE | Yes
16 | PATIENT WEIGHT, KGS | ≥ 225 kgs
17 | SIDERAIL
18 | Percentage of overall length | Specify
19 | Prevents being tucked or entrapped in the side rails and senses objects between the bed frames. | Yes
20 | CASTERS | 5 preferred
21 | Diameter, cm | Specify
22 | Functions | Brake, swivel lock, steer
23 | TRANSPORT HANDLES | Preferred
24 | CENTRAL BRAKE SYSTEM | Yes
25 | CORNER BUMPERS | Yes
26 | HEADBOARD & FOOTBOARD | Removable
27 | IV POLE
28 | Storage | Yes
29 | Mounts | 4 or more
30 | BATTERY BACK UP | Yes
31 | MANUAL CRANK FUNCTION | Yes
32 | SPEAKERS | Yes
33 | HOLDERS/SOCKETS
34 | Drainage bag holders | Yes, specify nos
35 | Oxygen Cyllinder Holder | Yes
36 | Equipments Sockets | Yes, specify nos
37 | Accessory electrical outlets | Yes, specify nos
38 | CONNECTIVITY
39 | USB connectivity (for smart phones & tablets) | Preferred
40 | Connectivity with the Nurse Call System and Hospital Data integration management System | Yes
41 | C-ARM COMPATIBLE | Yes
42 | INTEGRATED MATTRESS | Powered and programmable
43 | Pressure-reducing mattress | Yes; Specify
44 | Surface Risk Level | medium
45 | Mattress Base Cover | Waterproof PVC/waterproof Polyurethane coated ticking
46 | Maximum Patient Load | Specify
47 | Dimensions | Specify
48 | Radiotransluscent surface | Yes
49 | Integrated X-ray Cassette | Yes
50 | Anti-microbial | Yes
51 | Heat and moisture free | Yes
52 | SPECIAL FEATURES
53 | Cardiac Chair positioning | Yes
54 | Head Of Board 30° position indicator | Yes
55 | Integrared Patient Scale | Yes; Specify capacity
56 | Bed exit alarm | Yes
57 | Orthostatic hypotension prevention | Yes; Specify
58 | Early Mobility therapy | Yes; Specify
59 | Other integrated therapy | Specify
60 | Turn Assist- Mattress will inflate from one side and deflate from the other side for turning the patient easily in bed | Specify
61 | ELECTRICAL FEATURES
62 | Number of Motors | 6, or Specify
63 | Electrical Ground Connection | for bed frame
64 | DC device electrically isolated | Yes
65 | PHYSICAL SPECIFICATIONS
66 | Dimensions | Specify
67 | Weight | Specify
68 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Company Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | 4219 1800 01100 | 0 | BEDHEAD UNIT HORIZONTAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | System Description: The system shall have the following features:
3 | It shall be constructed of extruded anodized aluminum alloy sections.
4 | Anodized aluminum facials shall be removable to provide access to individual components mounted within the unit.
5 | The unit shall be contoured and ergonomically designed.
6 | Front panels shall be finished with high-pressure plastic laminate face. Color of the laminate face shall match the color of the patient bed, bedside cabinet and over bed table.
7 | The Horizontal Bedhead Unit shall include a Bed Locator.
8 | The Horizontal Bedhead Unit and Bed Locator shall comprise the services outlined on the corresponding detailed loaded drawings
9 | The outlets shall be Color coded. Coding to be coordinated with the Hospital’s electro-mechanical consultant or contractor.
10 | The outlets shall be ergonomically spaced so that they can comfortably carry items such as oxygen regulators with humidifiers and suction regulators with suction jars without colliding or obstructing other service outlets. Connection of all outlets and co
11 | The bedhead unit shall include accessory tracks for mounting various accessories.
12 | The bedhead unit installation and design shall meet the international quality and safety standards for medical gases, pipeline system and bedhead units adopted by the hospital.
13 | Approximate overall dimensions (D x H): 10 x 35 cm. For approximate Width dimension, refer to the corresponding Loaded Drawings.
14 | A wall mounted rail shall be installed under each bedhead unit. It shall be sturdy enough to carry several items such as baskets and physiological monitors.
15 | The bedhead unit supplier shall submit detailed shop drawings showing items in plan and elevation views with the color schemes matching the patient’s furniture, for the Client / Engineer approval prior to shipment and delivery of bedhead unit items to sit
16 | Shop drawings shall show dimensions and details of related services including wiring and piping diagrams. They shall also show a complete list of components, attachments and accessories for approval.
17 | The bedhead unit installation and design shall meet the standards and norms adopted by the electro-mechanical consultant (NFPA, HTM, etc.).
18 | Suppliers shall install medical gas outlet terminals and electrical socket outlets using the same type already installed in the existing hospital.
19 | Outlet provisions and outlet installations for patient monitors, data, telephone and nurse call systems shall be fully coordinated with the electromechanical consultant.
20 | The bedhead unit supplier shall ensure compatibility of the delivered bedhead units with patient beds for the purposes of connectability of patient overbed lighting system, nurse call system, telephone and TV system to bed siderail controls, wherever appl
21 | The bedhead unit shall be pre-piped and pre-wired; factory assembled electrical and mechanical components. It shall also be factory tested and delivered to site with singular terminal connections.
22 | Compliance with standards & legislation:
23 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
24 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
25 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
26 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
27 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
28 | Special Site Preparation Requirements:
29 | Bidders shall coordinate with the civil and electromechanical contractors to provide complete site preparation requirements.
30 | Bidders shall provide complete shop drawings.
1 | The supplier shall supply and install a horizontally wall-mounted patient bedhead service trunking, complete with provisions for power supply, nurse call, medical gases and other facilities.
Company Stamp |  |  |  | Company Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | 4229 3500 00800 | 0 | BLANKET WARMING SYSTEM AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Inflation time | 90 sec
2 | Warming up time | 4 minutes
3 | Length of Air Hoses | 1500 mm
4 | Temperature | 20°C - 42°C, step of 1°C
5 | Pressure | 0 – 60 mb. Step of 5 mb
6 | Alarm | Audio and visual
7 | Hand Controller | Specify
8 | Warm-Air Mattress, Wide | Length 200 cm, Width 60 cm.
9 | Warm-Air Mattress, Small | Length 100 cm, Width 50 cm
10 | Strap | Yes
11 | Hose seal | 2 pcs
12 | Hose set complete | 2 pcs
13 | Itemized price for all accessories | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | 4214 0000 10300 | 0 | CABINET MEDICATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | يحتوي على 4 أدراج بحد أدنى. | نعم
1 | مدهون بدهان حراري غير قابل للصدأ | نعم
1 | يحتوي على قفل. | نعم
1 | شكل عصري وعملي | نعم
1 | وجود فهرسة | نعم
1 | تقديم ضمان شامل لا يقل عن 5 سنوات | نعم
1 | الابواب زجاجية | نعم
1 | الابعاد | طول 100 سم, عرض 50 سم ,ارتفاع 180 سم
1 | يستخدم لحفظ الادوية. | نعم
Company Stamp |  |  |  | Company Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | 4219 1904 00000 | 0 | CABINET NARCOTIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FINISH | Enamelled
1 | NUMBER OF DOORS
1 | door type | piano hinge
1 | locks | Yes separate for each door
1 | DIMENSION | 350mm x 250mm x 200mm H. approximately
1 | OTHER SPECIFICATION
1 | ALARMS
1 | visual | Yes
1 | audible | Yes
1 | POWER
1 | vac
1 | hz
1 | APPLICATION | Storage of Narcotics
Company Stamp |  |  |  | Company Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | 4218 1903 00100 | 0 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | THE SYSTEM MUST BE ABLE TO CONTINUOUSLY MEASURE PARAMETERS | YES
3 | SV/SI (STROKE VOLUME / INDEX) | YES
4 | CO/CI (CARDIAC OUTPUT / INDEX) | YES
5 | HR (HEART RATE) | YES
6 | CARDIAC OUTPUT TECHNOLOGY & ALGORITHM | YES
7 | SHOULD HAVE MULTIPLE SCREENS DISPLAYING VARIOUS PARAMETERS AND THE PHYSIORELATION SCREEN | YES
8 | SHOULD DISPLAY THE PERIOPERATIVE GOAL DIRECTED THERAPY SCREEN, | YES
9 | CONTINUOUSLY MEASUREMENT OF HAEMDYNAMIC PARAMETERS | YES
10 | SYSTEM SHOULD AUTO-CALIBRATE USING PATIENT DEMOGRAPHICS | YES
11 | SVV (STROKE VOLUME VARIATION) – FLUID RESPONSIVENESS INDICATOR | YES
12 | CALCULATE SVR (SYSTEMIC VASCULAR RESISTANCE / INDEX) | YES
13 | CALCULATE DO2/DO2I (DELIVERED OXYGEN) | YES
14 | THE SYSTEM SHOULD BE ABLE TO DISPLAY AND PERFORM FLUID RESPONSIVENESS TEST (SV AND PLR) WHEN OPERATED | YES
15 | THE SYSTEM SHOULD BE USABLE FOR ALL DEPARTMENTS - ICU, OR , ER AND WARDS | YES
16 | THE TECHNOLOGY USED SHOULD HAVE OPTION OF TOTALLY NON-INVASIVE TO FACILITATE THE USE IN EMERGENCY SITUATIONS | YES
17 | EACH PARAMETER SHOULD BE DISPLAYED WITH IT'S NORMAL RANGE AND THE RANGE SHOULD BE ADJUSTABLE IF REQUIRED | YES
18 | THE MONITOR SHOULD BE COMPACT WITH EASY SETUP. | YES
19 | THE MONITOR SHOULD BE CONNECTABLE TO A SOFTWARE PROVIDES PARAMETER TRENDING, PASSIVE LEG RAISING TEST, PRINTING, ...ETC | YES
20 | THE MONITOR SHOULD BE TRANSPORTABLE AND EASY TO USE ON STAND OR TROLLEY | YES
21 | THE MONITOR SHOULD HAVE A BUILT-IN BATTERY AND STORAGE MEMORY FOR 72 HOURS PER PATIENT | YES
22 | BIDDER SHALL LIST ALL THE AVAILABLE ACCESSORIES | YES
23 | COMPLIANCE WITH STANDARDS & LEGISLATION: | YES
24 | SHOULD HAVE A FDA APPROVAL AND/OR CE MARK & SFDA REGISTRATION, WHERE APPLICABLE. LIST ANY OTHER INTERNATIONAL STANDARDS (CE, UL, TUV, CSA), IF ANY. | YES
25 | ALL ELECTRICAL CONNECTIONS AND PLUGS SHOULD BE HOSPITAL GRADE AND FOLLOW INTERNATIONAL, LOCAL AND HOSPITAL REQUIREMENTS. | YES
26 | PROVIDE HARD/SOFT COPIES OF THE OPERATION AND MAINTENANCE MANUALS AS PER THE TENDER TERMS AND CONDITIONS | YES
27 | ALL OTHER BASIC ACCESSORIES DEEMED NECESSARY THAT ARE NOT MENTIONED IN THIS SPECIFICATION BUT ARE REQUIRED FOR FULL FUNCTION AND HIGHEST CLINICAL OUTCOMES AND OUTPUT OF THE EQUIPMENT MUST BE INCLUDED. | YES
28 | SHOULD BE ACCURATE AND USER INDEPENDEDNT TO GET ALL PARAMETERS AND REDUCE THE NEED OF CLOSE CONTACT WITH PATIENT | YES
29 | SCREEN SOULD BE 10.4 INCHES APPROX COLOUR LCD TOUCH SCREEN MONITOR, WINDOWS OS | YES
30 | AUDIBLE ALARM WHEN PARAMETERS GO OUT OF RANGE AND THE ALARM SHOULD BE ABLE TO BE MANUALLY SILENCED | YES
31 | VENDOR MUST PROVIDE FREE DEVICE WITH SERTAIN NUMBERS OF SENSORS | YES
32 | 220 VOLTS / 60 HZ | YES
33 | CART MUST BE PROVIDED WITH BASKET | YES
34 | PRINTER | OPTIONAL
1 | NONINVASIVE HEMODYNAMIC MONITORING SYSTEM | YES
Company Stamp |  |  |  | Company Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | 4218 1900 00900 | 0 | CART CRASH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Cardio-respiratory resuscitation system | (1) Top Cavity Tray, (3) Locking Side Bins, (1) Adjustable Defibrillator Tray, (1) Suction Pump Shelf, (1) Backboard with Front Assembly Kit, (1) Oxygen Tank Holder, (1) I.V. Pole with Cart Mount, (1) Storage/Gel Bin, (1) Pack Plastic Security Seals = 100 pieces, All drawers have Code Blue drawer pulls.Two brake casters positioned in the front of the cart to provide stability.
2 | Accessories for Cart
3 | Oxygen supply system | Yes
4 | No.of cylinder | 2 with press. Reducing valve
5 | Oxygen flow meter | 0 - 15 L/min
6 | Suction Unit | Prefered Laerdal / Specify
7 | Built-in AC/DC charger | Yes
8 | Large knob controls vacuum regulation & unit power | Yes
9 | Flows, LPM | 30 LPM
10 | Vacuum power | 500+ mmHg
11 | Disposable canister and patient tubing | Yes
12 | Medical grade power extention | Built-in, at the back
13 | Laryngoscope (for pedia and adult) with set Blades ( Miller sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
14 | Laryngoscope (for pedia and adult) with set Blades ( Mickintosh sizes 0, 00, 1,2,3,4)(with batteries) | Prefered Heine / Specify
15 | Guedel airway set (size 0-4) | Yes
16 | Oxygen therapy mask set | Yes
17 | Resuscitation bag (for adult + pediatric) | Yes
18 | Mask | Yes
19 | Suction handle | Yes
20 | Tubing | Yes
21 | Tourniquet | Yes
22 | Fasteners (velcro) | Yes
23 | Forceps - Magil / Artery | Yes
24 | Scissors | Yes
25 | Dressing | Yes
26 | Suture | Yes
27 | Penlight | Yes
28 | Bite stick | Yes
29 | Nasal canula set | Yes
30 | Stop watch | Yes
31 | Aneroid Sphygmomanometer with all sizes cuffs (Adult & Pediatric) | Yes, aneroid
32 | Infusion pole | Yes
33 | ECG patient lead set | Yes
34 | ECG Electrodes | Yes
35 | ECG Chart recorder paper | 5 roll /each unit
36 | Additional , 1. NIBP hose 2.Pulse oxymetry cable 3.Pacing electrodes+cable | Yes
37 | Disposables | Yes
38 | Defibrillator tray | Yes
39 | Side tray for holding a portable suction machine | Yes
40 | Adult and pediatric Defibrillation pads. | Yes
41 | Arterial clamp | Yes
42 | Stethoscope | Yes
43 | Medication tray-3" (76 mm) Drawer tray approx.20" x 15" ( 508 x 381 mm) with 4 dividers, 2ea. Front to rear rails, sheet of blank labels, 5 security bags. | Yes
44 | PHYSICAL SPECIFICATIONS
45 | Dimensions | Specify
46 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | 4214 0000 00800 | 0 | CART DRESSING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Design | Heavy duty, compact design & high quality finishing
3 | CONSTRUCTION | Rectangular tube frame, Made from S/S or Equiv. that is corrosion resistance & easy to clean. Rugged construction to with stand mechanical shocks caused by movement over uneven surfaces
4 | COMPARTMENT | Please Specify .
5 | WORKING TRAY | pull-out & made from S/S
6 | DRAWERS | 5 utility drawers with partitions to store dressing instruments, gauze etc?.
7 | CUPBOARD | One included
8 | ADJUSTABLE SHELF | Please Specify
9 | WASTE RECEPTACLE | Removable with locking mechanism
10 | TRAY | Swivel mounted & Made from S/S or Equiv
11 | BUCKET WITH LID | Available within fixture & made from S/S or equiv.
12 | AMPOULE OPENER | Included
13 | CASTORS | 4-antistatic castors with 2 efficient braking mechanism
14 | Table Top surface | Not less than 900 mm x 600 mm approx. Made from S/S or Equiv.
15 | Table top height | Not less than 1500 mm approx.
16 | BUMPERS | Revolving Bumpers
17 | Push Handles | Full-width hand rail on both sides
18 | Guard rail | 3-sided made from S/S or equiv.
19 | Accessories & Options | Any available options shall be quoted separately
20 | Catalogues | Original/colored detailed catalogue stating all the above mentioned specifications should be submitted with offer
21 | OTHER SPECIFICATION
1 | Safety standard | FDA approval or CE Marking
Company Stamp |  |  |  | Company Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | 4115 0000 07100 | 0 | CENTRIFUGE 12 TUBES TABLE TOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Provides Thin Layer Cell Preparation from Body Fluids & CSF
2 | FEATURES
3 | Process 12 Specimens at one time | Yes
4 | Lid Interlock | Yes
5 | Window View during operation | Yes
6 | Safety of Mechanical & Electronic components from Spillage | Yes
7 | Easy Disinfection | Yes
8 | Protected Program Memory | Specify
9 | Tilting Feature in Ideal Position | Yes
10 | OPERATING FEATURES
11 | Operates only when lid is locked | Yes
12 | Speed Range | 200 - 2000 RPM
13 | Timer Range | 1 - 99 min
14 | Acceleration Control keys to protect fragile samples | Specify
15 | Keypad (Touch) | Yes
16 | ALARMS
17 | Alerts for Lid not Locked | Yes
18 | Alerts for Imbalance | Yes
19 | Alerts for outside speed tolerance | Yes
20 | Timed out Alarms to remove Specimen | Yes
21 | OTHER SPECIFICATION
22 | Noise Level | Specify
23 | Heat Dissipation | Specify
24 | Cytospin Includes Starter Kits & Sealed Head/Rotor | Yes
25 | Meets IEC 61010 and UL Safety Standards | Yes
26 | Integrate with Current & Future KFMC HIS System | Yes
27 | Additional Distinguish Features if any | Specify
28 | PHYSICAL SPECIFICATIONS
29 | Dimensions | Specify
30 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | 4219 2101 00400 | 0 | CHAIR BLOOD DONATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit must be securely balanced in all operational positions.
2 | The unit must be constructed with easily sanitized materials
3 | Suitable for right or left hand use.
4 | The blood-drawing chair shall have two skin foam cradle armrests located at both sides.
5 | Arm rest adjusts to any size/height patient.
6 | The armrest should lock into place at the desired location.
7 | Arm extensions must provide visibility of the elbow region, be large enough to securely support venipuncture, and be easily retracted in case of patient collapse
8 | Adjustable leg cushion with self blocking mechanism
9 | The chair should allow for rapid and easy adjustment from sitting to reclining, supine and Trendelenburg position - in the event of the patient fainting or going into shock
10 | The unit shall have power adjustable (electric motors) movements to allow safe, comfortable and precise patient positioning at the press of a button on the hand control unit.
11 | Capability to carry up to 200 Kg
12 | The unit casters should allow easy repositioning and lock to secure safety in operation
13 | The unit shall have non slip rubber feet
14 | Color shall be coordinated with ID concept / Engineer.
15 | All surfaces should be easily cleaned and sanitized, to have:
16 | Powder-coated frame for durable, high quality, finish
17 | Premium medical grade treated vinyl upholstery
18 | Replaceable clear protective cover over foot area
19 | Upholstery should possess superior bacterial protection making it ideal for healthcare applications.
20 | Compliance with standards & legislation:
21 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
22 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
23 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
24 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
25 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Company Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | 4111 5815 00600 | 0 | COUNTER CELL BLOOD POC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | METHOD | Volumetric impedance/ Flow Cytometry
3 | PARAMETERS | 22 hematology parameters, 3 separate histograms: WBC, LYM, MID, GRA, LYM%, MID%, GRA%, RBC, HCT, MCV, RDWc, RDWs, HGB, MCH, MCHC, PLT, PCT, MPV, PDWc, PDWs, P-LCC, P-LCR%
4 | Reagent Stability after open | 60 days
5 | Sample type | Venous or capillary, Open & Closed Tube
6 | Point of Care | Yes, portable
7 | Reagent Pack | >=200 Test Reagent &Waste Bottle
8 | SAMPLE VOLUME | ~20 μl of whole blood
9 | CYCLE TIME, sec | <= 1 minute
10 | SAMPLES/HOUR | 60 tests/hour
11 | VISUAL OUTPUT | LCD backlight, graphic display Display, 10.1˝, TFT Touch Screen, 1280 x 800 resolution
12 | PATIENT ID SYSTEM | Yes
13 | COMPUTER INTERFACE | 1 x LAN 4 x rear USB ports and 2 x side USB ports WiFi/4G dongle support
14 | USB Barcode Reader | Sample ID, Control ,blood data upload
15 | Biohazard-free on-board integrated reagent pack and waste | Reagent and Waste in One Bottle ,
16 | DATA STORAGE | >= 10,000 results including histograms
17 | QC PROGRAMS | QC functions: The trend function allows to display the selected records from the database on Levey-Jennings charts Extended QC mode: Secondary QC limits for all parameters
18 | ALERT INDICATORS | Yes
19 | CALIBRATION | Auto calibration
20 | VOLTAGE | 220V, max. 15A,
21 | FREQUENCY | 60Hz
22 | Battery Operated | Yes , Integrated
23 | REGULATORY COMPLIANCES | ISO, CE, FDA, (Specify in offer)
24 | SAFETY STABDARD | UL, IEC, CSA or other International Safety Standards (Specify in offer)
1 | FDA/CE APPROVED | Yes [submit copy with offer]
Company Stamp |  |  |  | Company Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | 4219 1607 00700 | 0 | CURTAIN SINGLE USE SLIDING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
3 | Width | 3 Meters Approx.
4 | Length | 2.5 Meter Approx.
5 | Colour | Specify
6 | Including Track | Yes
7 | Hanging system | Specify
8 | OTHER SPECIFICATION | Specify
1 | Type | Standard Curtains
2 | Disposable | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | 4217 0000 00800 | 0 | CYLINDER O2 SIZE D WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
3 | Regulator | Yes
4 | Humidifier & mask | Yes
5 | Connections | Yes
6 | OTHER SPECIFICATION
1 | TYPE | CYLINDER O2
2 | SIZE | D
Company Stamp |  |  |  | Company Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | 4214 0000 02300 | 0 | CYLINDER O2 SIZE E WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
3 | Nominal contents (litres) | 680L
4 | PIN INDEX | YES
5 | CAPACITY | SPECIFY
6 | ALUMINUM TROLLEY | INCLUDED
7 | OXYGEN REGULATOR | INCLUDED WITH FLOW METER 15 L/Min
8 | OTHER SPECIFICATION | SPECIFY
1 | OXYGEN CYLINDER E-SIZE | YES
2 | ALUMINUM CYLINDER | YES
Company Stamp |  |  |  | Company Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | 4115 0000 10700 | 0 | DECREASING SYSTEM PATHOGENS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The system should help improve the safety of the blood supply by reducing the infectious levels of disease-causing agents in platelets and plasma while still maintaining quality blood components.
2 | The device must use riboflavin (vitamin B2), a naturally occurring, non-toxic compound, combined with ultraviolet light, to provide effective pathogen reduction and white cell inactivation.
3 | It shall use a single platform capable of treating platelets and plasma using the same compound and light delivery system.
4 | The system shall demonstrate:
5 | An extra measure of safety in the face of emerging infectious disease.
6 | Reduced risk from screened and unscreened pathogens and white blood cells
7 | An effective alternative to bacterial screening of platelets, with no delay in product release
8 | Flexibility and ease of use for optimal operational efficiency
9 | Accurate reporting and traceability through fully integrated data capture
10 | The system shall consist of three main components:
11 | A disposable kit including an illumination/storage bag and sterile riboflavin solution
12 | An Illuminator to provide UV light and agitation for the process
13 | A software to integrate and manage data reporting and storage
14 | Treating platelets and plasma shall occur in three steps:
15 | The product is to be transferred to the illumination/storage bag
16 | Riboflavin solution is to be added and mixed with the product
17 | The mixture is then to be exposed to UV light for about five to ten minutes
18 | The device shall be easy to operate, flexible, and need limited handling and minimal product loss with the process help keep your operations efficient.
19 | The device shall have a fully integrated data capture and storage software system ensures accurate reporting and simplified data management.
20 | Connectivity with LIS , and HIS
21 | Compliance with standards & legislation:
22 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
23 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
24 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
25 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
26 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Company Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | 4218 1900 11900 | 0 | DEFIBRILLATOR UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Automatic/Semi Automatic
2 | DEFIBRILLATOR FEATURES
3 | Manual override | Yes
4 | Voice prompting | Yes
5 | Number of Wave forms | More than 8
6 | Unit with adult & pediatric paddles | Yes
7 | Energy sequence, AED mode, J | as per AHA guidelines
8 | Energy settings, manual mode, J | Step up 2-200 (or higher)
9 | Protocol configured | Specified
10 | Output waveshape | Truncated exponential or rectilinear
11 | Automated External Defibrillation Option | Yes
12 | Shock advisory. | Yes
13 | Synchronized cardioversion. | Yes
14 | Non-Invasive blood pressure monitoring | Yes
15 | Non-invasive pacing | Yes
16 | Side stream Capnography. | Yes
17 | Temperature monitor. | Yes
18 | 12-lead ECG monitoring. | Yes
19 | SpO2 monitoring. | Yes
20 | Using CPR metronome Built-in | Yes
21 | Audible tone guide without distracting vocal critique | Yes
22 | Sunview button | Yes
23 | Data display shall be arranged in such way that ease the reading and quick identification of abnormal reading. | Yes
24 | Trending and storage features for all parameters in both Defib/Monitormode. | Yes
25 | The unit shall comply with the most recent AHA guidelines. | Yes
26 | Discharge waveform shall provide the most effective energy delivery with the lowest possible damage; vendor shall submit all documents supporting their claim. | Yes
27 | Unit shall be provided with versatile arrhythmia detection algorithm.Vendor shall specify how many leads are being analyzed simultaneously. | Yes
28 | System shall be equipped with advisory messages and/or a feedback technique that provide the caregiver with information about the number and depth of compressions during CPR. | Yes
29 | The unit shall have full –energy Biphasic Manual Defibrillators. | Yes
30 | The equipment shall be easily switched from AED to manual mode. | Yes
31 | The unit shall be light weighted and have ergonomic design. | Yes
32 | The unit shall be easy to operate with clearly labeled big buttons for Analyze, Charge and Discharge. | Yes
33 | The unit shall be capable to recharge to the maximum energy in less than 5 seconds. | Yes
34 | The unit should incorporate hard adults paddles with integrated pediatrics paddles. | Yes
35 | Parameters Monitored | ECG, Pacer
36 | Pacemaker | Yes
37 | MONITOR
38 | The unit shall be provided with the following hardware features:
39 | Large ≥ 8” high-contrast TFT color display with resolution ≥ 800 x 600 | Yes
40 | Thermal array printer | Yes
41 | AC adaptor | Yes
42 | Patient cables and probes/sensors and all necessary accessories. | Yes
43 | Status indicator that should visually warn the users about the readiness of the system even if the battery is fully depleted and AC power is off. | Yes
44 | The unit shall be provided at least with the following alerting/alarming features
45 | The system shall provide an auditable signal for QRS | Yes
46 | Audiovisual alarms for any parameter(s) out of range | Yes
47 | Accurate Battery gauge and near battery depletion | Yes
48 | Indication of AD/DC power | Yes
49 | Arrhythmia detection | Yes
50 | System shall perform a daily automatic self test or on-demand and shall report (print) the readiness of the system, and log the result of testing in an internal memory. | Yes
51 | All hardware and software problems shall be logged with date and time stamp; the system shall be capable of recording at least 3000 events. Accessing and deleting events shall be authenticated. | Yes
52 | Unit shall recognize the type of the pads and limit the energy accordingly. (Pediatrics shall have a limited energy level that is different from that of adult. | Yes
53 | The unit should provide defibrillation energy selection, and demand/fixed rate pacing, and selection of ECG leads monitoring. | Yes
54 | The unit paddles should provide Charge/Discharge controls. | Yes
55 | The unit should provide AC/DC operation, with Li-Ion/Li-Pol batteries. | Yes
56 | The batteries should be fully charged 100% within < 3 hours. When fully charged, the unit should provide > 200 discharges or 6 hours of continuous monitoring with full parameters. | Yes
57 | The unit should be suitable for crash cart purposes. | Yes
58 | The printer should provide summary report and code markers. | Yes
59 | All external components should be securely mounted. | Yes
60 | The unit should be secure and provide adequate protection against moving and electrically energized parts. | Yes
61 | Switches, knobs, and other controls should be designed for conditions of heavy use. | Yes
62 | The controls (e.g., switches, knobs) should be visible and clearly identified, and their functions should be self-evident. Device design should prevent misinterpretation of displays and control settings. | Yes
63 | Controls should be sealed against penetration of liquids. | Yes
64 | The unit should be easy to clean, disinfect, and/or sterilize. | Yes
65 | Vendor shall provide itemized price for all option, disposable come with machine. | Yes
66 | Monitor ECG acquisition | Dispo. Defeb. Elect/ multifunctunal/quick patch
67 | ECG display Type / Message display Type | LCD/LCD
68 | Heart rate display Gain | Yes /autogain/manual
69 | Spo2 display with probe | Yes, Nellcor, finger, neonatal
70 | Disposable /Reusable | Yes
71 | Electrodes | Solid gel
72 | Conductive area | 82 cm2 approx.
73 | Shelf life | Not less than 1 8 months
74 | DOCUMENTATION /DATA STORAGE/ANALYSIS
75 | Chart recorder > 100mm | Yes
76 | ECG/voice recording | Standard/Optional
77 | Capacity, min | 60/90
78 | Playback | Yes
79 | Solid-state memory | Yes
80 | Information stored | Analysis/events
81 | Capacity, min | Specify
82 | Playback | Yes
83 | Event record database storage | Yes
84 | Data management software | Yes
85 | Analysis
86 | - Auto or manual | Auto/Manual/Both
87 | - Segment analyzed, sec | <3
88 | - Analysis time, sec | 2020-06-12
89 | - V tach rate threshold, bpm | >120
90 | - VF amplitude threshold, mV | >0.08
91 | BATTERY | QTY : 2 sets
92 | Integrated battery charger | Yes
93 | External battery charger | Yes
94 | Type | rechargable
95 | Power supply | Yes
96 | Integral/replaceble | Integral, with built in charger
97 | Charging method | Line cord/external/plugin
98 | Charge time, hr | < 3 Hrs
99 | PHYSICAL SPECIFICATIONS
100 | Dimensions | Specify
101 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | 4218 1900 01700 | 0 | DETECTOR BLOOD FLOW |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DOPPLER WITH 5 MHz PROBE | 1
2 | DOPPLER WITH 8 MHz PROBE | 1
3 | TYPE | Portable / Handheld
4 | APPLICATION | High level vascular assessment, ideal for the detection of peripheral arterial disease (probe sizes neonatal, infant, child, adult)
5 | PROBE FREQUENCY (INTERCHANGEABLE) | 5, 8, 10 MHz ( Probe 10 Mhz compatible with the device)
6 | PROBE CONSTRUCTION | Durable Long Lasting Sealed Probes
7 | PROBE SHAPE / LENGTH | Pencil / 109 mm (approx)
8 | BUILT-IN SPEAKER (200 mW) | High fidelity speaker
9 | POWER CONSUMPTION : | 500 mW or less
10 | BI-DIRECTIONAL LCD DISPLAY | Yes
11 | WAVEFORM CALIBRATION FUNCTION | Yes
12 | GAIN CONTROL | Yes
13 | OBSTETRIC CAPABILITY | Yes/ Neonatal
14 | BATTERY
15 | - Type of battery | Specify
16 | - Operating time, hrs | 250 one min tests
17 | If, rechargeable battery
18 | - Battery life | Specify
19 | - Charger | Yes
20 | AUTO SHUT-OFF : | After 3 min. if left ON
21 | BATTERY CHECK INDICATOR | Yes, (LED).
22 | EXTERNAL OUTPUTS: | Phone-220 mV Impedance 270Ω, Record 500 mV/kHz (3.5 mm jack) Impedance 20 kΩ
23 | HEADSET / HEADPHONE | Yes
24 | PROBE HOLDER | Yes
25 | USAGE | Detecting peripheral pulses
26 | PRINTER | Specify
27 | Computer Interface Or PC-based | Yes
28 | ACCESSORIES :
29 | - Large Carry Bag | with complete Assessment kit
30 | - Probes | Yes
31 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Company Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | 4218 1900 01800 | 0 | ECG 12 LEAD WITH CART |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Cart | Yes
3 | LEADS
4 | Lead switching | Automatic/Manual
5 | Sensitivity, mm/mV | 5, 10, 20., other specify.
6 | Calibration signal | Automatic
7 | FREQUENCY RANGE, HZ
8 | Diagnostic | 0.05-150
9 | Filtered | Notch, EMG, Baseline, Wander, filters
10 | INPUT IMPEDANCE,
11 | Mega ohms | one hundred
12 | CMRR @ 60 Hz, dB | 100-140 approx
13 | Leads-off indicator | Yes
14 | RECORDER
15 | Channels | twelve
16 | Recording | YES A4 PAPER
17 | Lead marker | Automatic
18 | Timing marker | Yes
19 | Event marker | Yes
20 | Chart speed, mm/sec | 5, 10, 25, 50.
21 | LCD MONITOR 10 OR MORE | Yes ADJSTABLE
22 | No. of traces | 3, 6, 12
23 | PREVIEW SCREEN | Yes
24 | No. Waveforms stored | Min. 30
25 | ECG transmission | Yes
26 | Interpretation | Yes
27 | ECG measurements | Yes
28 | Auxiliary output | Yes
29 | Auxiliary input | Yes
30 | DEFIBRILLATOR SYRORNAIZ | YES
31 | SINGLE AVREGING | YSE
32 | Overload protection | Yes
33 | BATTERY OPERATION | Yes
34 | Battery type | Built in rechargeable
35 | Operating time, hr | Specify
36 | ACCESSORIES | All standard accessories
37 | PRINTER | Built-in
38 | A4 paper | Yes
39 | SAFTY AND PERFORMANCE | IEC 60601-2-51
40 | POWER REQUIREMENTS | 220V, 60Hz
41 | OTHER SPECIFICATIONS | Spare patient cable, chest electrodes, clamp electrodes for limb leads, with ICU Analysis. SOFT WARE Patient cable holder.FDA, CE, ISO, Approved. HL7
1 | CONFIGURATION | Mobile
Company Stamp |  |  |  | Company Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | 4214 0000 03200 | 0 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Flow Range | 0 - 15 L / Min
3 | Color Coded | specify
4 | Pressure Compensated | Yes
5 | Glass Envelop | Included with indicator ball
6 | Nipple | Yes
7 | Probe | As Per Site
1 | Flow Meter | Wall Mounted
Company Stamp |  |  |  | Company Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | 4214 0000 02800 | 0 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Pressure compensated
2 | Flow tube | Impact resistant
3 | Flow indication tube | Single
4 | Lower scale | Expanded
5 | Flow range | Up to 15 L/min, approx
6 | Humidifier cup volume | Reusable
7 | Flowmeter compatibility to all hospital wall outlets | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | 4110 3005 00000 | 0 | FREEZER ULTRA LOW TEMPERATURE -20 TO -40 C |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TEMPERATURE
2 | Temperature Control | Microprocessor
3 | Temperature Range | (-20⁰ C to -40⁰ C)
4 | Temperature Accuracy | ± 1⁰C
5 | Temperature Adjustable | Yes
6 | Temperature Display | Yes
7 | Factory Preset Temperature | Specify
8 | Average Temperature Stability | Specify
9 | Cooling & Air Circulation Method | Forced Cooling
10 | Temperature Recorder with Back up Battery | Yes
11 | TECHNICAL SPECIFICATION
12 | Capacity | Approx. 23 to 24 Cu. Ft.
13 | Shelves | Yes, Six Stainless Steel Shelves
14 | Exterior Dimension & Material Type | Electrostatic Powder Coated Stainless Steel
15 | Interior Dimension & Material Type | Stainless Steel
16 | Door Type | Single, Solid Stainless Steel with Key Lock
17 | Door Gasket | Magnetic
18 | Enlarge Digital Display | Yes / Touch Screen
19 | ALARM
20 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temeperature, Door Ajar, Power Failure & Condensor
21 | Microprocessor with Back up Battery | Yes
22 | Adjustable Alarm Volume | Yes
23 | Remote Alarm Connected to Central Alarm System | Yes
24 | Visual / Audio Alarms | Yes, Flash & Buzzer
25 | Safety Lock to Prevent Unauthorised Access | Yes
26 | WIRELESS MONITORING SYSTEM
27 | Integrated Wireless Data Logging System | Yes
28 | Digital Display | Yes, 5" TFT Display or more
29 | USB Dowload Data | Records Temperature, Humidity, Alarm & Door Ajar Information
30 | Ports | Two USB / Ethernet
31 | Computer Interface | Yes, Software to Monitor & Data Recording
32 | CASTERS | Heavy Duty for Easy Mobility with Leveling
33 | INTERIOR LIGHTS | LED
34 | OTHER SPECIFICATION
35 | Automatic Defrost | Yes
36 | High Density CFC-Free Urethane Foam Insulation | Yes
37 | Medical Grade Hemetically Sealed Compressor | Yes
38 | Resetable Door Open Counter | Yes
39 | Upgradeble Memory | Yes
40 | Digital Calibration | Yes
41 | Noise Level (Quiet) | ˂ 60 dB
42 | Access Port | Additional Ports to map Temperature
43 | ACCESSORIES INCLUDED
44 | Integrated Electromagnetic Access Control | Yes
45 | Mounting Kit | Yes
46 | USB | Yes
47 | PHYSICAL SPECIFICATIONS
48 | Dimensions | Specify
49 | Weight | Specify
50 | INSTALLATION & PRE - INSTALLATION IF REQUIRED, SHOULD BE DONE BY THE COMPANY, (Civil, Electrical, Plumbing Works etc….) | Yes, attach separate scope with details
51 | LINE POWER
52 | AC Voltage / Phase | 220 V / Single Phase
53 | Current | 13 A for power socket
54 | Frequency | 60 Hz.
55 | Plug Type | 3 Pin British
56 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | 4217 0000 01200 | 0 | GLUCOMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Should be IVD marked (In Vitro Diagnostic Medical Device).
3 | Measuring range: 20-500mg/dl Approx.
4 | Must give accurate results within Approx. 15 seconds.
5 | Must store at least the 20 results.
6 | Must be supplied with 200 strips.
7 | Must be supplied with a virtually pain-free blood sampling tool (lancing device).
8 | Specify the shelf life of strips (vial closed, after opening the vial)
9 | Must be supplied with long-life batteries.
10 | Unit must have an automatic shut off to protect battery life
11 | Unit must have low battery warning
12 | Must have a case for storage.
13 | Compliance with standards & legislation:
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Microprocessor-controlled, hand held glucose meter to measure blood glucose levels.
Company Stamp |  |  |  | Company Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | 4229 5100 02600 | 0 | HEADLIGHT FIBEROPTIC SURGICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Headlight with ball joint and vertical height adjustment | Yes
3 | front selection movable in all directions | Yes
4 | holder with ball joint | Yes
5 | fiber optic light cable, diameter 3.5 mm Length 230cm | Yes
6 | fixation clip | Yes
7 | Power | 220v 60 hz
8 | OTHER SPECIFICATION | All access. needed
1 | Light source 300 watt | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | 4115 0000 17200 | 0 | KIT SPILL CHEMICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Features and contents :
3 | Lightweight, durable, hard-sided, wall mounting bracket .
4 | Equipped with tools, spill neutralization agent and healthcare disinfectant
5 | Pack for a single blood or urine/vomit spill
6 | Certified to World Health Organization.
7 | Detailed, illustrated, laminated instruction sheets
8 | Refills available
9 | Spill Kit Content :
10 | Clean-up Items
11 | Chemical neutralizer
12 | Forceps or tongs
13 | Broom and dustpan
14 | Paper towels or suitable absorbents
15 | Sharps bin
16 | Plastic bags
17 | Personnel Protection Items
18 | Goggles
19 | Disposable laboratory coats
20 | Gloves
21 | Respiratory mask.
22 | Mask .
23 | Gas cartridge .
24 | Shoe covers
25 | Instructions Quick Guides:
26 | Chemical spill response procedures
27 | Emergency contact numbers
28 | Chemical spill signage
29 | Inventory of chemical spill kit
30 | A sample of items, including its parts shall be submitted for approval before an order is placed ,Approved samples shall be used as standards of finish and workmanship.
1 | CHEMICAL SPILL KIT For Emergency cases in laboratory use.
Company Stamp |  |  |  | Company Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | 4218 1900 03100 | 0 | LARYNGOSCOPE ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | BLADES
3 | Adult | 4 Different
4 | Pediatric | 4 Different
5 | ILLUMINATION | xenon
6 | Spare lamps | 5
7 | CHARGER, | Yes
8 | Rechargeable Battery | Yes
9 | Power, vac, hz. | 220, 60
10 | Wall mounted | to be improvised
11 | ALL ACCESSORIES (ITEMIZEDPRICE) | Yes
12 | OTHER SPECIFICATION | FDA, CE, ISO Approved
1 | TYPE, CONVENTIONAL OR FIBER OPTIC | MacIntosh
Company Stamp |  |  |  | Company Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | 4218 1900 03300 | 0 | LARYNGOSCOPE PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | BLADES
3 | Pediatric (Different Sizes) | 2 Different Sizes (0,0 / 0,1 / 1,1 / 1,2 )
4 | Pediatric (Different Sizes) | 2 Different Sizes 0,1,2,3
5 | ILLUMINATION | With Incorporated Fiber Optic Light Carrier inside the Blade (Xenon Light)
6 | Spare lamps | 6
7 | Rechargeable Battery Shell | With High Power LED Technology, more than 50,000 lux, Lithium-ion batteries
8 | Reduction Sleeve | For rechargeable batteries
9 | Handle Sleeve | Yes
10 | Charging Unit | For 2 Rechargeable batterie (110-240 VAC, 60 Hz)
11 | Bag for All Laryngoscopes | Splash - Protected
12 | Other Specifications | FDA, ISO Approved.
1 | Laryngoscope Set : | Neonatal & Infant
Company Stamp |  |  |  | Company Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | 4214 0000 04200 | 0 | LIGHT EXAM MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | On/Off switch, easily accessible on the light stand or light head
3 | Stability and precise positioning in every desired position.
4 | Weight balancing system in the rotating joint
5 | Smooth surfaces for perfect hygiene
6 | Light intensity: approx. 40,000 lux at 80 cm
7 | Quick lamp replacement
8 | Color temperature ≥ 3500K
9 | Temperature increase in the head area ≤ 2° C
10 | Angle of rotation: unlimited
11 | Minimal heat dissipation
12 | 3 white LED lights
13 | Sterilizable hand grip
14 | Main and battery-operate
15 | Shadowless light
16 | Swiveling castors with locking feature
17 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
1 | Mobile examination light on sturdy and stable base
Company Stamp |  |  |  | Company Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | 4115 0000 21900 | 0 | MIXER BLOOD ROLLER TYPE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Rolling only and rolling and tilting modes of operation should be allowed
3 | Fixed speed at ~ 40 rpm
4 | Roller length ~ 35 cm
5 | Vertical movement ~ 16 mm
6 | To incorporate ~ 5 rollers for a total capacity of ~ 15 tubes
7 | Individual rollers should be removable to accommodate larger bottles
8 | Compliance with standards & legislation:
9 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
10 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
11 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
12 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
13 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Rolling and tilting tube mixer for homogeneous, gentle mixing of hematology blood sample tubes
Company Stamp |  |  |  | Company Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | 4218 1900 04300 | 0 | MONITOR CAPNOGRAPHY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | measurements | EtCO2,FiCO2
3 | respiration rate | 0 bpm 150 bpm
4 | spo2 | yes
5 | printer | yes
6 | disply | color screen
7 | wave form trends | yes
8 | trends storage | yes 24hr.,specify
9 | accuracy | specify
10 | alarms | yes, visual and audible,specify
11 | Battery | yes
12 | Battery operating time | hr? specify
13 | power Supply | 220 VAC / 60 HZ
14 | All standard accessories required. | Yes ,Provide Itemized Price
1 | application | monitoring CO2 measurements for neonatal, pediatric and adult patients
Company Stamp |  |  |  | Company Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | 4218 2005 00300 | 0 | OPTHALMOSCOPE DIRECT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Apertures | specify
3 | Corrective lenses range | -35 to + 40
4 | number of lenses | specify
5 | Rechargeable battery | yes
6 | other specifications | FDA or CE marked
1 | Illumination | specify
Company Stamp |  |  |  | Company Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | 4214 0000 04500 | 0 | OTOSCOPE TABLETOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | 3.5V Bulb type (Xenon) | Yes.
3 | Magnification more than 2.5x | Yes.
4 | Insuflation port | Yes.
5 | Insufflation Bulb | Yes.
6 | instrument housing | Specify.
7 | Smart design | Yes.
8 | Diameter of ear-reusable specula | 2.5, 3, 4, and 5mm approx.
9 | Disposable Tip for adult and pediatric | 200 tips for each approx.
10 | Three Spare bulb | Yes.
11 | Battery rechargeable | LI-ION
12 | Charger table top | Yes.
13 | Hard case | Specify.
14 | Power Supply 220 VAC / 60 HZ | Yes.
15 | FDA or CE approved | Yes.
1 | Fiber optics Otoscope | Yes.
Company Stamp |  |  |  | Company Signature

## ITEM 42
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
42 | 4214 0000 04800 | 0 | OXIMETER PULSE HAND HELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | DISPLAYS | SpO2, pulse rate, pulse strength and/or signal, low battery, Volumes, alarm limits.
3 | Type | LCD/ with backlit LED SHOWING PARAMETERS & READINGS
4 | Sp02 RANGE, % | 0-100
5 | Accuracy, ,% | < ± 2 %
6 | PULSE RATE, bpm | 30-240
7 | PERFUSION INDEX AND/OR SIGNAL STRENGTH INDICATOR | YES
8 | RESPONSE TIME, sec | Pulse to Pulse or time ( Specify)
9 | SETTLING TIME | Pulses or time ( Specify)
10 | START-UP TIME, sec or pulses | <10 sec
11 | ALARMS:
12 | Audiovisual | Yes
13 | Visual | Yes, Low battery
14 | ALARM OVERRIDE | Yes
15 | Reactivation method | Yes, after 2 mins.
16 | Volume control | Yes
17 | SELF-TEST MODE | Yes
18 | Memory: up to 72 hours of data storage | YES
19 | PROBE TYPES | YES REUSABLE (1 ADULT, 1 PEDIATRIC, AND 1 NEONATE)
20 | Patient range | Adult / Neonate/ Pediatric
21 | Cable length, m | Not less than1.5 m
22 | BATTERY BACKUP :
23 | Rechargeable battery | Yes, Specify the type
24 | Low battery notice | Yes
25 | Rechargeable time, hr | Specify
26 | Battery life, hrs | Specify
27 | The unit should operate on batteries as well as AC electricity, with automatic battery charging while in use. | YES
28 | The unit shall operate on batteries | YES
29 | battery type | Specify
30 | VAC | 220 V
31 | CURRENT | 13 A
32 | FREQUENCY | 60 Hz.
33 | PLUG TYPE | 3 Pin British
34 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
35 | IMPORTANT NOTE: | Please specify whether accessories are standard or optional, otherwise they will be considered standard.
36 | MDMA | Yes
1 | TYPE | Portable / Stand-alone/transport/HAND HELD FOR ADULT/PEDIATRIC/INFANT
Company Stamp |  |  |  | Company Signature

## ITEM 43
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
43 | 4115 0000 24400 | 0 | PIPETTOR ADJUSTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Adjustable volume pipette set to include:
3 | 1 x 0.1-2.5μl
4 | 1 x 0.5-10μl
5 | 1 x 5-50μl
6 | 1 x 10-100μl
7 | 1 x 100-1000μl
8 | 1 x 1000-5000μl
9 | Shall state accuracy and precision for each of the above sizes
10 | One channel
11 | High accuracy and precision
12 | Adjustable tip ejector fitting most tips
13 | Shall include user calibration set
14 | Shock, UV-light and autoclaving resistance
15 | To include a stand that holds up to 6 pipettes
16 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
17 | Shall list all the necessary accessories for proper installation and functioning
1 | Light weight and easy to use
Company Stamp |  |  |  | Company Signature

## ITEM 44
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
44 | 4220 3700 00500 | 0 | PROTECTION X-RAY APRON VEST & SQUIRT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Back Flap for better weight balance
3 | Built-in Black Bel tin Vest and Skirt for better lumbar support
4 | Partial overlap to increase protection
5 | Professionally designed for an exceptional fit and comfort
6 | VEST Specification:
7 | Each front panel: .5 mm Pb (1.0 mm Pb at overlap)
8 | Back panel: .25 mm Pb
9 | SKIRT Specifications
10 | Each front panel: .25 mm Pb (.5 mm Pb at overlap)
11 | Specify all available colors
12 | Medium size
13 | Features ;
14 | High Quality alloy.
15 | Easily cleaned front
16 | Non-slip backing
17 | Stain-resistant binding
18 | Permit hanging
19 | Protective multi-layered
20 | Provides protection secondary (scatter) radiation.
21 | radiological lab inspected and certified.
22 | Souple, drapes comfortably
23 | Lightweight, easy handling
24 | Life cycle and store time :7-10 years specify
1 | Lumbar Vest & Skirt apron
Company Stamp |  |  |  | Company Signature

## ITEM 45
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
45 | 4218 1900 06500 | 0 | PUMP INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | QUICK START OF INFUSION | YES
3 | FLOW RANGE ML/HR | 0.1 - 999 ML/HR
4 | ACCURACY | 5%
5 | DRUG/DOSE CALC | YES
6 | VTBI, ML | 0.1 TO 9999
7 | MAX PRESSURE, PSI | SPECIFY MAX
8 | PUMPING MECHANISM | SPECIFY MECHANISM
9 | IV SET | YES
10 | FREE-FLOW PROTECTION | YES
11 | MANUAL AND PROGRAMMABLE BOLUS | PLEASE SPECIFY BOLUS RATE
12 | ALARMS & INDICATORS
13 | EMPTY BAG | YES
14 | FLOW ERROR | YES
15 | DOOR OPEN | YES
16 | CONTINUOUS PRESSURE MONITORING PSI (ON THE FRONT SCREEN ALL THE TIME) | SPECIFY
17 | DRUG DOSE CALCULATION | ML/H, µG/H, µG/MIN, µG/KG/H, µG/KG/MIN, MG/H, MG/MIN,MG/KG/MIN, MG/KG/H, IU/H, IU/MIN, IU/KG/H, IU/KG/MIN.
18 | FLEXIBLE PRESSURE ALARM SETTINGS | IN LINE PRESSER
19 | HIGH PRESSURE/OCCLUSION | YES
20 | PRE- OCCLUSIONALARM | YES
21 | CIRCUIT MALFUNCTION | YES
22 | INFUSION NEAR END | YES
23 | INFUSION COMPLETE | YES
24 | LOW BATTERY | YES
25 | DEPLETED BATTERY | YES
26 | USER PROMPT | YES
27 | AUDIBLE ALARM | YES
28 | VOLUME CONTROL | YES
29 | MOMENTARY SILENCE | YES
30 | PPM DUE DATA REMINDER | YES
31 | FLUID RESISTANT | YES
32 | NURSE CALL JACK | YES
33 | DATA LOG | YES
34 | EVENT LOG | YES
35 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED)
36 | NUMBER OF EVENTS | 1500
37 | PRINT OUT | YES
38 | DISPLAY | LED AND LCD WITH COLOR CABLITIES IS A PLUS & BIG SCREEN INTERFACE IS ALSO A PLUS
39 | INFUSED VOLUME DISPLAY | YES , DIGITAL
40 | WEIGHT WITH BATTERY KG | SPECIFY WIGHT
41 | TIME /DATE | YES
42 | POWER SOURCE, VAC | 220VOLT / 60 HZ.
43 | POWER SUPPLY | INTERNAL OR EXTERNAL (IF EXTERNAL MUST BE LIFETIME WARRANTY)
44 | POWER CABLE | DURABLE POWER CABLE
45 | BATTERY | YES, RECHARGEABLE BUILT IN
46 | BATTERY | NI-MH
47 | OPERTING TIME , HR | SPECIFY HOURS @ 10 ML/HR
48 | RECHARGE TIME , HR | PREFERRED LOWER CHARGING HOURS
49 | POLE MOUNTING | PREFERRED
50 | OTHER SPECIFICATIONS | FDA / CE
51 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED).
52 | STACKABLE IN THE SAME DOCKING STATION | INFUSION PUMPS / SYRINGE PUMPS / ENTRAL FEEDING PUMPS.
53 | DOSE ERROR REDUCTION SYSTEM (SMART TECHNOLOGY) | YES
54 | LIBRARY SIZE | YES CONFIGURABLE
55 | NO. OF DRUG ENTITIES/CARE AREA | YES PLEASE SPECIFY
56 | WIRELESS CONNECTIVITY | UPGRADABLE
57 | TIME RETAINED | FIRST IN FIRST OUT
58 | ACCESSORIES FOR THE PUMP | CLOSED SYSTEM
59 | ALL SETS HAVE SELF -SEALING NEEDLE FREE INJECTION SITES | YES
60 | ALL SETS SHOULD BE LATEX DEHP & PVC FREE | YES
61 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
62 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
63 | ACCESSORIES FOR INFUSION PUMP | CLOSE SYSTEM
64 | THE VENDOR MUST PROVIDE A FIXED PRICE FOR THE ALL ABOVE MENTIONED ACCESSORIES VALID FOR 3 YEARS | YES
65 | PUMP HL7 READY | YES ( PLEASE SPECIFY )
66 | PROVIDE CONSUMABLE DEAL ( PUMPS FREE AND PRICE ON SETS ) | OPTIONAL DEAL SECTOR MIGHT CHOOSE
67 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
68 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
69 | THE WORKSTATION SHOULD BE MOUNTABLE , WALL UNIT (FIXATION TO WALL RAILS OF DIFFERENT MANUFACTURER).MOVABLE POLE-BASED UNIT.
70 | SMART ALARM SYSTEM BY AUDIBLE, VISUAL AND DISPLAYED MESSAGE. | OCCLUSION, END OF INFUSION, EMPTY BATTERY, NO RATE SET, DEFECT ON PUMP. PRE-ALARMS: LOW BATTERY, NEAR-END OF INFUSION , PLEASE SPECIFY
71 | MOBILE STAND TO FIX THE PUMP ON INSTEAD OF PENDANT OR WALL FOR SPECIAL SITE NEEDS | ITEMIZED PRICE OPTIONAL
72 | PRIMARY INFUSION SET (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 120 SETS WITH EACH PUMP ITEMIZED
73 | LIGHT SENSITIVE PRIMARY INFUSION SET WITH NO INJECTION PORT FOR NEONATAL TPN (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 40 SETS WITH EACH PUMP ITEMIZED
1 | PROGRAMMING PATINET ID / NURSE ID ON THE PUMP DURING ADMINISTRATION | YES ( PLEASE SPECIFY )
Company Stamp |  |  |  | Company Signature

## ITEM 46
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
46 | 4218 1900 07000 | 0 | PUMP SYRINGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Variable rate ranging from 0.1 to 500 mL/hr or better, with 0.1 mL/hr increments
3 | 3 % accuracy or better
4 | Variable volume-to-be-infused from 1 to 1,000 mL or similar
5 | Digitally displayed parameters to include:
6 | Infusion rate
7 | Battery / AC operation
8 | Running indicator
9 | Alarming condition when active, with indication of alarm type or code
10 | Back pressure monitor / indicator
11 | Event history
12 | Capability to accept different syringe types and sizes with automatic syringe detection and identification
13 | Syringe compatibility and auto detection shall include but not be limited to all sizes of the following (3 to 60 mL):
14 | BD
15 | Terumo
16 | Monoject
17 | Braun
18 | Fresenius
19 | Variable bolus rate
20 | Up to maximum flow rate
21 | Bolus infused volume indicator during bolus activation
22 | Protected access
23 | Audiovisual alarms shall include but not be limited to the following:
24 | Syringe installation and integrity (detection)
25 | Line disconnection (sudden drop in back pressure)
26 | Occlusion pressure pre-alarm
27 | Occlusion pressure
28 | Near end of infusion alarm
29 | End of infusion
30 | Volume limit pre-alarm
31 | Volume limit
32 | KVO (1 ml/hr; if other, specify)
33 | Low battery pre-alarm
34 | Discharged battery
35 | Internal malfunction
36 | Syringe unlocked
37 | Plunger disengaged
38 | Empty syringe
39 | Data log capability and data port for data transmission, display and printing. Any required software for such function shall be included.
40 | Logged data to include ( ≥1 year) :
41 | Settings
42 | Alarms
43 | Errors
44 | Key presses
45 | Error codes
46 | Amount infused
47 | Safety features shall include but not be limited to:
48 | Self test at start-up
49 | Nurse call interfacing capability
50 | Splash proof design
51 | Auto priming
52 | Adjustable alarm volume. No permanent silencing shall be possible.
53 | Keypad lock
54 | Impossibility to improperly install infusion set
55 | Free flow prevention system
56 | Last parameter setting retention
57 | IV stand mounting accessory shall be included
58 | Battery autonomy of 3 hrs or more when fully charged. specify:
59 | Battery type and characteristics (voltage and current capacity)
60 | Autonomy at 10 mL/hr
61 | Recharging time from depleted to 90%
62 | Compliance with standards & legislation:
63 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
64 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
65 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
66 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
67 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
1 | Microprocessor controlled with digital LCD alphanumeric display of parameters and alarms
Company Stamp |  |  |  | Company Signature

## ITEM 47
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
47 | 4115 0000 26500 | 0 | RACK BLOOD SEDIMENTATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Designed for Westergren and Wintrobe tubes
3 | Leveling bubbles and adjustment screws allow tubes to be held firmly in a perpendicular position
4 | Reduce the chance of bio hazardous contamination
5 | Light weight
6 | Approx. Dimensions: 12"W x 12"H x 2"D
1 | Ten tube capacity
Company Stamp |  |  |  | Company Signature

## ITEM 48
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
48 | 4110 3011 00700 | 0 | REFRIGERATOR MEDICATION 700L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TEMPERATURE
2 | Temperature Control | Microprocessor
3 | Temperature Range | 1 to 12 ⁰C
4 | Factory Preset | 4 ⁰C
5 | Accuracy | ± 1 ⁰C
6 | Average Temperature Stability | ± 3 ⁰C
7 | Temperature Adjustable | Yes
8 | Automatic Defrost | Yes
9 | Cooling & Air Circulation Method | Forced Cooling
10 | FEATURES
11 | Capacity | 25 Cu.Ft. Approx
12 | Exterior Dimension | Specify
13 | Interior Dimension | Specify
14 | Interior Finish | Stainless Steel
15 | Door Type | Single Glass Door
16 | Gasket | Magnetic
17 | Access Ports | Additonal Probes to Map Temperature
18 | Digital Display | Specify
19 | ALARM
20 | Automatic Alarm ( Selectable, Adjustable, Resetable ) | Includes Hi/Lo Temperature, Power Failure, Door Ajar & Condensor
21 | Remote Alarm Connected to Central Alarm System | Yes
22 | Microprocessor with Backup Battery | Yes
23 | Visual / Audio Alarms | Yes
24 | Key-Operated Alarm & Setpoint Security | Lock in Temperature & Setpoint Security
25 | OTHER SPECIFICATION
26 | High Density CFC-Free Urethane Foam Insulation | Yes
27 | Number of Adjustable Shelves & Type | Specify
28 | Medical Grade Hermetically Sealed Compressors | Yes
29 | Upgradeble Memory | Yes
30 | Digital Calibration | Yes
31 | Serial Port | USB
32 | Noise Level | Specify
33 | Interior Led Lights | Yes
34 | ACCESSORIES INCLUDED
35 | Temperature Chart Recorder | Yes
36 | Casters | Heavy Duty for Easy Mobility with Leveling
37 | USB | Yes
38 | Mounting Kit | Yes
39 | Integrated Electromagnetic Access Control | Yes
40 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Company Signature

## ITEM 49
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
49 | 4214 0000 05600 | 0 | REGULATOR SUCTION HIGH WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MOUNTING | Wall / Pneumatic
2 | STANDARD FEATURES AND TECHNICAL SPECIFICATIONS
3 | Autoclaveable Bottles | Yes
4 | Suction Regulator Modes | Continuous & Intermittent
5 | Adjustable Timing Cycles | At least from 3 to 30 sec
6 | Fully Adjustable Negative Pressure | From 0 to -200 mmHg or more
7 | Flow Rate at Continuous Mode | At least 90 L/min
8 | Overflow shut off valve/trap | Yes
9 | British Standard Probe | Yes
10 | Compatible with KFMC Standards | Yes
11 | Suction Bottles and Canisters | Disposable / reusable
12 | PHYSICAL SPECIFICATIONS
13 | Dimensions | Specify
14 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 50
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
50 | 4214 0000 06100 | 0 | SCALE CHAIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Clinical, chair
2 | TYPE | Electronic
3 | MEASUREMENT UNITS
4 | Weight | Metric
5 | Height | Metric
6 | CAPACITY | 250 kg. Approx.
7 | ACCURACY | 0.1 g
8 | CALIBRATION ADJUST | Yes
9 | Power | 220V 60Hz
10 | TARE WEIGHT ADJUST | Yes
11 | OTHER SPECIFICATIONS. | Chargeable Battery with charger
Company Stamp |  |  |  | Company Signature

## ITEM 51
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
51 | 4214 0000 06000 | 0 | SCALE PATIENT WITH HEIGHT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Measure Weight and Body Mass
2 | DISPLAY | Digital, LCD
3 | UNIT DISPLAYED | Kg & lbs
4 | TRANSDUCER TYPE | Load cell / Strain gauge
5 | BODY MASS INDEX FUNCTION | Yes
6 | MAX WEIGHT | Not less than 270 Kg
7 | WEIGHT CHANGE | 0.05 Kg
8 | ACCURACY | ± 2%
9 | ZERO ADJUSTMENT | Automatic
10 | CALIBRATION ADJUSTMENT | Yes
11 | HEIGHT SCALE OPTION | Specify
12 | - Max height | Not less than 200 cm
13 | BATTERY | Yes
14 | - Type | Rechargeable
15 | - Charger | Internal
16 | - Operating time /Charging time | Specify/Specify
17 | The scale should be highly accurate with no need of adjustment or periodic calibration. Accuracy shall be better than or equal to 1 %. Provide detailed characteristics | Yes
18 | Unit shall include a measuring column made from corrosion resistant hard aluminum and shall have a measuring range of approximately 750 mm to 2,000 mm in divisions of approximately 5 mm. | Yes
19 | Measuring rod shall have both inches/cm increments. | Yes
20 | Calibration weight set to be included | Yes
21 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 52
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
52 | 4229 5100 05700 | 0 | SCOPE SET RIGID ENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
3 | COTTLE Nasal Speculum, length 13 cm, with set screw, 35 - 45 mm | Yes
4 | COTTLE Nasal Speculum, lenght 13 cm, with set screw, 55 mm | Yes
5 | COTTLE Nasal Speculum, length 13 cm, with set screw, 75 mm | Yes
6 | KILLIAN-STRUYCKEN Nasal Speculum, 15 cm, with set screw, 90 mm. | Yes
7 | Tonsil-Syringe, "ULTRA-ASEPT", glass barrel interchangeable, LUER-Lock, 5 cc. | Yes
8 | Surgical Handle, No. 3, length 12.5 cm, for blades | Yes
9 | Handle No. 4 | Yes
10 | Surgical Handle No. 7 , for blades | Yes
11 | FREER Elevator, double-ended, length 20 cm. approx. | Yes
12 | MASING Elevator, double-ended, graduated, sharp and blunt, 22.5 cm | Yes
14 | Double Hook, sharp, 13 -15 cm, 10 - 12 mm wide | Yes
15 | REYNOLDS Dissecting Scissors, delicate tips, curved, 15 cm approx. | Yes
16 | METZENBAUM-LAHEY Scissors, curved, delicare, 13 - 15 cm | Yes
17 | ADSON Tissue Forceps, MICRO model, 1 x 2 teeth, 12 - 15 cm | Yes
18 | ADSON Tissue Forceps, 1 x 2 teeth 12 cm | Yes
19 | Tissue Forceps, 1 x 2 teeth, length 13 cm, standard size | Yes
20 | Tissue Forceps, 1 x 2 teeth, length 14.5 cm, delicate size | Yes
21 | ADSON-BROWN Tissue Forceps, non-traumatic, fine side grasping teeth, length 12 cm | Yes
22 | Dressing Forceps, length 13 cm, delicate size approx. | Yes
23 | Dressing Forceps, length 13 cm, standard size approx. | Yes
24 | COTTLE Lower Lateral Forceps, with set screw, bayonet, 12 - 15 cm | Yes
25 | GRUENWALD Nasal Dressing Forceps, bayonet, 20 cm | Yes
26 | TAKAHASHI Nasal Forceps, straight, working length 9 - 12 cm | Yes
27 | BLAKESLEY Nasal Forceps, straight working lenght 11 cm, size 2 | Yes
28 | BLAKESLEY Nasal Forceps, size 3 working lenght 11 cm, size 3 | Yes
29 | HEYMANN Nasal Scissors, medium standard model, working lenght 9.5 cm | Yes
30 | CRAIG Septum Forceps, special matte finish, working length 9- 11 cm, straight | Yes
31 | Suction Tube, with cut-off hole and stylet, LUER, workinge lenght 11 cm, 9 Fr. | Yes
32 | Suction Tube, with cut-off hole and stylet, LUER, working length 11 cm, 10 Fr. | Yes
33 | Suction Tube, with cut-off hole and stylet, LUER, working lenght 11 cm, 12 Fr. | Yes
34 | Suction Tube, with cut-off hole and stylet, LUER, working length 11-14 cm, 15 Fr. | Yes
35 | Suction Elevator with stylet, length 19.5 cm. | Yes
36 | Osteotome, flat, 19 cm 3 mm wide | Yes
37 | Osteotome, flat, 19 cm, 4 mm wide | Yes
38 | Osteotome, flat, 19 cm, 7 mm wide | Yes
39 | COTTLE Mallet, lenght 18 cm approx. | Yes
40 | Bipolar Forceps, insulated, bayonet shape, blunt, tips 1 mm, length 19 cm | Yes
41 | Insulated Suction Cannula, for nose and epistaxis, angular, distal with non-insulated horn for coagulation, O.D. 3.5 mm, working length 12 cm | Yes
42 | Insulated Suction Cannula, for nose and epistaxis, angular, distal with non-insulated horn for coagulation, O.D. 4 mm, working length 12 cm | Yes
43 | Insulated Suction Cannula, for nose and epistaxis, angular, distal with non-insulated horn for coagulation, O.D. 4.5 mm, working length 12 cm | Yes
44 | Nasal Snare, 25 cm | Yes
46 | Haemostatic Forceps, straight, length 14 cm Qty:2 | Yes
47 | HALSTEAD Mosquito Forceps, 12.5 cm, curved approx. Qty:2 | Yes
48 | Haemostatic Forceps, curved, length 14 cm | Yes
49 | HALSTEAD Mosquito Forceps, 1 x 2 teeth, 12.5 cm, straight. 12.5 cm, straight approx. Qty:2 | Yes
50 | Needle Holder, tungsten carbide inserts, 15 cm | Yes
51 | Needle Holder, tungsten carbide inserts, 17 cm | Yes
52 | MAYO Scissors, straight, length 15-17 cm | Yes
53 | MAYO Scissors, curved, length 15-17 cm | Yes
54 | Dressing Forceps, curved, with ratchet, length 21 cm | Yes
55 | Towel Forceps, with serrated inserts, for paper cover sheets, length 9.5 cm Qty:10 | Yes
57 | Nasal Snare Wires | Yes
58 | Monopolar High Frequency Cord | Yes
59 | Bipolar High Frequency Cord | Yes
60 | Cup Medicine, stainless steel, 50 cc, 47 mm high, diameter 57 mm | Yes
61 | Cup Medicine, stainless steel, 100 cc, 30 mm high, diameter 80 mm Qty:2 | Yes
62 | Plastic Container for Sterilizing and Storage | Yes
2 | Telescope Handle, round, standard model, length 11 cm, for use with Telescope 30°-120° with diameter 4 mm and length 18 cm | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 53
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
53 | 4214 0000 06300 | 0 | SPHYGMOMANOMETER ANEROID MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | INFLATION PRESSURE, mm Hg | 300 mm Hg
3 | CUFF SIZES | Adult, Pediatric, Velcro cuff
4 | SYSTOLIC, mm Hg | 0-240 approx.
5 | DIASTOLIC, mm Hg | 0-130 approx.
6 | OTHER SPECIFICATIONS
1 | METHOD | ANEROID
Company Stamp |  |  |  | Company Signature

## ITEM 54
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
54 | 4214 0000 06500 | 0 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Pole mounted for Adult and Pediatric
2 | CONFIGURATION | Configured
3 | STANDARD PARAMETERS FOR MONITOR | NIBP, SpO2, temperature
4 | NIBP | Yes
5 | Measurement technique | Oscillometric
6 | Cuff Inflation Adjustment Range for Adult/Pediatric | Automatic
7 | Cuff Pressure Range for Adult/Pediatric, mmHg | 0 to 300
8 | Determination Time, sec | 20 - 40
9 | Pressure Reset Levels | Resolutions of 5 mmHg up to maximum limits
10 | Systolic Determination for adult | Max:300 mmHg and Min:30 mmHg
11 | Diastolic Determination for adult | Max:195mmHg and Min :10 mmHg
12 | Map Determination (Medium Arterial Pressure for adult) | Max:215 mmHg and Min :15 mmHg
13 | Pulse Rate Range for adult/pediatric, bpm | 30 - 200
14 | Over Pressure Cut off for adult/pediatric, mmHg | 300 - 330
15 | PULSE OXIMETRY | Yes
16 | Accuracy (+/- 1 standard deviation) | Adult/Neonate 70 - 100 % +/-3.0 digits; under 70% unspecified
17 | Saturation range, % | 0 - 100
18 | Pulse rate range, bpm | 20 - 250
19 | PROBE TYPES | Specify probe type. Disposable/reusable (Compatible with KFMC existing Type)
20 | Patient range | Adult / Neonate/ Pediatric
21 | Cable length, m | Not less than 2.5
22 | TEMPERATURE | Yes
23 | No. of inputs | 1
24 | Probe type | Oral, axillary, rectal
25 | Pole mountable (Included pole) | Yes
26 | ALARMS & TRENDING
27 | ALARMS | Yes
28 | - Systolic | Preset at 30 mmHg low and 180 mmHg High
29 | - Diastolic | Preset at 15 mmHg low and 130 mmHg High
30 | Silence active alarm | Yes
31 | TRENDING | Monitor stores 24 hrs of data; up to 100 reading within time period
32 | -Parameters | All
33 | -Graphical/ tubular | Yes
34 | BATTERY | Yes, Built in with integrated charger
35 | - Operating time | Not less than 2.5 Hrs
36 | PRINTER | Yes, Specify the paper type
37 | PHYSICAL SPECIFICATIONS
38 | Dimensions | Specify
39 | Weight | Specify
40 | ACCESSORIES
41 | - All sizes cuff for Adult | Yes
42 | - All sizes cuff for Pediatric | Yes
43 | - Spo2 probe/adult (reusable & disposable) | Yes
44 | - Spo2 probe/pedia (reusable & disposable) | Yes
45 | - Temperature probes | Yes
46 | - Paper roll | 1 box
47 | The Device Should Integrate with the current and future HIS. | Yes / Specify
Company Stamp |  |  |  | Company Signature

## ITEM 55
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
55 | 4214 0000 06600 | 0 | STAND IV |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | TYPE | Telescopic
3 | HOOKS | 4 - 5 Approx
4 | HEIGHT | Adjustable
5 | BASE | Mobile heavy base to withstand the hanged Items weights
6 | Castors | Qty : 5
7 | CAPACITY | 5 kg per hook
1 | FRAME | Stainless steel 304 preferred
Company Stamp |  |  |  | Company Signature

## ITEM 56
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
56 | 4214 0000 06700 | 0 | STETHOSCOPE ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | MATERIAL
3 | EAR TUBE | Stainless steel
4 | CHEST PIECE | Stainless steel
5 | BINAURAL
6 | ADJUSTMENT | To angle of ear
7 | CHEST PIECE | Adult and pediateric & Neonatal
8 | ADDITIONAL DIAPHARGM | 1 EACH
9 | EAR TIPS, TYPE | COMFORT SEALING
10 | OTHER SPECIFICATION
1 | TYPE | Double head
Company Stamp |  |  |  | Company Signature

## ITEM 57
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
57 | 4214 0000 06800 | 0 | STETHOSCOPE PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | MATERIAL
3 | EAR TUBE | Stainless steel
4 | CHEST PIECE | Stainless steel
5 | BINAURAL
6 | ADJUSTMENT | To angle of ear
7 | CHEST PIECE | pediateric & Neonatal
8 | ADDITIONAL DIAPHARGM | 1 EACH
9 | EAR TIPS, TYPE | COMFORT SEALING
10 | OTHER SPECIFICATION
1 | TYPE | Double head
Company Stamp |  |  |  | Company Signature

## ITEM 58
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
58 | 4214 0000 07100 | 0 | STOOL FOOT 2 STEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | MATERIAL | Specify the material type
3 | DIMENSION | ALMOST ( 14x19x6 ) WxLxH
4 | TOP SURFACE | SLIP RESISTANCE
5 | WEIGHT - ALONE | Minimum 2.5 Kg
6 | WEIGHT - STACKED | Maximum 200 Kg
7 | OTHERS | EASY CLEAN - INFECTION CONTROL REGULATION
1 | FOOT STOOL - two step
Company Stamp |  |  |  | Company Signature

## ITEM 59
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
59 | 4618 1810 01100 | 0 | STRETCHER EMERGENCY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Stretcher Type | Hydraulic
3 | Position Control | Dual sided foot pedal
4 | Trendelenburg | 12 degree APPROX.
5 | Reverse Trendelenburg | 12 degree APPROX.
6 | Emergency Trendelenburg Foot or Handle pedal | Yes, specify.
7 | Height Adjustment | Yes, specify range
8 | Back Rest | Adjustable Gas Spring Assisted
9 | Restraining straps | Yes
10 | Frame Material | Rigid tubular framework with epoxy painted steel.
11 | Patient platform | X-Ray translucent
12 | Safety Side rails | Integral Fold Down
13 | X-Ray cassette access | Yes
14 | Oxygen Cylinder Holder with Cylinder | Accept Size E
15 | Bumper | Yes
16 | Central Locking Bracking System | Four Corner Activation
17 | IV Pole | Yes
18 | Spare IV rod | Yes
19 | Sockets number
20 | Large Storage Area | Yes
21 | Maximum Patient Weight | 180 KG OR more
22 | Four 20 cm Casters | All Lockable Antistatic
23 | Mattres | Yes
24 | Thickness cms | 6.cm or more
25 | Monitor Shelf | Yes
26 | 5th Wheel | Yes
1 | Area of Use | Emergency
Company Stamp |  |  |  | Company Signature

## ITEM 60
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
60 | 4219 1800 03700 | 0 | STRETCHER RECOVERY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | FRAME MATERIAL | Epoxy painted steel
3 | STRETCHER TYPE | Hydraulic
4 | FOOT PEDAL OPERATED | Yes, both sides
5 | HEIGHT ADJUSTMENT | Yes
6 | TRENDELENBURG | 15 approx.
7 | REVERSE TRENDELENBURG | 15 approx.
8 | HEAD ADJUSMENT | Yes
9 | STRETCHER SURFACE | Radio translucent
10 | X-RAY CASSETTE HOLDER | Yes, full length
11 | SIDERAILS | Yes
12 | OXYGEN TANK HOLDER | Yes
13 | RESTRAINING STRAPS | Yes
14 | BUMPERS | Yes, all corners
15 | IV POLE SOCKETS | 2 or 4
16 | CASTORS | 4 or 5, lockable
17 | DIAMETER | 20 cm approx.
18 | MAXIMUM PATIENT WEIGHT | Approx 200 Kg
19 | ACCESSORIES | All neceessory accessories
20 | MATTRESS | Yes, durable and easy to clean
21 | PLATFORM DIMENSION IN MM | 2000 X 600 approx.
22 | OTHER SPECIFICATION
1 | AREAS OF USE | Emergency / Re-covery bay
Company Stamp |  |  |  | Company Signature

## ITEM 61
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
61 | 4229 5100 07000 | 0 | SUCTION UNIT HIGH VACUUM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | The pump shall possess (or exceed) the following technical specifications:
3 | Double piston pump continuous rating type.
4 | Pressure guage up to 800 mm Hg
5 | Pressure regulator .
6 | All Castors with brakes .
7 | Over flow protection.
8 | Silicon Tubings.
9 | Suction Bottle Capacity - 2 x 2500 ml minimum (with safety valve)
10 | Bottles graduation scale 100ml - material polysulphone . sterelisable at 134°C .
11 | Vacuum control : ≥ 50 LPM specify
12 | Variable (regulating valve) vacuum pressure from 0 to approximately – 750 mm Hg, with negative pressure gauge indicating actual pressure.
13 | Low noise operation (≤ 45 dB @ 1 meter)
14 | Incorporated filter (specify type)
15 | Foot switch operation .
16 | Hand control .
17 | The offer shall include all the accessories and parts necessary for the full and efficient operation of the system. A detailed list of such standard accessories (with part numbers and quantities) shall be included.
1 | Stable mobile stand made of lightweight, heavy-duty, antirust metal with 4 antistatic swiveling castors and wide base to prevent accidental tripping.
Company Stamp |  |  |  | Company Signature

## ITEM 62
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
62 | 4229 3522 00100 | 0 | SUCTION UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Portable, Capable of connecting to bed
2 | TRACHEAL USE
3 | Pre-Hospital (EMS) | Yes
4 | Intra-Hospital (crash carts) | Yes
5 | PUMP TYPE | Piston
6 | COLLECTION CANISTER | Disposable
7 | Limit line, mL | 800-1000 mL
8 | Capacity measured, mL | 800-1000 mL
9 | Graduations | Yes
10 | Overflow protection | Yes
11 | Connector labels | Yes
12 | Filter Type | Bacterial, HEPA
13 | TUBING
14 | Internal diameter, mm (in) | 8 (0.31)
15 | Length, m (ft) | 1.8 (6)
16 | VACUUM
17 | Range, mm Hg | 50 to 600 mm Hg
18 | Time to 300 mm Hg, sec | <3
19 | Flow rate at max vacuum, L/min | 27
20 | Vacuum Level Control | Regulator, variable
21 | Vacuum Gauge Type | LED indicators
22 | Vacuum Gauge Scale, mmHg | 50 to 600 mm Hg
23 | PHYSICAL SPECIFICATIONS
24 | Dimensions | Specify
25 | Weight | Specify
33 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
34 | IMPORTANT NOTE | Please specify whether accessories are standard or optional, otherwise they will be considered standard
35 | External Charger | Yes/ Specify
36 | Wire Canister bracket | Yes/ Specify
37 | Rechargeable 12 Volts DC battery | Yes/ Specify
38 | AC and DC Power cords | Yes/ Specify
39 | AC/DC Adapter Charger | Yes/ Specify
40 | Disposable Canisters | Yes/ Specify
41 | Soft Pouch carrying case | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 63
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
63 | 4214 0000 14400 | 0 | TABLE EXAM GENERAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Urodynamic Treatment Table/ Electric
2 | APPLICATION | • Urodynamics• Pelvic Muscle Rehabilitation• Anorectal Manometry• Pelvic Exam
3 | TECHNICAL SPECIFICATIONS
4 | • Shall be ergonomics with improved foam to enhance patient comfort and feel | Yes
5 | • Shall be One touch movement locking mechanism for precise positioning | Yes
6 | • Shall be three motors to control lift, reclining back and pelvic tilt | Yes
7 | • Shall be easy access for elderly and patients with physical difﬁculties | Yes
8 | Colour | Standard / Specify
9 | • Inﬁnite height adjustment from | 16” (40 cm) min. to 41” (105 cm) max.
10 | Single foot switch control | Yes
11 | Working Load | ≥ 180 kg
12 | Seat tilt range | 0º - 36º (approx)
13 | Motorized Height & seat &back recling back adjustment | Yes
14 | Lift Mechanism: | Electric
15 | Retractable Wheels | 4
16 | PHYSICAL SPECIFICATIONS
17 | Dimensions | Specify
18 | Cushion Width, cm | Specify
19 | Head Section Length, cm | Specify
20 | Mid Section Length, cm | Specify
21 | Leg Section Length, cm | Specify
22 | Total Length, cm | Specify
23 | Weight | Specify
24 | Standard Accessories
25 | 2- Arm rests 2- Leg rests 2- Leg rest holders 2- Fold up footrests 1- Seat cushion without cut out 1- Cushion with cut out 1- Paper towel holder + maneuvering handle 1- Three motor foot control 1- Funnel & funnel holder 1- Neck pillow | Yes / Specify
Company Stamp |  |  |  | Company Signature

## ITEM 64
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
64 | 4214 0000 07800 | 0 | THERMOMETER DIGITAL HANDHELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Should use disposable probes for measurement
3 | Fast, accurate readings
4 | Temperature accuracy: 0.1°C
5 | Automatic shut-off feature
6 | Large LCD display for reading
7 | Should signal after temperature taken
8 | Safe to use, with no glass and mercury-free
9 | Should have an auto memory to show last 10 temperatures taken
10 | Should include rechargeable batteries with charging desk unit
11 | Unit should operate >1,000 uses on full battery
12 | Unit should have a self test capability
13 | Unit shall have visual indicators
14 | Final temperature stops flashing
15 | low-battery
16 | low and high temperatures
17 | Should have a suitable case
18 | Compliance with standards & legislation:
19 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
20 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
1 | Digital Thermometer safe for all ages and shall be for oral, underarm, or rectal use.
Company Stamp |  |  |  | Company Signature

## ITEM 65
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
65 | 4218 1900 01500 | 0 | THROMBOSIS PREVENTION VIEN DEEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Compression pressure shall be achieved by means of incorporated pump (no external connection to compresses air line shall be necessary) | YES
3 | Features: Specify . | YES
4 | Single Operation Button | YES
5 | One push operation selects | YES
6 | adjusts pressure manually | YES
7 | Monitor/alarm Specify. | YES
8 | One pump for calf and thigh cuffs . | YES
9 | Calf and thigh cuffs available | YES
10 | Pressure and control parameters should be specified: | YES
11 | Pressure regulation method | YES
12 | Approximate applied pressures should be as low as possible between 30 and 60 mm Hg for the calf (varies per body part) and not more than 200 mm Hg as a rule. | YES
13 | Approximate cycle time range of 30 to 60 sec. | YES
14 | Adjustable pressure or pre-set pressure. | YES
15 | Display real-time pressure or indicators for set pressure. | YES
16 | Equipped with high pressure mechanical release valve. | yes
17 | Shall allow one limb or two limb operation. | yes
18 | Equipped with overpressure, leaking, kinked tubing, internal electronic malfunction low pressure and high pressure alarms. | yes
19 |  | yes
20 | Peripheral circulatory assist units designed for sequential compression. | YES
21 | The unit shall display important operating parameters and other alert messages and alarms. Specify display characteristics and list all displayed information. | YES
22 | Specify high pressure safety cutoff points (preferably more than one level, fully independent) and breaking method | YES
23 | Safety features should include: | YES
24 | Splash proof front panel design | YES
25 | Compression halt on alarm | YES
26 | Maintenance-free operation | YES
27 | 1 meter anti kink tubing can be extended for convenient and safe pump positioning. | YES
28 | secure snap lock connections to pump to prevent accidental disconnection. | YES
29 | The unit should be portable, self contained incorporating mounting bracket for IV pole and bed rails | YES
30 | The unit shall be able to connect to all types of garments available on the market (open-system) using universal tubing connectors (luer lock or snap-on) | preferred
31 | Compliance with standards & legislation: | YES
32 | The system must comply with the Electrical safety standards for electrical safety IEC-60601 | YES
33 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any. | YES
34 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements. | YES
35 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions | YES
36 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included. | YES
37 | SLEEVE THIGH SMALL | QTY : 30
38 | SLEEVE THIGH MEDIUM | QTY : 30
39 | SLEEVE THIGH LARGE | QTY : 30
40 | SLEEVE KNEE SMALL | QTY : 30
41 | SLEEVE KNEE MEDIUM | QTY : 100
42 | SLEEVE KNEE LARGE | QTY : 60
43 | PLUG | BRITISH
44 | LANGUAGE | ENGLISH
45 | CERTIFICATES | APPROVED FROM REGULATORY BODY
46 | OTHERS | PROVIDE PROVE OF USE FROM DIFFERENT HOSPITALS LOCALLY OR INTERNATIONALLY
47 | UNIT OF MEASUREMENT | SI : INTERNATIONAL STANDARD
48 | VENDOR MUST PROVIDE ITEMIZED PRICES FOR ALL SLEEVES KIND AND SIZED TO BE SELECTED | MUST
49 | VENDOR MUST PROVIDE FOC MACHINE WITH SUGGESTED NUMBER OF SLEEVES TO BE PURCHASED | YES
1 | Microprocessor controlled operation and safeties | YES
Company Stamp |  |  |  | Company Signature

## ITEM 66
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
66 | 4115 0000 32500 | 0 | TIMER INTERVAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | Electronic interval timer with 4-digit-display for minutes and seconds
3 | Big LCD Display
4 | Count-down-function
5 | Alarm at the end of the set time
6 | Bidder to specify the attachement method
7 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
1 | Digital Timer
Company Stamp |  |  |  | Company Signature

## ITEM 67
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
67 | 4214 0000 08200 | 0 | TROLLEY WITH BASKET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Compact and ergonomic design to securely accommodate the entire equipment. | Yes
2 | Modular structure built with aluminium profiles. | Yes
3 | Push handle . | Yes
4 | Storage basket. | Yes
5 | Two shelves | Yes
6 | Drawer with lock. | Yes
7 | Support for fixation | Yes
8 | Basket for storing . | Yes
9 | Standard rail . | Yes
10 | Foot switch support . | Yes
11 | Dimensions ( W * H * D ) | 630mm x 940mm x 650 mm
12 | Four wheels, two of them with brake. | Yes
13 | Mains power outputs. | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 68
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
68 | 4220 1700 00900 | 0 | ULTRASOUND UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1.1 | (FDA 510 (K) Clearance OR CE Mark) & SFDA Registration | Yes
2 | PRODUCT DESCRIPTION
2.1 | Range | Portable- Laptop, disgn
2.2 | Purpose | Used for general clinical routine applications, in addition to primary health care centres, ICU, ER, OR.
3 | CONSOLE DESIGN
3.1 | Laptop | Yes
3.2 | Total Weight - (incl. battery) | ≤ ( 7 ) kg .
3.3 | Easy Acess mode keys | Yes
3.4 | Screen Monitor | ≥ (15) inch. LCD or LED with High Resolution (Touch Screen is Preferred)
3.5 | Security Lock | Yes
3.6 | Built-In Battery | Yes
3.7 | Battery scanning time | 1 hour at least
3.8 | Power Requirement | 220-240 V / 60 Hz
4 | COMPUTING CAPABILITIES
4.1 | Operation System | Windows or equivalent
4.2 | Integrated Hard Desk | At least ( 100 ) GB (SSD)
4.3 | Fast System boot up | < 50 seconds
5 | CLINICAL PRESETS
5.1 | Abdomen and small parts | Yes
5.2 | OB / Gyn, Breast | Yes
5.3 | Intra Operative & Musculoskeletal | Yes
5.4 | Vascular, Cardiac & Transcranial | Yes
5.5 | Urology | Yes
5.6 | Pediatrics & Neonatal | Yes
6 | MULTI FREQUENCY PROBES
6.1 | Linear array | 5 - 12 MHz (±3)
6.2 | Convex array | 1 - 5 MHz (±2)
6.3 | Active Probe Ports | 1 in the main console
7 | OPERATING MODES
7.1 | Brightness Mode (B-Mode) (2D) | Yes
7.2 | Motion Mode (M-Mode) | Yes
7.3 | Color Doppler mode | Yes
7.4 | Pulsed wave Doppler (PW) | Yes
7.5 | High PRF Doppler Mode | Yes
7.6 | Color Flow Doppler Mode (CFM) | Yes
7.7 | Power Doppler Mode | Yes
7.8 | Continuous Wave Doppler | Yes
8 | DOPPLER
8.1 | Type | pulsed wave in all related probes
8.2 | Frequency Display | Yes
8.3 | Doppler baseline shift after freeze | Yes
8.4 | Velocity Display | Yes
8.5 | User – defined annotation | Yes
8.6 | GrayScale Levels | Yes
8.7 | Transmit focal zones | Select between 1 to 4
8.8 | Digital reconstructed zoom | 16 levels
8.9 | Adaptive Gain Compensation | Yes
8.1 | Wall filters (clutter filter) in colour Doppler and PW Doppler | Yes
9 | CINE FEATURES
9.1 | Cine Storage length | Minimum 1 Sec
9.2 | Capacity | Specify
10 | SCANNING PARAMETERS
10.1 | Imaging Depth, cm | Not less than 28 cm
10.2 | Dynamic Range | Up to 150 dB
10.3 | TGC Control | Yes
10.4 | Frame Rate 2D | Up to 350 fps
10.5 | Frame rate Color Doppler | Up to 100 fps
10.6 | GrayScale Levels | 256 ( 8 bits )
10.7 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
10.8 | Digital Calipers | Yes
10.9 | Preprocessing | Yes
10.1 | Postprocessing | Yes, (Measurment on stored images)
10.11 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
11 | SCANNING FEATURES
11.1 | Speckle Reduction Imaging | Yes, Organ Specific
11.2 | Trapeziodal Imaging / Virtual Convex | Yes
11.3 | B-Mode Steering | Yes
11.4 | Compound Imaging | Yes
11.5 | Extended field of view imaging (Panoramic Imaging) | Yes, Min 30 cm
11.6 | tissue Harmonic Imaging | Yes
11.7 | TGC (Time Gain Compensation) | Yes
73 | Automatic Doppler Calculations | Yes
74 | Adjustable Transmit Focus | Yes
75 | Dynamic Receive Focus | Yes
76 | Selectable Dynamic Range | Yes
77 | Vascular access or needle guide software package | Optional (vascular and interventional applications)
78 | ZOOM / PAN
79 | Real- time image | Yes
80 | Frozen image | Yes
81 | Split Screen | Yes
82 | NETWORK & SECURITY
83 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicaple
84 | Wireless Connection | Yes
85 | Full DICOM Connectivity | (Verify, Print, Store, Modality work list, multi frame, storage commitment, Modality Performed Procedure Step, MPPS, Media Exchange). Query / Retrieve, Structured Reporting
86 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
87 | Storage and Playback of M-mode, CW and PW Spectral Doppler. | yes
88 | Full Connectivity to USB Printers | Yes
89 | Printing of Patient Reports Via a USB Connection. | Yes
90 | Recalculation for stored images | Yes
91 | Possibility to Export on different Formats | Raw data, PC formats, DICOM Viewer
92 | HIPAA Patient Data Security | yes
93 | DATA PROTECTION AGAINST DELETION & export | Different privilege levels required
94 | Ultrsound System security Virus protection Or equivilant | Yes
95 | ACCESSORIES
96 | Secondary Battery | Yes
97 | Carrying Case | Yes
98 | Cart should include the following: Integrated Power Supply, Probe Ports, wheels with locking braking system. | Yes
99 | Add-on specifications on top of base specs | Vendor input
100 | Phased array; 2 - 5 MHz (±2) for echocardiography, transcranial imaging, trauma (FAST)
101 | Cardiac Analysis packages (scanning, stress echo, IMT, intracardiac (ICE), Auto EF ) ) including increase rate to more than 1000 fps
102 | Upgradability to cardiac 4D (Live 3D) Echo Software and hardware
103 | Upgradability to cardiac 4D (Live 3D) TEE Software and hardware adult and pediatric
104 | Anatomical Intelligence
105 | xMATRIX array
106 | Please specify Add-on specifications on top of base specs
1 | GENERAL STANDARD
Company Stamp |  |  |  | Company Signature

## ITEM 69
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
69 | 4218 1900 08800 | 0 | VENTILATOR ICU ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | ICU Adult, Pediatric and Neonatal Patients
2 | CONTROLS | Microprocessor
3 | - Tidal volume, mL | 2ml - 2500 ml
4 | - Inspiration pressure, cm H20 | 0 - 100
5 | - Ventilation rate, Breaths / min | 0- 150
6 | - Insp flow L/min | 0 - 180
7 | - Insp time. Sec | 0.1 - 10 Sec Increments of 0.1
8 | - Exp time. Sec | 0.1 - 30 Sec
9 | - Insp Hold / plateau | Yes
10 | - Expiratory Hold | Yes
11 | - I : E Ratio | Specify
12 | - Fi02. percent | 21 - 100% Increments of 1%
13 | - Manual Oxygen Breath | For 2 min
14 | - Manual breath | Yes
15 | - PEEP / CPAP. cm H20 | 0 - 40
16 | - Sensitivity range, cm H20 / L / min | 0 - 35
17 | - Trigger mechanism
18 | Flow | Yes
19 | Pressure | Yes
20 | Biosignal Feedback | Preferable
21 | - Leak Test | Yes
22 | - Leak Compensation | 100% invasive & 200% non-invasive
23 | - Manual Calibration | Yes
24 | - Apnea time | 5-60 Sec Increments of 5Sec
25 | Neonatal Option:
26 | - Flow Sensor | Yes 20 with Each Machine
27 | - Flow Sensor Cable | Yes 2 Each Machine
28 | Diagnostic Procedure
29 | - NIF | up to -80 cm H20 / L / min
30 | - Auto PEEP/ PEEPi | 40 cm H20
31 | - RSBI | 0- 200
32 | - P 0.1 | Yes
33 | - Static Compliance | Yes
34 | Self Diagnostics | Yes
35 | Capnograph | Yes
36 | Transesophegal Pressure Monitoring | Yes/ Specify
37 | Heliox Compatible | Yes/ Specify
38 | Reusable Expiratory cassette | Yes 2 Each Machine
39 | Reusable Flow Sensor | Yes 4 Each Machine
40 | OPERATING MODES
41 | - Pressure - controlled | Yes
42 | - Assist / control (Volume / Pressure) | Yes/ Specify
43 | - IMV / SIMV (Volume / Pressure) | Yes/ Specify
44 | - Volume -controlled | Yes
45 | - MMV | Yes
46 | - APRV / Bivent / Bi level | Yes
47 | - Pressure support | 0-60 cm H2O
48 | - Pressure Regulated | Yes
49 | - Combind mode (Assist + Spontaneous) | Yes
50 | - Automode | Yes
51 | - ATC | 0-100 %
52 | - CPAP | Yes
53 | - NIV | Yes
54 | - NIV / Nasal CPAP | Yes
55 | - NIV / Pressure Control | Yes
56 | - NIV / Pressure Support | Yes
57 | - High flow Oxygen Therapy | Yes
58 | - Apnea-backup vent | Yes
59 | - Proportional Support | Yes
60 | Preferably to have a biosignal triggering mechanism to assist the breathing of patient in synchrony with his own effort | Specify
61 | All consumables and disposables required by the techniques specified above shall be quoted in itemized price | Yes/Specify
62 | HUMIDIFICATION SYSTEM
63 | - Compatible Humidifier system with KFMC ventilator circuits | Yes
64 | - Temperature Prop | Yes
65 | - Temperature Cable Compatible with KFMC Disposable Ventilator Circuit | Yes
66 | - 10 pieces Chamber per machine | Yes
67 | - Alarm for High & Low Temp | Yes
68 | Durable Trolley | Yes
69 | INTERFACING
70 | - Output ports | Yes RS 232/USB
71 | - Remote alarm/ display | Yes
72 | - Report generation | Yes / SIMV
73 | - Display | Yes or Via RS232/USB
74 | - Hard Copy | Yes or Via RS232/USB
75 | - Archival disk | Yes or Via RS232/USB
76 | - Network | Yes or Via RS232/USB
77 | - Able to save the set and measurement number | yes for 48 hour
78 | ALARMS AND DISPLAY
79 | Alarms
80 | Volume
81 | - Low & High VT | Yes
82 | - Low & High MV | Yes
83 | Pressure
84 | - High PIP | Yes
85 | - Low & High PEEP | Yes
86 | - Pressure Limited | Yes
87 | Respiratory Rate
88 | - High RR | 15-100 BPM
89 | - Apnea | Yes
90 | FIO2
91 | - Low FIO2 | 5% below the set
92 | - High FIO2 | 5 % above the set
93 | - Gas supply failure | Yes
94 | - Power failure | Yes
95 | - Vent inoperative | Yes
96 | - Low battery | Yes
97 | Display Type | Specify
98 | - Screen size | 15-17 inch (Touch screen)
99 | Data Displayed | Digital Number , Graphics waveform, Loops and Trends
100 | Other Alarms | PEEP / Patient trig. / I:E ratio / Inv.I:E ratio / Spontaneous
101 | PHYSICAL SPECIFICATIONS
102 | Dimensions | Specify
103 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 70
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
70 | 4217 0000 02900 | 0 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To Transport Critical Care Adult, Pediatric and Neonatal Patients
2 | CONTROLS | Microprocessor
3 | - Tidal volume, mL | 219179
4 | - Inspiratory pressure, cmH20 | 0 - 100
5 | - Ventilation rate, Breaths / min | 0 - 150
6 | - Inspiratory flow, lpm | 1 - 180
7 | - Insp time. Sec | 0.1 - 10 with increments of 0.1
8 | - Exp time. Sec | 0.1 - 30
9 | - Insp Hold / plateau | Yes
10 | - Expiratory Hold | Yes
11 | - IE ratio | Specify
12 | - Fi02, % | 21 - 100 with Increments of 1
13 | - Manual Oxygen Breath | For 2 min
14 | - Manual breath | Yes
15 | - PEEP / CPAP. cmH20 | 0 - 35
16 | - Sensitivity range, cmH20 / L / min | 0 - 35
17 | - Trigger mechanism | flow / pressure
18 | - Capnography (adult, pediatric, neonate) | Yes
19 | - Leak Test | Yes
20 | - Leak Compensation | 100% invasive & 200% non-invasive
21 | - Manual Calibration | Yes
22 | - Apnea time, secs | 5-60 with of Increments of 5
23 | Neonatal Option:
24 | - Flow Sensor | Yes with 20 Pcs Each Machine
25 | - Flow Sensor Cable | Yes 2 Each Machine
26 | - Nasal CPAP | Yes
27 | Patient Transport Capability | Yes
28 | OPERATING MODES
29 | - Pressure - controlled | Yes
30 | - Assist / control ( Volume / Pressure) | Yes / Specify
31 | - IMV / SIMV ( Volume / Pressure) | Yes / Specify
32 | - Volume -controlled | Yes
33 | - Pressure support, cmH2O | 0 - 60
34 | - Pressure Regulated | Yes
35 | - Combination modes (Assist + Spontaneous) | Specify
36 | - CPAP | Yes
37 | - APRV /Bilevel mode | Yes
38 | - NIV | Yes
39 | - Apnea-backup vent | Yes
40 | DISPLAY
41 | Type | Specify
42 | Data Displayed | Digital Number , Graphics waveform, Loops and Trends
43 | INTERFACING
44 | - Output ports | Yes RS 232/USB
45 | - Remote alarm/ display | Yes
46 | - Report generation | Yes / SIMV
47 | - Display | Yes or Via RS232/USB
48 | - Hard Copy | Yes or Via RS232/USB
49 | - Archival disk | Yes or Via RS232/USB
50 | - Network | Yes or Via RS232/USB
51 | ALARMS
52 | Volume
53 | - Low & High VT | Yes
54 | - Low & High MV | Yes
55 | Pressure
56 | - High PIP | Yes
57 | - Low & High PEEP | Yes
58 | - Pressure Limited | Yes
59 | Respiratory Rate
60 | - High RR, bpm | 15 - 100
61 | - Apnea | Yes
62 | FIO2
63 | - Low FIO2 | 5% below the set
64 | - High FIO2 | 5 % above the set
65 | - Gas supply failure | Yes
66 | - Power failure | Yes
67 | - Vent inoperative | Yes
68 | - Low battery | Yes
69 | Self Diagnostics | Yes
70 | Other Alarms | PEEP / Patient trig. / I:E ratio / Inv.I:E ratio / Spontaneous
71 | Expiratory cassette Reusable (2 backup for each machine) | Yes
72 | PHYSICAL SPECIFICATIONS
73 | Dimensions | Specify
74 | Weight | Specify
75 | Disposables | Disposable Transport circuit with full accessories (Exhalation valves)
76 | Durable Trolley | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 71
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
71 | 4219 1800 03100 | 0 | WARMER BABY BOTTLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | heating element . | Specify
3 | Elliptical shape, compact design . | Specify
4 | Thermostatic controlling and automatic | Yes
5 | thermal cutoff . | Yes
6 | Thermostat keeps constant temperature . | Yes
7 | Infection Control . | Yes
8 | heating plate & removable bracket, easy for cleaning . | Yes
9 | No. of bottle .(4) | Yes
10 | Temperature range . | Specify
11 | Power Supply 220 VAC / 60 Hz | Yes
1 | Efficient, safe, economic, durable PTC | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 72
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
72 | 4229 3500 02700 | 0 | WARMER BLOOD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile, IV pole mounted
2 | MAXIMUM TEMPERATURE SETTING, °C | Up to 42°C
3 | TECHNICAL SPECIFICATIONS
4 | Heat Exchanger
5 | - Aluminum heat exchanger | Yes
6 | - Technology | Specify
7 | - Maximum Temperature Setting, °C | Up to 42°C
8 | - Increments, °C | 1
9 | - Warm-up time, mins | < 2
10 | Maximum Flow, mL/min, 10°c input to ≥ 37°c output | ≥ 150
11 | Priming Volume, mL | 20 - 150
12 | Display | LCD
13 | - Temperature range, °C | 33 - 42
14 | - Increments, °C | 1
15 | Air Vent or Trap | Manual bubble trap on standard flow sets; Autoventing bubble trap on all high-flow sets.
16 | APPLICATIONS
17 | - IV | Yes
18 | - Heater power, W | 800-850 W (Approx)
19 | - High Temperature Cutoff, °C | 43 °C
20 | - Automatic air elimination on all high flow sets | Yes
21 | - Electronic temperature control | Yes
22 | - Audible, visual & alarm test High Temperature alarms | Yes
23 | PHYSICAL SPECIFICATIONS
24 | Dimensions | Specify
25 | Weight | Specify
Company Stamp |  |  |  | Company Signature

## ITEM 73
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
73 | 4215 2200 15700 | 0 | WASTE BIN CHEMICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
2 | The bin shall have scratch, corrosion and abrasion resistant qualities
3 | It shall have a yellow color and marked with the international infectious substance symbol.
4 | Waste container of approx. 60 to 80 L capacity
5 | Pedal activated operation is preferred, to minimize the risk of cross infection
6 | It shall have a frame constructed from hot dipped coated galvanized material to prevent corrosion
7 | The unit shall be fire safe
1 | This unit should be designed for pharmaceutical and chemical waste collection
Company Stamp |  |  |  | Company Signature

## ITEM 74
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
74 | 4219 2210 06500 | 0 | WHEELCHAIR STANDARD SIZE 20 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Aluminium wheelchair | Yes
2 | Powder coating finish black | Yes
3 | Foldable backrest, aluminium folding mechanism | Yes
4 | Flip back armrest | Yes
5 | Detachable footrest, heel loop | Yes
6 | Nylon black upholstery | Yes
7 | Front castor 8"x1" solid tire | Yes
8 | Steel brake | Yes
9 | Rear wheel 24" x 1 3/8" PU tire , , | Yes
10 | Quick release rear axle | Yes
11 | aluminium hand rim | Yes
12 | brakes | Yes
13 | Anti tipper | Yes
14 | 5 Cm seat cushion | Yes
15 | Capacity 120 Kg | Yes
16 | Wieght 14 Kg | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 75
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
75 | 4220 1800 01300 | 0 | X-RAY DIGITAL MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1.1 | (FDA 510 (K) Clearance OR CE Mark) & SFDA Registration | Yes
2 | MOBILE X-RAY UNIT
2.1 | Microprocessor Controlled | Yes
2.2 | Forward & Reverse Motor Driven | Yes, Self Propelled Handle Drive Control
2.3 | Max incline, degree | ( 5-7 ) Degree is Preferable
2.4 | Speed | Adjustable from ( 0.5 - 1.3 ) m/s
2.5 | AC line | Yes
2.6 | Line-voltage Compensator | Yes / Automatic
2.7 | Type of Power Sources | Both AC & Battery, The Generator Should Work on AC in Case of Battery Failure (Preferable)
2.8 | Power Requirements | 220 Volt / 60 Hz.
2.9 | Single or Dual Battery System | Two Independent Systems Provide Power for Driving Motors & Imaging Exposure (Preferable)
2.1 | Battery Type | Lead acid sealed lead / NiCad
2.11 | Battery Charging Time - ( From Empty to Full Charged ) | Not more than ( 10 ) hrs
2.12 | Low-Battery Indicator | Yes
2.13 | Battery Power Storage | Specify number of maximum oprating hours
2.14 | Power Cord Length, m | Not less than ( 2.5-5 ) m or More
2.15 | Wired Handheld Switch & Cordless Remote control Switch | Yes, Exposure By Using Remote Control for Optimal Radiation Protection
2.16 | Manual Movement in Case of Battery or Motor Failure | Yes
2.17 | Remote Control Operating range | Not Less than ( 7 ) m
2.18 | Anti-Collision Protection System or Equivalent | Yes
2.19 | System Weight | Specify
2.2 | Mobile X-ray can be Moved Forward & Backward by using the Bedside Drive Controls Located on the Collimator or on the Articulated Arm | Preferred
3 | X-RAY TUBE
3.1 | Rotating Anode | Yes
3.2 | Anode Heat Storage Capacity | Not Less Than ( 120 ) kHU
3.3 | Maximum output Voltage, kVp | Approx. ( 140 ) KV , ( ±10 )
3.4 | Nominal Focal Spot Value, mm | ( 0.8 / IEC ) or Dual Focal Spots
3.5 | Nominal Anode Input power | Not Less Than ( 30 ) KW
3.6 | Tube Movement & Angulation | Yes
3.7 | Clear view while moving (telescopic movement of the column or dropdown) | Yes
3.8 | Shortest Exposure Time | 2 ms or less
3.9 | Working Column Movement Range | Rotating ( 180 ) Degree
4 | X-RAY GENERATOR
4.1 | High Frequency X-Ray Generator | Yes
4.2 | Nominal power Rating | Not Less Than ( 30 ) KW
4.3 | Kv Range | Approx. ( 40 - 140 ) Kv ( ±10 ) step of 5 (as minimum)
4.4 | mA range | Not Less Than ( 300 ) mA
4.5 | mAs range | Min.: ≤ 0.4 , Max.: ≥ 320
4.6 | Increments | Step Up / Step Down / Continuous
4.7 | Digital Indicator meters | Yes, (If applicable) For KVp, mAs, mA, Dose
5 | X-RAY COLLIMATOR
5.1 | Type | Manual Collimator
5.2 | Rotation | Approx. ( ±90 ) Degree
5.3 | Aluminum filter, mm | ≤ (3.5) mm
5.4 | SID Range, cm | Min.: ≤ 68 , Max.: ≥ 200
5.5 | Centering indicator (light /laser) | Yes, for Quick & Easy Indication of SID
6 | LARGE DETECTOR ASSEMBLY
6.1 | Waterproof | Yes
6.2 | Shock Sensor | Yes
6.3 | Portable Digital Detector | Wireless
6.4 | Semiconductor material | Amorphous silicon ( A-Si )
6.5 | Scintillator | Cesium iodide ( Csl )
6.6 | Detector Active Area Size | ≥ ( 35 x 43 ) cm, Approx.
6.7 | Image Depth / A/D Conversion / Acquisition Depth | ( 14 ) Bit
6.8 | Pixel size | 150 Micron or less is Preferred
6.9 | Matrix Size / Resolution | Not Less Than ( 2022 X 2022 )
6.1 | Detective Quantum Efficiency ( DQE ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
6.11 | Modulation Transfer Function ( MTF ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
6.12 | Detector Weight Including Battery | Less Than or equal ( 4 ) kg
6.13 | Max. Load Capacity On Lying Position | More Than ( 100 ) kg
6.14 | Rechargeable Battery (Built in or Removable) | Lithium-ion , in Case of Removable Second Battery is Required or equivalent
6.15 | Charging Time | Not More Than ( 4 ) hours for 100% Charge
6.16 | Battery Operation Time | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
6.17 | Charging Station | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
6.18 | Rigidity of Detector Housing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
6.19 | Detector Handle | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
6.2 | Detector Sharing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at (90) images per hour
7 | PROCESING & VIEWING WORKSTATION (Operator Console)
7.1 | Acquisition Workstation | Built in
7.2 | LCD Touch Screen Monitor | Not Less Than ( 15 ) inch
7.3 | Radiation Dose Monitoring ( DAP ) | indicate
7.4 | Had Disk Storage | Not Less Than ( 3000 ) images
7.5 | Automatic Programming ( APR ) | Yes
8 | IMAGE POST PROCESSING FUNCTIONS
8.1 | Window Width & Leveling | Yes
8.2 | Gray Scale Invert / Annotation | Yes
8.3 | Image Rotate / Free Rotation | Yes
8.4 | Electronic L/R Marker | Yes
8.5 | Patient Edit / Emergency Exam | Yes
9 | NETWORKING & SECURITY
9.1 | DICOM Compatible | Yes
9.2 | DICOM Store and DICOM structured dose report | Yes
9.3 | Image auto transfer | Yes
9.4 | Query, retrieve, Send, Receive | Yes
9.5 | Modality worklist | Yes
9.6 | Storage commitment | Yes
9.7 | DICOM print | Yes
9.8 | Patient edit | Yes
9.9 | DICOM MPPS | Yes
9.1 | DICOM viewer on CD "Burn exam on CD with DICOM viewer" | Yes
9.11 | HIPAA Patient Data Security | Yes
9.12 | DATA PROTECTION AGAINST DELETION & export "Different privilege levels " | Yes
9.13 | Security Package | Yes
9.14 | Virus Protection | Yes
9.15 | The ability to interact with the processing workstation from remote PC to view and process studies within the same hospital | Yes
10 | ACCESSORIES
10.1 | Two light weight Aprons (0.5) Equivalent of ( Non Leaded Material ) | ( Two ) - (medium & large)
10.2 | One Moveable Detector holder | ( One ) with Vertical, Horizontal & Swivel Movement
11 | Add-on specifications on top of base specs | Vendor input
11.1 | Extra detector (small)
11.2 | Extra detector (large)
11.3 | Above specs with telescopic column movement for X-Ray tube (if available)
11.4 | Extra detector (small) with exchange price
11.5 | Extra detector (large) with exchange price
11.6 | One Mobile Apron Hanger, and must hold 5 aprons
11.7 | Please specify Add-on specifications on top of base specs
1 | GENERAL STANDARD
Company Stamp |  |  |  | Company Signature