## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | ALARM & MONITOR RADIATION AREA DIGITAL
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | ANALYZER CO2
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | CABINET AUTOMATED MEDICATION MOBILE
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | CABINET DRYING & STORAGE FLEXIBLE ENDOSCOPE
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | CABINET DRYING ENDOSCOPY
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | CABINET UNIT DOSE PET LEAD LINED
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | CALIBRATOR DOSE 55T PET
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | COMPRESSOR AIR HOME CARE
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | CONCENTRATOR OXYGEN ADULT 10L
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | CONCENTRATOR OXYGEN ADULT 5L
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | DEFIBRILLATOR UNIT
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | DISINFECTOR ENDOSCOPY
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | ECG 12 LEAD WITH CART
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | ELECTROSURGERY WITH ARGON & VESSEL SEALING
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | ENDOSCOPY WORKSTATION
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | FLOWMETER AIR
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | HUMIDIFIER
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | IRRIGATION SYSTEM SURGCAL
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | LIFT PATIENT BARIATRIC
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | LIFT PATIENT MOBILE
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | METER SURVEY RADIATION
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | MONITOR PHYSIOLOGICAL TRANSPORT
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | MONITOR RADIATION SYSTEM EXTERNAL PROBE
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | NEBULIZER AIR
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | OXIMETER PULSE FINGER TIP
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | PUMP DUTY CHEMICAL
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | PUMP INFUSION
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | PUMP INFUSION EPIDURAL
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | PUMP INSULINE
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | PUMP PCA
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | PUMP SYRINGE
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | SCAN DEVICE BLADDER
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | STRETCHER EMERGENCY
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | SUCTION UNIT PORTABLE
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | TRANSILLUMINATOR VEIN PORTABLE
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | VASCULAR DOPPLER HAND HELD
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | VENTILATOR BIPAP AUTO HOME CARE
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | VENTILATOR CPAP AUTO HOME CARE
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | WHEELCHAIR LIGHT WEIGHT SIZE 20

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4220 1800 01800 | ALARM & MONITOR RADIATION AREA DIGITAL | 1 | MEDICAL GROUP | 1 | [10 empty] | 0
2 | 4111 3119 00000 | ANALYZER CO2 |  | MEDICAL GROUP | 4 | [10 empty] | 0
3 | 4219 1909 00400 | CABINET AUTOMATED MEDICATION MOBILE |  | MEDICAL GROUP | 1 | [10 empty] | 0
4 | 4229 5100 48100 | CABINET DRYING & STORAGE FLEXIBLE ENDOSCOPE | 2 | MEDICAL GROUP | 7 | [10 empty] | 0
5 | 4229 6300 00300 | CABINET TEE PROBES STORAGE |  | MEDICAL GROUP | 1 | [10 empty] | 0
6 | 4220 1800 03100 | CABINET UNIT DOSE PET LEAD LINED | 1 | MEDICAL GROUP | 1 | [10 empty] | 0
7 | 4220 3700 01700 | CALIBRATOR DOSE 55T PET | 1 | MEDICAL GROUP | 1 | [10 empty] | 0
8 | 4219 1704 32200 | COMPRESSOR AIR HOME CARE |  | MEDICAL GROUP | 80 | [10 empty] | 0
9 | 4227 1702 00700 | CONCENTRATOR OXYGEN ADULT 10L |  | MEDICAL GROUP | 25 | [10 empty] | 0
10 | 4227 1702 00100 | CONCENTRATOR OXYGEN ADULT 5L |  | MEDICAL GROUP | 565 | [10 empty] | 0
11 | 4218 1900 11900 | DEFIBRILLATOR UNIT |  | MEDICAL GROUP | 4 | [10 empty] | 0
12 | 4229 6300 00500 | DISINFECTOR ENDOSCOPY | 2 | MEDICAL GROUP | 2 | [10 empty] | 0
13 | 4218 1900 01800 | ECG 12 LEAD WITH CART |  | MEDICAL GROUP | 7 | [10 empty] | 0
14 | 4229 5104 01500 | ELECTROSURGERY WITH ARGON & VESSEL SEALING | 3 | MEDICAL GROUP | 2 | [10 empty] | 0
15 | 4115 0000 35600 | ENDOSCOPY WORKSTATION |  | MEDICAL GROUP | 1 | [10 empty] | 0
16 | 4214 0000 02900 | FLOWMETER AIR | 4 | MEDICAL GROUP | 50 | [10 empty] | 0
17 | 4214 0000 03200 | FLOWMETER O2 15 L/Min WALL MOUNTED | 4 | MEDICAL GROUP | 300 | [10 empty] | 0
18 | 4214 0000 03700 | HUMIDIFIER |  | MEDICAL GROUP | 20 | [10 empty] | 0
19 | 4229 3500 04000 | IRRIGATION SYSTEM SURGCAL | 3 | MEDICAL GROUP | 2 | [10 empty] | 0
20 | 4219 1800 01800 | LIFT PATIENT BARIATRIC |  | MEDICAL GROUP | 2 | [10 empty] | 0
21 | 4219 1800 02000 | LIFT PATIENT MOBILE |  | MEDICAL GROUP | 2 | [10 empty] | 0
22 | 4220 3900 00700 | METER SURVEY RADIATION | 1 | MEDICAL GROUP | 2 | [10 empty] | 0
23 | 4218 2401 00400 | MONITOR PHYSIOLOGICAL TRANSPORT |  | MEDICAL GROUP | 2 | [10 empty] | 0
24 | 4220 1800 55500 | MONITOR RADIATION SYSTEM EXTERNAL PROBE | 1 | MEDICAL GROUP | 1 | [10 empty] | 0
25 | 4227 1802 02300 | NEBULIZER AIR |  | MEDICAL GROUP | 200 | [10 empty] | 0
26 | 4217 0000 01700 | OXIMETER PULSE FINGER TIP |  | MEDICAL GROUP | 600 | [10 empty] | 0
27 | 4228 1700 00400 | PUMP DUTY CHEMICAL |  | NON MEDICAL | 1 | [10 empty] | 0
28 | 4218 1900 06500 | PUMP INFUSION | 5 | MEDICAL GROUP | 4500 | [10 empty] | 0
29 | 4218 1900 16100 | PUMP INFUSION EPIDURAL | 5 | MEDICAL GROUP | 100 | [10 empty] | 0
30 | 4218 1900 06900 | PUMP INSULINE |  | MEDICAL GROUP | 25 | [10 empty] | 0
31 | 4229 3500 02000 | PUMP PCA | 5 | MEDICAL GROUP | 200 | [10 empty] | 0
32 | 4218 1900 07000 | PUMP SYRINGE | 5 | MEDICAL GROUP | 1350 | [10 empty] | 0
33 | 4218 1900 07600 | SCAN DEVICE BLADDER |  | MEDICAL GROUP | 5 | [10 empty] | 0
34 | 4214 0000 06500 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | MEDICAL GROUP | 14 | [10 empty] | 0
35 | 4618 1810 01100 | STRETCHER EMERGENCY |  | MEDICAL GROUP | 8 | [10 empty] | 0
36 | 4229 3522 00100 | SUCTION UNIT PORTABLE |  | MEDICAL GROUP | 220 | [10 empty] | 0
37 | 4218 1900 08100 | TRANSILLUMINATOR VEIN PORTABLE |  | MEDICAL GROUP | 6 | [10 empty] | 0
38 | 4229 5100 10700 | VASCULAR DOPPLER HAND HELD |  | MEDICAL GROUP | 4 | [10 empty] | 0
39 | 4227 2208 01000 | VENTILATOR BIPAP HOME CARE |  | MEDICAL GROUP | 95 | [10 empty] | 0
40 | 4227 2208 01100 | VENTILATOR CPAP AUTO HOME CARE |  | MEDICAL GROUP | 260 | [10 empty] | 0
41 | 4219 2210 03400 | WHEELCHAIR LIGHT WEIGHT SIZE 20 |  | MEDICAL GROUP | 100 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | MEDICAL GROUP | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
|  | NON MEDICAL | 0 | 0 | [5 empty] | If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4220 1800 01800 | 1 | ALARM & MONITOR RADIATION AREA DIGITAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MONITOR DIGITAL AREA 000.0-9999 WITH DETECT FAIL ALARM FOR GM GAMMA RADIATION DETECTION (QTY 1):
2 | Monitor should provides a fast response to pulsed fields by utilizing an external ion chamber covering any five consecutive decade-range between 0.1 mR/hr and 1000 R/hr. It should featured a wall-mount chassis and a four digit LED display that is readable from 9 meters (30 feet) away. | Yes
3 | DIGITAL SURVEY METER WITH GM PANCAKE PROBE (QTY 2):
4 | Measure X-ray, Beta & Gamma Radiation | Yes
5 | LCD Display | Yes
6 | Rechargeable batteries | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 4111 3119 00000 | 0 | ANALYZER CO2 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CO2 Analyzer Incubator Analyzer monitors CO2 for verifying incubators.
2 | Measures CO2 from 0-20%
3 | CO2 is measured by a custom dual wavelength infra-red sensor with reference channel.
4 | Quick verification of CO2 incubator levels
5 | Built in gas moisture removal
6 | Greatly improved accuracy and response time
7 | Lightweight and has an easy to read LCD display with LED backlight when needed.
8 | The internal pump has a flowrate of roughly 100 cc/min and features a built-in gas dryer tube that removes moisture.
9 | has ports for two temperature probes.
10 | Battery Charger
11 | Sample Tube Kit
Company Stamp |  |  |  | Company Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 4219 1909 00400 | 0 | CABINET AUTOMATED MEDICATION MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MANUFACTURER | Please specify
2 | Model | Please specify
3 | Specification | automated decentralized medication management systems used to dispense drugs and medical supplies in patient care areas (e.g., nursing floors). provide ready access to medications in patient care areas and to control drug distribution tory
4 | Drug/Supply Storage Configurations | Please Explain
5 | Types of dispensing | Multi drugs/Multi dose, Single drug/ Multi dose, etc.
6 | Types dispensed as a single unit dose | Oral Solids, Liquids, Ampules, Syringes, Vials, Large & Small volume, etc.
7 | Security Features | Biometric ID, optical/magnetic card, user ID/password, pin/pin encryption, and etc.
8 | Additional Security Options | Preferred
9 | Barcode reader | Required
10 | User Interaction in dispensing unit | Touch screen, mouse, keyboard
11 | Screen Size, Type in dispending unit | Required, built in PC (all in one)
12 | Operating system | Windows 10
13 | HIS interface | Required
14 | System Compatibility | The offered system must be fully compatible with current KFSH&RC automated decentralized medication management systems
Company Stamp |  |  |  | Company Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 4229 5100 48100 | 2 | CABINET DRYING & STORAGE FLEXIBLE ENDOSCOPE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope storage capacity | Min. 10 scopes, Vertical position.
2 | Dry cycle completion time. | Max. 3 hours
3 | Sustain the bacteriological status of the endoscope after washing in an AEWS for a period of | At least 30 days
4 | Simple integrated data management with the Endoscope washer software (device and consumables via the main screen or admin pc). | Required
5 | Means of storing data. | Required and specify
6 | Rechargeable wireless barcode readers | Required
7 | Detect and alarm any disconnection to endoscope | Required
8 | Fit all flexible endoscopes currently located in KFSHRC (as per attached scope list). | Required
9 | H X W X D, Cm (In) | Please Specify
10 | Run on medical Air or built-in compressor | Required and specify
11 | Compatible with the pass through tracking system | Required and specify
12 | Full compliance with relevant parts of BS EN 16442 | Required
13 | Dual filtered air maintains positive cabinet pressure | Required
14 | HEPA filtered air delivered to each scope channel | Required
15 | Clear toughened glass doors allow easy scope recognition | Required
16 | Temperature and humidity (for each channel) monitoring within the cabinet | Required
17 | Secure mounting for scope control section and light guide plug | Required and specify
18 | Air changes in cabinet 95% recirculated | Required and specify
19 | Air cleanliness in cabinet > ISO Class 7 | Required and specify
20 | Cabinet temperature monitoring | Required
21 | Filter condition indicator (Filter Blockage) | Required
22 | Door open alarm Door securely locked - access restricted | Required
23 | Air flow to each endoscope monitored | Required
24 | Traceability - all data stored and can be connected to the endoscopy network | Required and specify
Company Stamp |  |  |  | Company Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 4229 6300 00300 | 0 | CABINET TEE PROBES STORAGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | High Level Disinfection by UV-c | Yes
2 | Time Required | 180sec
3 | Chemical free | Yes
4 | Effective | Bactericide,Viricide,Fungicide, Mycobactericide & Sporicide.
5 | Automatic | Yes
6 | Size | Cylindrical Enclosure
7 | Diameter | 55 cm
8 | Height | 202 cm
9 | Weight | 110 kg
10 | Power | 1430 Watt
11 | Supply Voltage | 230V @60HZ
12 | Intensity | 6.2 A
13 | NO CONSUMABLES | Yes
14 | Report print electronic version | Yes
15 | Wheels / Can be shifted easily | Yes
16 | Environmental Friendly | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 4220 1800 03100 | 1 | CABINET UNIT DOSE PET LEAD LINED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LEAD-LINED PET Unit Dose Cabinet 0.5 inch Lead:
2 | -Dimensions (40”H x 30”W x 24”D (101.6x76.2x60.9cm) | Yes
3 | -Lead Shielding: 0.5” thick | Yes
4 | -Finish: Stainless steel | Yes
5 | -Countertop: Stainless steel | Yes
6 | -Countersunk calibrator chamber mount | Yes
7 | -Waste Penetration fits shielded sharps container | Yes
8 | -Supports 511 L-Blocks | Yes
9 | TABLE TOP SHIELD 511 L-BLOCK WITH 4" LEAD GLASS:
10 | -Dimensions: 20 ¼”h x 14” d (51 x 35.5 x 35.5 cm) | Yes
11 | -Lead Glass Density: 8” x 8” x 4” (20.32 x 20.32 x 10.16cm) | Yes
12 | - Weight: 195 lbs. (89 kg) | Yes
13 | - The fixed angle design of the lead glass allows easier viewing | Yes
14 | - Contains two sheets of ¼” window glass to protect lead glass from scratching and chips (industry exclusive) | Yes
15 | - Lead shielding thickness: 1.0” (2.54 cm) | Yes
16 | BRICK KIT FOR USE WITH CABINET PET UNIT DOSE AND TABLE TOP SHIELD 511 L-BLOCK:
17 | - Lead Brick Enclosure: (5 cm) lead | Yes
18 | - Dimensions: 12” H x 28” W x 21 7/8” D (30.5 x 71.1 x 55.6 cm) outer | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 4220 3700 01700 | 1 | CALIBRATOR DOSE 55T PET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DOSE CALIBRATOR PET (QTY 1):
2 | Dose Calibrator includes:
3 | Ionization Chamber | Yes
4 | USB/PC Communications | Yes
5 | Built in Software | Yes
6 | Roll Printer | Yes
7 | Ticket Printer | Yes
8 | Display Touch 8 inch | Yes
9 | Voltage input 220V AC with British plug | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 4219 1704 32200 | 0 | COMPRESSOR AIR HOME CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Noise Level | Less than 54 decibels
4 | Power Consumption | 150 Watts max or less
5 | Patient category | Adult, pediatric, neonatal
6 | Oxygen bleed-in | Up to 10 LPM, integrated FiO2 monitor preferred
7 | Power requirements | 100-240V; 50/60Hz, Otherwise 220V
8 | Plug Standard | British type
9 | Flow Range | 0-60 LPM, increments for adult and ped
10 | Max Pressure | 0-50 PSI, preferred
11 | temperature output | ≥ 32°
12 | accessories | 2 humidifier champers, universal humidifier adaptors, per unit, 2 patient setups
Company Stamp |  |  |  | Company Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | 4227 1702 00700 | 0 | CONCENTRATOR OXYGEN ADULT 10L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Oxygen Concentration | 1-3 LPM 95% ±3% - 4 LPM 92% ±3% - 5 LPM 90% ±3%
4 | Maximum Oxygen Flow | 8-10LPM or better
5 | Noise Level | Less than 54 decibels
6 | Power Consumption | 350 Watts max or less
7 | Humidifier Bottle Support | yes
8 | Filter System | HEPA or Equivalent
9 | Air Cooling System
10 | Flow Meter Accuracy | +- 0.5LPM
11 | Audible/Visual alarm: | Low/High Pressure, Low Concentration, Temperature Alarm and Power Failure
12 | Power requirements | 100-240V; 50/60Hz, Otherwise 220V
13 | Plug Standard | British type
14 | ADDITIONAL FILTER | YES
15 | Backup/Emergency | Each concentrator has 1 cylinder E size (light to carry) with integrated regulator (PREFERRED) MANDATORY
Company Stamp |  |  |  | Company Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | 4227 1702 00100 | 0 | CONCENTRATOR OXYGEN ADULT 5L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Oxygen Concentration | 1-3 LPM 95% ±3% - 4 LPM 92% ±3% - 5 LPM 90% ±3%
| Flow Range | Adjustable 1 LPM increments for adult, 0.25-0.5 LPM increments micro flowmeter for Neonates
4 | Maximum Oxygen Flow | 5 LPM or better
5 | Noise Level | Less than 54 decibels
6 | Power Consumption | 350 Watts max or less
7 | Humidifier Bottle Support | YES
8 | Filter System | HEPA or Equivalent
10 | Flow Meter Accuracy | +- 0.5LPM
11 | Audible/Visual alarm: | Low/High Pressure, Low Concentration, Temperature Alarm and Power Failure
12 | Power requirements | 100-240V; 50/60Hz, Otherwise 220V
13 | Plug Standard | British type
14 | Backup/Emergency | Each concentrator has 1 cylinder E size (light to carry) with integrated regulator (PREFERRED) MANDATORY
15 | ADDITIONAL FILTER | YES
Company Stamp |  |  |  | Company Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | 4218 1900 11900 | 0 | DEFIBRILLATOR UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONTROL / SETTING RANGES
2 | AED And Manual Defibrillation
3 | Internal Energy Selection, J | 5-50
4 | External Energy Selection, J | 50 – 360 Biphasic
5 | FiO2, % | 21 - 100
6 | Synchronizer
7 | Lead Selection | Lead I, II, III, Other optional
8 | Pacing mode | Demand, Fixed Rate
9 | Pacing Rate, PPM | 20 - 150
10 | Pacing Output Current, mA | 0 - 140
11 | Pulse Width, mSec | More than 20
12 | Paddle Controls | Charge, Discharge, Energy Selection
13 | Heart rate selection | ECG leads / SpO2
14 | Automatic NIBP timing, Hours | 0.5 - 12
15 | MONITORING PARAMETERS
16 | Delivered Energy
17 | ECG Waveform
18 | Heart Rate
19 | SpO2
20 | NIBP, Systolic
21 | NIBP, Diastolic
22 | End Tidal CO2
23 | Sync Marker
24 | Sweep Speed, mm/Sec – Configurable
25 | Lead Configuration, I,II,III, optional 5 and 12 Leads
26 | Trace Freeze
27 | External Pacemaker
28 | Annotation, Time, Date, Lead, Gain, Heart Rate, Operating Mode
29 | Discharge Waveform Shape | Specify
30 | HARDWARE SPECIFICATION
31 | Pediatric Paddle
32 | Adult Paddle
33 | Disposable Electrodes
34 | Discharge Time, mSec Less than 20 mSec
35 | Display, Color LCD TFT
36 | Display Size, More than 5 Inches
37 | Low / High Frequency Pressure
38 | ECG monitoring Through Paddles
39 | ECG Leads Type, Snapped
40 | ECG Recorder
41 | Auto Manual Print
42 | Battery Slots, Two | Preferred
43 | Battery Type – Lithium Ion, Nickel Metal Hydride
44 | Battery Backup Time – More than 60 Shocks @ 360 J
45 | Hot Swappable Battery
46 | Charging Time to Maximum Energy on Battery, Less than 10 S | Preferred
47 | Charging Time to Maximum Energy on Electricity, Less than 10 S
48 | Dedicated External Battery charger | Preferred
49 | Defibrillator’s ECG input Amplifier Protection
50 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
51 | Altitude : - 300m to 3000m from the sea level
52 | Dimension | Specify
53 | Temperature : 15 to 45 C
54 | Cart must contain at least four wheel in which two wheels must have built-in brakes, Accessories drawer, Hanger for leads
55 | ELECTRICAL SPECIFICATION
56 | Defibrillator must have dual voltage operational compatibility at automatic selection of frequency.
57 | Voltage Range : 100 ~ 240 Volts
58 | Frequency : 50Hz / 60Hz
59 | Maximum Current : ≤ 5 Amp
60 | Battery backup time must be 1 hrs. at-least
61 | ALARMS
62 | Low Battery Alarm
63 | Heart Rate Alarm
64 | Systolic Pressure Alarm
65 | Diastolic Pressure Alarm
66 | Oxygen Saturation Alarm, High /low
67 | EtCO2 Alarm, High/LOW
68 | Lead Disconnection Alarm
69 | Arrhythmia alarm with Interpretation
70 | Equipment malfunction alarm with code
71 | Battery Replacement / Maintenance Alarm | Preferred
72 | CONNECTIVITY
73 | Ethernet / LAN
74 | USB | Preferred
75 | WIFI Connectivity
76 | GPRS / 3G / 4G Connectivity (Paramedics/Ambulance Only) | Preferred
Company Stamp |  |  |  | Company Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | 4229 6300 00500 | 2 | DISINFECTOR ENDOSCOPY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | automated endoscope washing system (AEWSS)
2 | Technical Specification | Parameters
3 | H X W X D, Cm (In) | Please Specify
4 | Weight, Kg (Lb.) | Please Specify
5 | Ground Clearance, Cm (In) | Please Specify
6 | Made and assembled in which country? | Please Specify
7 | Double Chamber With Asynchronous Operation Maximizing Scope Availability And Turnaround Time. | Required
8 | Possibility of accessing the machine remotely (RAS) for trouble shooting, repair and upgrade. | Required
9 | connection to the scope | Easy (One click if possible).
10 | channel monitoring of the endoscope and blockage detection during washing cycles | Required and Specify
11 | scope leak test | Required and Specify
12 | AEWS should be Hands free ergonomic pass through design | Required
13 | Scope Compatibility: Compatible with all flexible endoscopes including ultrasound, TEE probes, TOE probes from all makes and sizes (as per attached scopes list). | Required
14 | Extensive cycle data management to be stored on both the machine and the server | Required and Specify
15 | cycle time for all types and features | Less than 40 minutes
16 | Number of washing cycle to provide consistent thorough soil and residue removal | twice or more
17 | No manual handling of scopes (hands free) in order to minimize accidental cross contamination and have a better and safer infection prevention | Required
18 | Pass-Thru System intended for the cleaning and high-level disinfection of Endoscopes | Required and Specify
19 | Provide high level disinfection of heat sensitive semi-critical endoscopes and related accessories if needed in conjunction with AAMI, 2010 standard. | Required
20 | Test Strips to be used after the disinfection cycle to ensure that the used disinfectant is at or above the minimum recommended concentration. | Required
21 | Operating Temperatures | Ambient 25 to 35 ℃
22 | The endoscope hookups are automatically disinfected during the cycle and are provided free of charge for all kinds of scopes available at KFSH (as per scope attached list) at the time of the tender release (three sets each) | Required
23 | Excess chemicals to be vented through special ventilation opening | Required
24 | System to run on Reverse Osmosis water | Required
25 | Detergent solution to be low-foaming and formulated for use with medical instruments and safe and will not harm the scopes in any way including shape or form in conjunction with AAMI, 2010 standard. | Required
26 | Thorough flushing of all endoscope channels | Required
27 | See through washing window to enable the user to see the scope being washed and the other side of the decontamination area IE clean area staff can clearly see the staff in the dirty area | Required
28 | Touchscreen Monitor, (Rechargeable), Foot Switch to Allow hands free door opening and closing, built-in UPS to safe grade the cycle data and database | Required
29 | Alcohol Reservoir container to be available upon the request of the hospital | Optional
30 | Overflow Drain in the washing basin. | Required
31 | Sample Port Overflow | Required
32 | Minimum Basin Level Sensor indicator | Required and specify
33 | Basin Lights to see the scope easily | Required
34 | Wireless Barcode Scanner : to scan operator barcode tag and to read the scope barcode | Required
35 | Print Service automatically prints a certificate of the duty cycle see more details below in disinfection program. | Required
36 | Status light indicators to indicate the status of the scope high level disinfection (HLD) status | Required
37 | Type of cycles: Leak Test Phase, Washing Phase, Disinfectant Phase, Rinse Phase, Alcohol Purge Phase (Optional), Air Purge Phase, Disinfection Cycle, Cycle Completion indicator. | Required and Specify
38 | Delayed Start up | Required
39 | Run Data Recorded (Database) directly to the server after cycle ends. | Required
40 | 1. For each disinfection program run, the following data must be at least recorded and stored on the computer’s hard drive: | Required
41 | 1. Machine serial number.
42 | 2. Basin No.
43 | 3. Run number.
44 | 4. Date and time.
45 | 5. Hookups used.
46 | 6. Cycle result.
47 | 7. Disinfectant temperature.
48 | 8. Disinfectant contact time.
49 | 9. Program run type.
50 | 10. Operator Load ID.
51 | 11. Operator Unload ID.
52 | 12. Endoscope model.
53 | 13. Endoscope serial number.
54 | 14. Patient ID.
55 | 15. Physician.
56 | 16. Assistant.
57 | 17. Process step times.
58 | 18. Errors.
59 | 19. Disinfectant MRC Pass/Fail.
60 | 20. Patient MRN or full name or both.
61 | Automated windows opening hands free, Mechanical (Manual) as backup | Required
62 | Comprehensive online trouble shooting guide | Required and Specify
63 | Secured administrator menu | Required
64 | Traceability software to be submitted free of charge | Required and specify
65 | Ability to add or remove users and scopes from the database | Required
66 | Computer storage to be redundant | Required and specify
67 | Ability to connect the washer to an external central server | Required and specify
68 | AEWS Air pressure | 6 bar
69 | AEWS Water pressure | 2 bar
70 | AEWS Water flow | 2 to 6 L/Min
71 | AEWS Water temperature operating range | 30 C + 10 C
72 | The system will come with three trays (Basins), accessories and connectors for every scope that KFSH&RC Riyadh has currently (as per attached scope list). | Required
73 | SOFTWARE SPECIFICATIONS (FREE OF CHARGE) COMPATIBLE WITH AEWS:
74 | Technical Specification | Parameters
75 | · Appointments planning: | Required
76 | a. Compatible with KFSH Cerner Millennium HIS.
77 | b. Planning views per room / per examiner.
78 | c. Graphical display of room availability.
79 | d. Capacity dashboard.
80 | e. Flag for examinations with anesthetist requirement.
81 | f. Patient information, medical information and history from the Hospital HIS.
82 | g. Examination types with preset materials and nomenclature.
83 | h. Admission and hospitalization management.
84 | i. Reprocessing:
85 | j. Workflow support with multimedia support and alerts.
86 | k. Shelf life management.
87 | l. Detailed information about selected endoscope.
88 | m. Custom procedures and periodic actions per endoscope type.
89 | n. Register all bacteriological and TOSI tests.
90 | o. View endoscopes status and locations in one tiled overview.
91 | p. Repair management.
92 | q. Repair approval flow support.
93 | Possibility of adding pictures or videos to help the new staff on a designated Step
94 | · Examination: | Required
95 | a. Linking the endoscope ID to the examination.
96 | b. Linking the Patient identity to the examination.
97 | c. Linking Medical staff identities to the examination.
98 | d. Patient Flow Monitoring.
99 | e. Registering instrument sets and consumed materials.
100 | f. Invoice nomenclature and consumed materials.
101 | · Reporting capabilities: | Required
102 | a. Scope usage/availability.
103 | b. Possibility to calculate the cost per procedure.
104 | c. Detailed daily action about workers.
105 | Detailed cycle data of all AERs and cabinets.
Company Stamp |  |  |  | Company Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | 4218 1900 01800 | 0 | ECG 12 LEAD WITH CART |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Instrument type | 12 Leads or more
2 | Channels | Multiple Channel, 10 Channel or more
3 | Automatic ECG feature
4 | Analysis Frequency | 500 Samples/S
5 | CMRR @ 60 Hz, dB | More than 90
6 | Frequency range, Hz: | Diagnostic 0.67-150
7 | Sensitivity, mm/mV: | 2.5, 5, 10, 20
8 | CONFIGURATION | Mounted On Mobile Cart
9 | Calibration signal: | Yes
10 | Filtered: | line frequency, low and muscle artifact/high frequency, Notch Filter
11 | Input impedance, megohms: | 100 or higher (Meets or exceeds ANSI/AAMI)
12 | Leads-off indicator: | Yes
13 | Integral printer | Yes
14 | Lead marker: | Yes
15 | Timing marker: | Yes
16 | Event marker: | Yes
17 | Chart speed, mm/sec: | 25, 50
18 | Number of Channels acquired simultaneously: | 12
19 | Number Channels printed simultaneously: | 3
20 | ECG TRANSMISSION: | Yes
21 | Transmission Type | Wireless LAN, Standard
22 | MICROPROCESSOR Programs: | Adult, pediatric, 12-lead, late potential, P-wave signal averaging
23 | Keyboard | Yes
24 | Storage: | PCMCIA, Flash Media or others for not less than 150 ECG
25 | PATIENT DATA: | Name, ID no., age, sex, height
26 | MEASUREMENTS: | PR, PQ, QT, ATC, P, QRS, T, HR;(if others, pls. specify)
27 | Arrhythmia ID: | Yes
28 | INTERPRETATION | Yes
29 | Analyzed waveform: | Yes
30 | VISUAL INDICATORS | Loose contact, lead fail, low battery, system status, artifacts
31 | DEFIBRILLATOR OVERLOAD PROTECTION: | Yes
32 | BATTERY OPERATION: | Prefered, Rechargeable
33 | Operating time on Battery, hr: | more than 1
34 | Multiple recording formats | Yes
35 | Transmits/receives data to/from ECG cart and data management system. | Yes
36 | 1 min full-size full disclosure for 1-lead printout | Yes
37 | 5 min half half-size full disclosure for 3-lead printouts; | Yes
38 | Bar-code reader. | Optional
39 | Diagnostic Software | Yes
40 | Display Size | 10" orMore
41 | Display type | Color LCD TFT
42 | Writer type | Thermal Dotarray
43 | Writer Resolution | 1000dpi
44 | Paper Type | Thermal Z fold
45 | Keyboard Type | Sealed Elastomer, IPX7
46 | Battery Backup | 2 Hrs Minimum
47 | Battery Type | Nickel Metal Hydride or Lithium Ion
48 | Integration With HIS
49 | Connectivity | WIFI, ETHERNET
Company Stamp |  |  |  | Company Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | 4229 5104 01500 | 3 | ELECTROSURGERY WITH ARGON & VESSEL SEALING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Output Configuration | Bipolar / Monopolar
2 | Display | Touchscreen
3 | Duty Cycle | 0.25
4 | Audible Alarms | <65 dB
5 | Visual Activation Indicator | Yes
6 | Volume Control | Yes
7 | Mounting | Cart
8 | Input Ratings | 100-240 , 50/60 Hz
9 | Max Power Output
10 | Cut at rated load | 400 Watts
11 | Coag at rated load | 360Watts
12 | Multi Optimized Modes | Yes
13 | Bipolar Resection in a saline solution | Yes
14 | Rapid Vessel sealing | Yes
15 | Low-smoke exposure | Yes
16 | Smooth cutting with min to moderate hemostasis | Yes
17 | Controlled cutting with hemostasis | Yes
18 | Fractionated Cutting with CUT and COAG interval | Yes
19 | Preset Programs | Yes
20 | Expanded Instrument Selection up to 4 | Yes
21 | Universal receptacle | Yes
22 | Easy receptacle exchange | Yes
23 | Simultaneous Coagulation | Yes
24 | Advance tissue Measurement Technology | Yes
25 | Argon Plasma Coagulator used with Electrosurgical Generator
26 | General Specification
27 | Country of Origin | USA
28 | CE Mark | Yes
29 | Output | Monopolar
30 | Duty cycle | 0.25
31 | Alarms | <65 dB
32 | Technical Specification
33 | Input Ratings | 100-240 , 50/60 Hz
34 | Argon Gas System
35 | Gas Supply | Argon
36 | Flow Range L/min | 0.1-8
37 | Filter pore size | 0.2µm
38 | Tank Capacity (cu.ft.) | 40
39 | Low Pressure Alarm | Yes
40 | Output
41 | Power Output | 1-120 Watts
42 | Automatic mode | 2-10 L/min
43 | Manual mode | 4-10 L/min
44 | Footswitch | Yes
45 | Handswitch | Yes
46 | Modes
47 | Argon Plasma Coagulation | Yes
48 | Coagulation and Devitalization | Yes
49 | Homogenous Coagulation | Yes
50 | Pulsed Activation for Coagulation | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | 4115 0000 35600 | 0 | ENDOSCOPY WORKSTATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MONITOR LCD HD 26":
2 | Screen Size | 26 Inch
3 | Display Device | TFT Active Matrix
4 | Resolution | 1920 x 1080 Full HD
5 | Contrast Ration | 1400:01
6 | Aspect Ratio | 16:09
7 | Viewing Angle | Ø 170 Degree
8 | Number of colors | 1 Billion
9 | Dimensions | >626 mm W, >395 mm H, >76 mm D
10 | Weight | < 8 Kg
11 | Power | Auto 100 to 240 V
12 | Inputs | SDI (2), Y/C, HD 15, DVI (2) Aux BNC
13 | Output | SDI (2), Video BNC connector with loop through output and auto 75 Ohm termination, Y/C, HD 15, DVI (2), Clone out 3G/HDSDI BNC
14 | Remote | RS 232C Connector D-Sub 9-pin, Parallel Remote RJ-45
15 | Picture Control | PIP, POP and clone out
16 | Earth, Grounding | Through power cable
17 | VIDEO CAMERA GASTROSCOPE:
18 | Dual Focus
19 | less halation and noise
20 | Narrow Band Imaging
21 | Help doctors differentiate contrast between blood
22 | Vessels and mucosa
23 | Field of view, Normal 140º, Near 140º
24 | Direction of view, Forward viewing
25 | Depth of field, Normal 5-100 mm, Near 2-6 mm
26 | Distal end outer diameter = 9.9 mm
27 | Insertion tube outer diameter 9.9 mm
28 | Working length = 1030 mm
29 | Channel inner diameter 2.8 mm
30 | Minimum visible distance 3.0 mm from the distal end
31 | Bending Section, Up 210º, Down 90º, Right 100º, Left 100º
32 | Total Length = 1350 mm
33 | Water proof one-touch connector
34 | Water Jet nozzle
35 | VIDEO COLONOSCOPE 170° FIELD OF VIEW:
36 | Dual Focus
37 | less halation and noise
38 | Narrow Band Imaging
39 | Help doctors differentiate contrast between blood vessels
40 | and mucosa
41 | Field of view, Normal 170º, Near 160º
42 | Direction of view, Forward viewing
43 | Depth of field, Normal 5-100 mm, Near 2-6 mm
44 | Distal end outer diameter = 13.2 mm
45 | Insertion tube outer diameter 12.8 mm
46 | Working length = L1680 mm I:1330
47 | Channel inner diameter 3.7 mm
48 | Minimum visible distance 4 mm from the distal end
49 | Bending Section, Up 180º, Down 180º, Right 160º,
50 | Left 160º
51 | Total Length = L: 2005 mm I:1655
52 | Water proof one-touch connector
53 | Water Jet nozzle
54 | VIDEO DUODENOSCOPE WITH DISPOSABLE DISTAL COVER:
55 | Square image shape and 15° backward viewing
56 | Facilitates fast and secure short guidewire locking with dual
57 | system at the distal end
58 | Enables a 1:1 transfer of pushing and rotating forces to the
59 | distal end of the Duodenoscope
60 | Narrow Band Imaging
61 | Single-Use Distal Cover
62 | Distal End Flushing Adapter
63 | Waterproof One-Touch Connector
64 | Field of view | 100 Degree
65 | Direction of view | Backwards 15 Degree
66 | Depth of Field | 5.0-60 mm
67 | Distal end outer diameter | 13.5 mm
68 | Insertion tube outer diameter | 11.3mm
69 | Working Length | 1,240 mm
70 | Channel inner diameter | 4.2 mm
71 | Maximum visible distance | 10 mm from the
72 | distal end
73 | Angulation of bending section, up 120, down 90,
74 | right 110, left 90
75 | Total Length 1,560mm
76 | 1.5 VIDEO PROCESSOR:
77 | NBI, Dual Focus,
78 | waterproof one touch connector to enable one touch connection
79 | superior image quality with modern image processing
80 | 16:9 and 16:10 output for an HDTV output monitor
81 | Compatible with analog and HD-SDI and DVI output
82 | Flex ray connection
83 | PIP
84 | Portable memory, Support DV output
85 | Power 100 – 240 60 Hz
86 | Size W370, H 90, D490mm
87 | Weight | < 11 Kg
88 | Protection class I
89 | White balance adjustment
90 | Color tone adjustment
91 | Automatic gain control
92 | Contrast, Normal, High, Low
93 | Iris, adjusted using the iris mode switch
94 | Fog free function
95 | Displaying the image information
96 | Memory backup, 5 Years
Company Stamp |  |  |  | Company Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | 4214 0000 02900 | 4 | FLOWMETER AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Air Flow | 0 – 15 LPM (Flowmeter only)
4 | Needle Valve | (Flowmeter Only)
5 | Gas Probes | As per hospital Standard
6 | IPX7 | YES
Company Stamp |  |  |  | Company Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | 4214 0000 03200 | 4 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | O2 Flow | 0 – 15 LPM (Flowmeter Only)
4 | Needle Valve | (Flowmeter Only)
5 | Gas Probes | As per hospital Standard
6 | IPX7 | YES
Company Stamp |  |  |  | Company Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | 4214 0000 03700 | 0 | HUMIDIFIER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Temperature Select | Chamber, Limb temp
4 | Temperature Range | 37 C to 42 C
5 | Over temperature safety | YES
6 | Flow Measurement | YES
7 | Disposal Chamber | YES
8 | Reusable Temperature sensor | YES
9 | Limb Heater Wire Support | YES
10 | Temperature Auto adjustment | YES
11 | Audible/Visual alarm: | Low/High Temperature, malfunction with Technical Codes
12 | Power requirements | 100-240V; 50/60Hz, Otherwise 220V
13 | Plug Standard | British type
14 | Temperature Display | YES
15 | IPX1 | YES
16 | TEMPERATURE PROBE 1.5M | YES / Optional Price
17 | ADULT VENTILATOR CIRCUIT HEATED | YES / Optional Price
Company Stamp |  |  |  | Company Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | 4229 3500 04000 | 3 | IRRIGATION SYSTEM SURGCAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Line Voltage | 220-240Vac, 50Hz
2 | Type of Pump | Peristaltic Roll Pump
3 | Operation | Continuous
4 | Activation | Footswitch or Control Panel
5 | Max Flow Rate | 500 ml/min
6 | Pre-set values | Yes, 30, 50 and 80%
7 | Irrigation Adapter for Olympus | Yes
8 | Irrigation Adapter for Fujifilm | Yes
9 | Irrigation Adapter for PENTAX | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | 4219 1800 01800 | 0 | LIFT PATIENT BARIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA certificate | yes
2 | FDA CLEARANCE | yes
3 | CE MARK (MDD) | yes
4 | CAPACITY, kg (lb) (Size and weight capacity must be sufficient to accommodate individual patients based on patient demographics) | more than 300 (661)
5 | Displayed on lift | Preferred
6 | Lifting mechanism | Battery powerd
7 | Boom, H x L, cm (in) | please specify
8 | Patient support | please specify (sling , rigid or contoured stretcher,Vest.)
9 | Material | Washable/sanitizable, please specify
10 | Suspension | 4-point, or stretcher,please spcify
11 | Lift range, lowest position, cm (in) | please specify
12 | Lift range, highest position, cm (in) | please specify
13 | Base size, max open | please specify
14 | Base size, closed | please specify
15 | H x W x L, cm (in) | please specify
16 | Underbed clearance needed, cm (in) | please specify
17 | Minimum storage space needed | please specify
18 | H x W x L, cm (in) | please specify
19 | FRAME MATERIAL, please specify | please specify
20 | SCALE FUNCTION | Preferred
21 | CASTERS | 6
22 | Size, cm (in) | 10 (4),please specify
23 | Brakes | required
24 | SAFE FOR BATHROOM USE (Lifts are often used to assist in positioning patients on commodes and in bathtubs) | please specify
25 | WEIGHT, kg (lb) | please specify
26 | BATTERY
26_1 | Type (number) | at least 2, with charging unit, please specify
26_2 | Life, hr | please specify
26_3 | Rechargeable | please specify
26_4 | Recharging time, hr | please specify
26_5 | Low-battery notice | please specify
27 | Manual back-up | required
28 | SWITCH TYPE | Remote and fixed push button,please specify
29 | CONTROLS
29_1 | Up/down | required
29_2 | Right/left | required
Company Stamp |  |  |  | Company Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | 4219 1800 02000 | 0 | LIFT PATIENT MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA certificate | yes
2 | FDA CLEARANCE | yes
3 | CE MARK (MDD) | yes
4 | CAPACITY, kg (lb) (Size and weight capacity must be sufficient to accommodate individual patients based on patient demographics) | more than 158.7 (350)
5 | Displayed on lift | Preferred
6 | Lifting mechanism | Battery powerd
7 | Boom, H x L, cm (in) | please specify
8 | Patient support | please specify (sling , rigid or contoured stretcher,Vest.)
9 | Material | Washable/sanitizable, please specify
10 | Suspension | 4-point, or stretcher,please spcify
11 | Lift range, lowest position, cm (in) | please specify
12 | Lift range, highest position, cm (in) | please specify
13 | Base size, max open | please specify
14 | Base size, closed | please specify
15 | H x W x L, cm (in) | please specify
16 | Underbed clearance needed, cm (in) | please specify
17 | Minimum storage space needed | please specify
18 | H x W x L, cm (in) | please specify
19 | FRAME MATERIAL, please specify | please specify
20 | scale function | prefered,
21 | CASTERS | 4
23 | Size, cm (in) | ≥8 (3),please specify
24 | Brakes | ≥1,please specify
25 | SAFE FOR BATHROOM USE (Lifts are often used to assist in positioning patients on commodes and in bathtubs) | please specify
26 | WEIGHT, kg (lb) | please specify
27 | BATTERY
27_1 | Type (number) | at least 2, please specify
27_2 | Life, hr | please specify
27_3 | Rechargeable | please specify
27_4 | Recharging time, hr | please specify
27_5 | Low-battery notice | please specify
33 | Manual back-up | required
34 | SWITCH TYPE | Remote and fixed push button,please specify
35 | CONTROLS
35_1 | Up/down | required
35_2 | Right/left | required
Company Stamp |  |  |  | Company Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | 4220 3900 00700 | 1 | METER SURVEY RADIATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
22 | DOSE CALIBRATOR REFERENCE SOURCE GE-68 :
23 | Dose Calibrator Reference Sources provide a safe and convenient method of calibrating instruments for measuring the accuracy of imaging solutions most commonly utilized by medical technicians. The Dose Calibrator Vial is manufactured by distributing the active element in 20 ml. of epoxy, comprising a density of approximately 1.0 g/cm3. Each standard is supplied in a 27 ml polyethylene vial. | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | 4218 2401 00400 | 0 | MONITOR PHYSIOLOGICAL TRANSPORT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Intended For High-Acuity Patient Monitors | Required
2 | PATIENT TYPE | Adult, pediatric, neonate
3 | MONITORING PARAMETER
4 | Required Parameters | ECG, SpO2, NIBP, respiration, temperature, IBP, ETCO2
5 | Optional Parameters | BIS, SvO2, tcpO2, tcpCO2
6 | ECG Number of leads | 12 ECG system or better
7 | ECG HR range, bpm | 30-250
8 | HR Accuracy | <5% or 5 bpm
9 | 12 Lead ECG Disclosure | Required
10 | Interpretation | Yes, with 12 SL 12-Lead ECG
11 | Arrythmia No. of leads analyzed | 4 Leads
12 | Arrhythmia Sampling rate: | at least 250 sample/second
13 | ST analysis | Yes
14 | ST No. of leads analyzed | 12 Leads
15 | Auto lead switch | 1 mV calibration (Preferred)
16 | Respiration Breath Rate, bpm | 0 - 30
17 | Respiration Method | ECG Impedance or etCO2
18 | Respiration Waveform displayed | Yes
19 | Respiration Threshold control | Yes
20 | IBP | Yes
21 | IBP No. of channels | 1 or More
22 | IBP Scales range, mm Hg | -30 to +340
23 | IBP Labels | Yes
| NIBP | Yes
| NIBP Measurement technique | Oscillometric
| NIBP Pressure Range | At least 20-300 mmHg
| NIBP Cuff size | adult, pediatric, neonate, thigh
| NIBP Hose connection | Any; Luer not recommended (please specify)
| PULSE OXIMETRY | Yes
| PULSE OXIMETRY Probe type | Nellcor, Masimo
| PULSE OXIMETRY Disposable/reusable | Reusable
| TEMPERATURE | Yes
| TEMPERATURE, No. of channels | 1 channels
| TEMPERATURE Probe type | Oral, axillary, rectal
| ETCO2 | Yes
| ETCO2 accuracy | ± 5%
ALARMS
| HR Alarms | Yes
| IBP Alarms | Yes
| Arrhythmia detect | Full Arrythmia
| Respiration Rate Alarm | Yes
| Systolic Pressure | Yes
| Diastolic Pressure | Yes
| Low SpO2 Saturation | Yes
| ETCO2 Alarm | Yes
| ALARMS Type | HR/PR, respiration, NIBP, SpO2, lethal arrhythmia
| Visual ALARM | Yes
| Audible ALARM | Yes
| TRENDING | Yes
| TRENDING Parameters | All
| TRENDING, Graphical/tabular | Both
| Trending Length of time, hr | Min. 72
| Retrieve Patient data After Discharging | Yes
| Silence active alarm | Yes
| ALARM Silence duration, sec | 60 or 120
| All-alarm suspend | Yes
| All-alarm suspend Silence duration, sec | userconfigured 60 or 120
| Alarm Volume Retriction feature | Yes
| Alarm Suspend Restriction | Yes
| NETWORK CONNECTIVITY - INTERFACE | Yes
| Interfaced to other devices, | ventilators, anesthesia units, intravenous pumps, pulse oximeters, and multiple medical-gas monitoring systems.
| NETWORKING | Yes
| NETWORKING, Configuration | Wireless or Hardwired
| NETWORKING Communication | Wireless , Ethernet, TCP/IP
| Network Architecture | Recommended Ethernet (if others, please specify)
| Max no. of devices on single network | up to 1024
| Networking Other clinical information systems | Hospital Information System with HL7 compliance
| Peripheral device interface | Yes
| DISPLAY SPECIFICATION | Yes
| Display Type | LCD, Medical Grade
| DISPLAY Size, (in) | (8) or more
| Display Traces | 8, Configurable
| Display USER INTERFACE | Trim Knob, remote control, touchscreen
| security management system | Yes
| PERIPHERALS | Yes
| Full disclosure of ECG waveforms. | At least 72 hours
| Recorder/Printer | Yes, Thermal and Laser
| Recorder Channels | Single
| Auto Print | Detect arrhythmias
| HARDWARE SPECIFICATION | Yes
| Battery | Yes
| Battery Slots | One or more
| Battery Type | Lithium Ion or Nickel Metal Hydride
| Battery Back up | Min. 3 hr
| Battery Life | 500 Cycles or 3 Years
| Monitor configuration | Transport
Monitor must have dual voltage operational compatibility at automatic selection of frequency.
| Voltage Range : 100 ~ 240 Volts | Yes
| Frequency : 50Hz / 60Hz | Yes
| Maximum Current : ≤ 5 Amp | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | 4220 1800 55500 | 1 | MONITOR RADIATION SYSTEM EXTERNAL PROBE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MONITOR RADIATION SYSTEM WITH ONE EXTERNAL PROBE:
2 | Energy Compensated Geiger Muller detector. | Yes
3 | Measuring range: 0.1 -40 micro Sv | Yes
4 | User friendly touch screen | Yes
5 | Ethernet/Wireless Connection | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | 4227 1802 02300 | 0 | NEBULIZER AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Noise Level | Less than 54 decibels
4 | Power Consumption | 150 Watts max or less
5 | Ultrasonic Nebulization
6 | Power requirements | 100-240V; 50/60Hz, Otherwise 220V
7 | Plug Standard | British type
8 | Particle Size | 0.5 to 5 Micron
9 | Max Pressure | 30 PSI
10 | Nebulizer Capacity | 5 – 8 ml
11 | Continuous Flow | 0 – 15 LPM
Company Stamp |  |  |  | Company Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | 4217 0000 01700 | 0 | OXIMETER PULSE FINGER TIP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class III Equipment, Type BF
3 | Parameters | SPO2 , Pulse Rate
4 | Saturation O2 | 40% to 100%
5 | SpO2 Accuracy | ±1 units or ±1%
6 | Pulse Rate Ranges | 30 BPM to 300 BPM
7 | PR Accuracy | ±2 BMP or ±1%
8 | Operating Temp. | 5 C to 45 C
9 | Display | Color OLED/TFT screen
10 | Turn Off Auto | 10 Sec without operation, Auto shutdown
11 | Audible/Visual alarm: | Low/High SPO2, Low/High Pulse Rate, Disconnection, Low battery
12 | Technical Code | yes
13 | Power requirements | Battery Operated 1.5V AAA
14 | Power Consumption | Less Than 30mA
15 | Battery Life | 48 Hrs continuous
16 | Non Invasive measurement method | yes
Company Stamp |  |  |  | Company Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | 4228 1700 00400 | 0 | PUMP DUTY CHEMICAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Portable Vacuum/Pressure Pump
2 | Chemical Duty
3 | 220 V/50-60 Hz
4 | 24 in. Hg vacuum
5 | 4 L/min Flow Rate Pressure
6 | For use with corrosive chemicals and solvents
7 | Tolerates solvent vapor and some moisture
8 | Easy-to-adjust vacuum and pressure regulators with individual gauges
9 | Includes filter unit to protect pump from liquid carryover
10 | UL and CE marked
11 | Includes silicone tubing
Company Stamp |  |  |  | Company Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | 4218 1900 06500 | 5 | PUMP INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEARANCE | PREFFERED
2 | CE MARK (MDD) | REQUIRED
3 | APPLICATION | blood products, epidurals, general care, ICU, IV solutions
4 | TIVA-/TCI-capable | REQUIRED
5 | CONFIGURATION
6 | Number of channels | EQUAL OR MORE THAN 1
7 | Pump mechanism | Stepper motor, lead screw/peristaltic
8 | Pole mounting | REQUIRED / 5wheels
9 | Max units per pole | MORE THAN 1
10 | DISPLAY | LCD,screen size data easy to read
11 | Data displayed | Alarms, drug name, pumping status, volume infused, VTBI
12 | PUMP CAPABILITIES
13 | SYRINGE-SIZE DETECTION | 3 to 50 or more
14 | Flow range, mL/hr | 0.1 to ≥999
15 | Increments, mL | 0.1-100 mL/hr
16 | KVO rate, mL/hr | 2020-01-05
17 | Accuracy, % | 5
18 | VTBI, mL | 1-9,999
19 | Automatic piggybacking | Required
20 | Bolus mode | Required
21 | Syringe delivery | Required
22 | MRI conditional | Preferred
23 | Fluid resistant | Required
24 | Front-panel lockout | Required
25 | Free-flow protection | Required
26 | Air-trapping capability, volume | Required
27 | Needleless IV connection | Required
28 | SYRINGE-SIZE DETECTION | Required
29 | ALARMS & INDICATORS | Required
30 | High pressure/occlusion | Required
31 | System malfunction | Required
32 | Empty syringe | Required
33 | Plunger disengaged | Required
34 | Syringe unlocked | Required
35 | Infusion near end | Required
36 | Infusion complete | Required
37 | Occlusion upstream | REQUIRED
38 | Occlusion downstream | REQUIRED
39 | Pressure, psi | ≤15
40 | Real-time display | REQUIRED
41 | Air in line | Required
42 | System malfunction | Required
43 | Set loaded improperly | Required
44 | Door open | Required
45 | Low battery | Required
46 | Depleted battery | Required
47 | User prompts | Required
48 | BAR CODE READER | Required
49 | AUDIBLE ALARM/ visual | Required
50 | Volume control | Required
51 | DOSE ERROR REDUCTION SYSTEM (smart technology) | Required
52 | Defaults to DERS on startup | Required
53 | Library size | >3000
54 | No. of care areas | ≥20
55 | No. of drug entities per care area | ≥100
56 | Wireless connectivity | Required
57 | Autodocumentation | Required
58 | Autoprogramming | Required
59 | Log-analysis software | Required
60 | DATA PORT | Required
61 | EVENT LOG | Required
62 | Number of events | 1 week of operation min
63 | Events stored | Key presses, program settings, alarms, volume infused, dose limit warnings
64 | POWER SOURCE
65 | Line power, VAC (Hz) | 100 TO 240 /60 HZ DUAL
66 | Battery | rechargeable batteries ,Lithium ion
67 | Life, hr @ flow (mL/hr) | ≥5 @ 25 with all channels running
68 | Recharge time, hr | <8
69 | H x W x D, cm (in) | small
70 | WEIGHT, kg (lb) | <4 kg
71 | List price | itemised price required
72 | PRIMARY INFUSION SET WITH NEEDLE FREE SYSTEM PRIMING VOLUME 18 ML, SMART SITE VALVE Y PORT, MALE LUER, DEHP-FREE LATEX-FREE | YES/ QTY: 56
73 | SECONDARY SET PRIMING VOLUME 7 ML BORE 2.7 MM HYDROPHOBIC AIR INLET FILTER (PVC) CONTAINS HANGER LATEX-FREE | YES/ QTY: 74
74 | INFUSION SET NEEDLE FREE VALVE PRIMING VOLUME 21ML SMALL BORE TUBING SEGMENT MALE LUER DEHP-FREE | YES/ QTY: 10
75 | BLOOD SET NON VENTED NEEDLE FREE VALVE PRIMING VOLUME 33ML LOW SORBING TUBING SEGMENT 2 MALE LUER LOCK DEHP-FREE | YES/ QTY: 8
76 | PRIMARY INFUSION SET PRIMING VOLUME 19 ML BORE 2.6 MM HYDROPHOBIC AIR INLET FILTER MALE LUER (PVC) LATEX-FREE | YES/ QTY: 6
77 | EXTENSION SET WITH LOW SORBING PRIMING VOLUME 1.6 ML BORE 0.9 MM FEMALE/MALE LUER (PVC) DEHP-FREE LATEX-FREE | YES/ QTY: 21
78 | EXTENSION SET WITH Y PORT PRIMING VOLUME 1.8 ML BORE 0.6 MM FEMALE/MALE LUER (PVC) LATEX-FREE | YES/ QTY: 1
79 | EXTENSION SET WITH Y PORT NEEDLE FREE SYSTEM PRIMING VOLUME 5 ML LARGE BORE 2.7 MM FEMALE/MALE LUER (PVC) LATEX-FREE | YES/ QTY: 11
80 | EXTENSION SET WITH T CONNECTOR NEEDLE FREE SYSTEM PRIMING VOLUME 0.2 ML SMALL BORE 1 MM FEMALE/MALE LUER (PVC) DEHP-FREE LATEX-FREE | YES/ QTY: 27
81 | MICRO BORE PUMP SET PRIMING VALUME 13 ML YELLOW STRIPED ROLLER CLAMPS WITH 2 MALE LUER LOCK DEHP-FREE LATEX-FREE STERILE | YES/ QTY: 1
82 | PRINMARY INFUSION SET NEEDLEFREE VALUME 22ML BORE 2.6 MM HYDROPHOBIC AIR INLET FILTER MALE LUER (PVC) LATEX-FREE | YES/ QTY: 2
Company Stamp |  |  |  | Company Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | 4218 1900 16100 | 5 | PUMP INFUSION EPIDURAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEARANCE | PREFFERED
2 | CE MARK (MDD) | REQUIRED
3 | CONFIGURATION
4 | LOCKABLE KEYPAD | REQUIRED
5 | DISPLAY | LCD,screen size data easy to read
6 | Data displayed | Dose, concentration, lockout interval, rate, patient requests, alarms
7 | PUMP CAPABILITIES
8 | Flow range | .1-999ml
9 | Increments | 0.1 mL/hr or equivalent mg or µg
10 | PUMP MECHANISM | PRESTALTIC
11 | Bolus dose | REQUIRED
12 | Increments | 0.1-25 mL
13 | Loading dose | REQUIRED
14 | Dose programmable | REQUIRED
15 | Concentrations | 0.1-100 mg/mL
16 | KVO rate, mL/hr | 2020-01-05
17 | Accuracy, % | 5
18 | Fluid resistant | REQUIRED
19 | Nurse call jack | prefferd
20 | Bar-code reader | REQUIRED
21 | SAFETY FEATURES | Lockbox, locking keypad, memory protection, tamper evident
22 | Free-flow protection | Required
23 | Lockout interval range, min | ≤5-100
24 | Accumulated dose limit | REQUIRED
25 | Controls access | Key or security code
26 | Volume capacity, mL | Any (≥30)
27 | Access | Key, lockbox
28 | SYRINGE-SIZE DETECTION | REQUIRED
29 | ALARMS & INDICATORS | REQUIRED
30 | Up/down occlusion | REQUIRED
31 | Flow error | REQUIRED
32 | Air in line | REQUIRED
33 | System malfunction | REQUIRED
34 | Empty reservoir | REQUIRED
35 | Infusion near end | REQUIRED
36 | Low battery | REQUIRED
37 | Depleted battery | REQUIRED
38 | Other | On line or battery power
39 | AUDIBLE ALARM/visual | required
40 | Volume control | required
41 | DOSE ERROR REDUCTION SYSTEM (smart technology) | Required for electronic pumps
42 | Pump defaults to DERS on startup | required
43 | Library size | >3000
44 | No. of care areas | ≥20
45 | No. of drug entities/care area | ≥100
46 | Wireless connectivity | REQUIRED
47 | Autodocumentation | REQUIRED
48 | Autoprogramming | REQUIRED
49 | Log-analysis software | REQUIRED
50 | DATA PORT | REQUIRED
51 | EVENT LOG | REQUIRED
52 | Time/date stamp | REQUIRED
53 | Number of events | ≥1200
54 | Events stored | Key presses, error codes, alarms, drug infused, amount infused, settings, boluses requested
55 | POWER SOURCE
56 | Line power, VAC (Hz) | 100 TO 240 /60 HZ DUAL
57 | Battery | rechargeable batteries ,Lithium ion
58 | Operating time, hr | ≥5 @ 10 mL/hr
59 | Recharge time, hr | <8
60 | H x W x D, cm (in) | small
61 | WEIGHT, kg (lb) | light < 4kg
62 | List price | itemised for components
Company Stamp |  |  |  | Company Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | 4218 1900 06900 | 0 | PUMP INSULINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA | REQUIRED
2 | PUMP FEATURES
3 | Bolus dose calculator | REQUIRED
4 | Bolus Speed Options | Standard: 1.5 units/minute, ◾Quick: 15 units/minute
5 | Blood glucose meter | REQUIRED
6 | Auto off | Preferred
7 | Hold option | REQUIRED
8 | Fluid resistant | Yes, IPX8 rating
9 | Usable in aircraft | Preferred
10 | DOSE INFORMATION
11 | Concentration | U-100 (fast and regular)
12 | Basal flow range, U/hr | 0 to 35 units per hour or the Max Basal Rate amount, whichever is lower
13 | Max U/day | 840
14 | Basal flow profiles | 8, ≤48 rate changes/day
15 | Bolus dose increments, U | ◾0.025, ◾0.05 and ◾0.1 units
16 | RESERVOIR
17 | Capacity, mL | 3.0 ml (300-unit)
18 | Connection | Luer
19 | DISPLAYED PARAMETERS
20 | Basal rate | Preferred
21 | Bolus dose | REQUIRED
22 | Accumulated dose | REQUIRED
23 | Backlight Display
| Type | LED
| Time out | 15 seconds (default), 30 seconds, one minute, three minutes
| Infusion time | REQUIRED
| Other | Suspend, bolus, bolus history, daily total history, complete diagnostic systems test, meter blood glucose value , real-time continuous glucose monitoring data
ALARMS & INDICATORS
| Data-entry error | REQUIRED
| Up/down occlusion | Yes Within 5 units
| Runaway infusion | Preferred
| System malfunction | REQUIRED
| Low reservoir | REQUIRED
| Empty reservoir | REQUIRED
| Infusion complete | REQUIRED
| Low battery | REQUIRED
| Depleted battery | REQUIRED
| Other | No delivery (triggered by an average of 2.4 units of missed insulin), auto off, programmed over delivery, generic reminder, blood glucose reminder, suspend before low (automatically stops insulin delivery when pre-set low limit is predicted to be reached
| AUDIBLE ALARM | REQUIRED
| Visual alarm | REQUIRED
| Volume control | REQUIRED
| Vibrate | Prefered
| DATA PORT | Kindly Specify
| EVENTS STORED | 90 days storage; 30 days viewable of 80 (alarms and alerts) and 25 (bolus and other events) per day (at maximum) for 30 days
PUMPING MECHANISM
| BATTERY | Pumps that accept disposable and rechargeable batteries preferred
| Operating time, hr | ≥72
| Rechargeable | REQUIRED
| Recharge time, hr | <3
| H x W x D, cm (in) | small
| WEIGHT, g (oz) | light
| Reservoir/syringe | REQUIRED Holds insulin and it inserts into pump. 3.0 ml (300-unit)
| battery charger | 110/220 auto 3 prong british
| Battery | AA NiMH (HR6) nickel-metal hydride rechargeable battery.
| Environmental Conditions | designed to operate between 41-104°F.
Company Stamp |  |  |  | Company Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | 4229 3500 02000 | 5 | PUMP PCA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEARANCE | PREFFERED
2 | CE MARK (MDD) | REQUIRED
3 | CONFIGURATION
4 | PCA dose button location | On pump or bolus cord
5 | DISPLAY | LCD,screen size data easy to read
6 | Data displayed | Dose, concentration, lockout interval, rate, patient requests, alarms
7 | PUMP CAPABILITIES
8 | Flow range | .1-999ml
9 | Increments | 0.1 mL/hr or equivalent mg or µg
10 | Bolus dose | REQUIRED
11 | Increments | 0.1-25 mL
12 | Loading dose | REQUIRED
13 | Dose programmable | REQUIRED
14 | Concentrations | 0.1-100 mg/mL
15 | KVO rate, mL/hr | 2020-01-05
16 | Accuracy, % | 5
17 | Fluid resistant | REQUIRED
18 | Nurse call jack | prefferd
19 | Bar-code reader | REQUIRED
20 | SAFETY FEATURES | Lockbox, locking keypad, memory protection, tamper evident
21 | Free-flow protection | Required
22 | Lockout interval range, min | ≤5-100
23 | Accumulated dose limit | REQUIRED
| Controls access | Key or security code
| Volume capacity, mL | Any (≥30)
| Access | Key, lockbox
| SYRINGE-SIZE DETECTION | REQUIRED
| ALARMS & INDICATORS | REQUIRED
| Up/down occlusion | REQUIRED
| Flow error | REQUIRED
| Air in line | REQUIRED
| System malfunction | REQUIRED
| Empty reservoir | REQUIRED
| Infusion near end | REQUIRED
| Low battery | REQUIRED
| Depleted battery | REQUIRED
| Other | On line or battery power
| AUDIBLE ALARM/visual | required
| Volume control | required
| DOSE ERROR REDUCTION SYSTEM (smart technology) | Required for electronic pumps
| Pump defaults to DERS on startup | required
| Library size | >3000
| No. of care areas | ≥20
| No. of drug entities/care area | ≥100
| Wireless connectivity | REQUIRED
| Autodocumentation | REQUIRED
| Autoprogramming | REQUIRED
| Log-analysis software | REQUIRED
| DATA PORT | REQUIRED
| EVENT LOG | REQUIRED
| Time/date stamp | REQUIRED
| Number of events | ≥1200
| Events stored | Key presses, error codes, alarms, drug infused, amount infused, settings, boluses requested
POWER SOURCE
| Line power, VAC (Hz) | 100 TO 240 /60 HZ DUAL
| Battery | rechargeable batteries ,Lithium ion
| Operating time, hr | ≥5 @ 10 mL/hr
| Recharge time, hr | <8
| H x W x D, cm (in) | small
| WEIGHT, kg (lb) | light < 4kg
| List price | itemised for components
Company Stamp |  |  |  | Company Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | 4218 1900 07000 | 5 | PUMP SYRINGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEARANCE | PREFFERED
2 | CE MARK (MDD) | REQUIRED
3 | APPLICATIONS | blood products, epidurals, general care, ICU, IV solutions
4 | TIVA-/TCI-capable | REQUIRED
5 | CONFIGURATION
6 | Number of channels | EQUAL OR MORE THAN 1
7 | Pump mechanism | Stepper motor, lead screw
8 | Pole mounting | REQUIRED / 5wheels
9 | DISPLAY | LCD,screen size data easy to read
10 | Data displayed | Alarms, event history, rate,dose,drug name,volume delivered,pressure ,
11 | PUMP CAPABILITIES
12 | SYRINGE-SIZE DETECTION | 5 TO 50 or more
13 | Flow range, mL/hr | 0.1-99.9
14 | Increments, mL | 0.1
15 | KVO rate, mL/hr | 2020-01-05
16 | Accuracy, % | 5
17 | Max pressure, psi | 20
18 | Pump-based priming | required
19 | MRI conditional | Preferred
20 | Bar-code reader | required
21 | Fluid resistant | required
22 | COMPATIBLE SYRINGES, mL | BD,MONO,TERO,OTHER
23 | BD Plastipak | 5 TO 60
| Monoject | 5 TO 60
| Terumo | 5 TO 60
| Others | preffered
| SYRINGE-SIZE DETECTION | Required
| ALARMS & INDICATORS | AUDIBLE AND VISUAL REQUIRED
| High /LOW pressure occlusion | Required
| System malfunction | Required
| SYRINGE NEAR END | Required
| Empty syringe | Required
| Plunger disengaged | Required
| Syringe unlocked | Required
| AIR IN LINE | Required
| Infusion complete | Required
| Low battery | Required
| Depleted battery | Required
| User prompts | Required
| AUDIBLE ALARM/VISUAL ALARM | Required
| Volume control | Required
| DOSE ERROR REDUCTION SYSTEM (smart technology) | Required
| Pump defaults to DERS on startup | Required
| Library size | >3000
| No. of care areas | ≥20
| No. of drug entities/care area | ≥100
| Wireless connectivity | REQUIRED
| Autodocumentation | REQUIRED
| Autoprogramming | REQUIRED
| Log-analysis software | REQUIRED
| DATA PORT | REQUIRED
| EVENT LOG | REQUIRED
| Time/date stamp | REQUIRED
| Number of events | ≥1,200
| Events stored | Key presses, error codes, alarms, rate, amount infused, program settings
| Time retained | ≥1 year
POWER SOURCE
| Line power, VAC | 100 TO 240 /60 HZ DUAL
| Battery | rechargeable batteries ,Lithium ion
| Operating time, hr | ≥5 @ 10 mL/hr
| Recharge time, hr | <10 hrs
| H x W x D, cm (in) | small
| WEIGHT, kg (lb) | light <4kg
| List price | itemised price required
Company Stamp |  |  |  | Company Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | 4218 1900 07600 | 0 | SCAN DEVICE BLADDER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Power source | Battery
2 | charger input | 110 V with American plug or dual
3 | Real-time 3D Ultrasound Bladder Scanner. | Yes
4 | SFDA | Yes
5 | One button controls the scanner start and stop. | Yes
6 | Automatic Urinary Bladder volume calculation in large digits | Yes
7 | Fully automatic multi-plane scans from one position on the body. | Yes
8 | Simple, intuitive software with Touch-Screen | Yes
9 | USB port to save on external memory device | Yes
10 | Integrated carry handle | Yes
11 | Integrated thermal printer | Yes
12 | Battery operated | Yes
13 | Calibration test tool | Yes
14 | Verification tool | Yes
15 | Easy-mode Feature | Yes
16 | Calibration License (if needed) | Yes ( Life Time)
17 | Mobile Cart /Trolley | Yes
18 | Scan Depth
19 | 10cm: Designed for Children. | Yes
20 | 16cm: Designed for Adults. | Yes
21 | 23cm: Designed for Obese. | Yes
Company Stamp |  |  |  | Company Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | 4214 0000 06500 | 0 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA Clearance | Required, Provide Certificate number
2 | SFDA Clearance | Required, Provide Certificate number
3 | CE Mark | Required, Provide Certificate number
4 | Non - Invasive blood pressure (NIBP) Specifications:
4_1 | Cuff pressure range: Automatically adjustable and monitored for Adults, Pediatrics, and Neonatal | Required
4_2 | Trend and graphs | Optional
4_3 | Data Storage / Retrieval (minimum 50 events) | Required
4_4 | BP cuffs size: Neonate, Pediatric, Adult, Large Adult, and Thigh | Required
4_5 | Default target cuff inflation: Adult/ Pediatric: 160 ± 15 mmHg | Required
4_6 | Default target cuff inflation: Neonate: 110 ± 15 mmHg | Required
4_7 | Target cuff inflation adjustment range: Automatically adjusted and sensed to maximum accuracy (Min ± 5 mmHg) | Required
4_8 | Blood pressure measurement range: Automatically adjustable for Systolic, MAP, and Diastolic readings, depending on the measurement mode | Required
4_9 | Readings Accuracy: ± 5 - 8 mmHg | Required
4_10 | Blood pressure determination time: shall be Minimum | Required
4_11 | Pulse rate determination: (Adult: 20-180 BPM ± 10%) (Neonate: 20-180 BPM ± 10%) | Required
4_12 | Pulse Accuracy: ± 3.5 % of the actual pulse rate | Required
4_13 | Over pressure cutoff (Automatically adjustable depends on the cuff size (shall not exceed 330 mmHg in adults and 165 mmHg in neonate), Audio visual alarm shall be triggered) | Required
4_14 | Adjustable alarm limits for all parameters to be measured, Audio Visual, and Volume-controlled | Required
5 | SPO2 Specifications:
5_1 | Oxygen saturation accuracy: maximum ,depends on the saturation limits | Required
5_2 | Nellcor compatible | Required
5_3 | Pulse rate range (Automatically sensed) | Required
5_4 | Saturation Pitch Indicator- pitch shall change with saturation, with selectable volume | Required
5_5 | Waveforms-pulse plethysmograph waveform (The monitor shall detect the attachment or detachment of the sensor, and also disconnect from patient within shortest possible time (5 sec. Max),audiovisual alarm shall be triggered) | Required
5_6 | Pulse detection (within 10-15 seconds of connecting to the patient, the monitor shall start detecting a pulse or enter a no signal state. Also, it shall detect loss of pulse from patient and enter a no signal state within 5-10 seconds. Audiovisual alarm shall be triggered) | Required
5_7 | Ability to use disposable / reusable probes | Required
5_8 | Barcode Scanner | Required
5_9 | Connectivity, USB, Ethernet, Wi Fi | Required, please specify
Company Stamp |  |  |  | Company Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | 4618 1810 01100 | 0 | STRETCHER EMERGENCY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safe working load: 225 Kg. or more | Required
2 | Push handles | Required
3 | Integrated transfer boards | Required
4 | Sleep deck length | Please specify
5 | Sleep deck width | Please specify
6 | Max. height (floor to deck) | Please specify
7 | Min. height (floor to deck) | Please specify
8 | Max head elevation: 90 degrees | Please specify
9 | 4 side centralized brakes/steer | Required
10 | Oxygen holder | Required
11 | Corner and side wall bumpers | Required
12 | Caster diameter: 8” | Required
13 | Trend. /Reverse Trend. ±16° or less | Required
14 | Min. 2 permanent IV poles | Required
15 | Foot side hydraulic control pedals | Required
16 | Side rails | Required
17 | Restraining straps | Required
Company Stamp |  |  |  | Company Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | 4229 3522 00100 | 0 | SUCTION UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Vacuum, MAX | 350 mmHg or more
4 | Vacuum Range | 80 – 350 mmHg
5 | Battery Run Time | Minimum 45 mins
6 | Battery Life | Three Years minimum
7 | Battery Type | Not SLA, Preferred Lithium Ion
8 | Weight | Not More than 4Kg
9 | Noise Level | Less than 60dBA
10 | Pressure Gauge | Analog or Digital
11 | Audible alarm: | Low Battery, Occlusion Pressure
12 | Power requirements | 100-240V; 50/60Hz
13 | Air Flow | 25LPM or more
14 | Canister Capacity | 800ml to 1200ml
15 | IPX7 | yes
Company Stamp |  |  |  | Company Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | 4218 1900 08100 | 0 | TRANSILLUMINATOR VEIN PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MODES:
1_1 | Universal | Yes
1_2 | Fine Detail | Yes
1_3 | Inverse | Yes
1_4 | Re-Size | Yes
2 | Active Vein Imaging Navigation (AVIN) Capability | Yes
3 | Non-Invasive Procedure | Yes
4 | No-Heat Generation | Yes
5 | No-Cross Contamination | Yes
6 | No-Disposable & Consumable parts required | Yes
7 | Range of Use | All Ages ( Pre-mature to Adult )
8 | Skin Pigmentation | All Skin Tone
9 | Completely Hands Free for the Caregiver with Eye On Patient (EOP) | Yes
10 | Certifications | FDA, CE & SFDA Approved
Company Stamp |  |  |  | Company Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | 4229 5100 10700 | 0 | VASCULAR DOPPLER HAND HELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards | EN60601, IEC 60601
2 | Device Classification | Class II Equipment, Type BF
3 | Battery Operated | User replaceable (Doppler Only)
4 | Frequency | 2 MHZ (Doppler Only)
5 | Transducers | Two Crystals Narrow Beam (Doppler Only)
6 | Audio Response | 300 Hz to 1 KHZ (Doppler Only)
7 | Fetal Heart Rate | 50 to 200 BPM (Doppler Only)
8 | HR Resolution | 1 BPM (Doppler Only)
9 | HR Accuracy | ±2 BPM (Doppler Only)
10 | IPX7 | YES
Company Stamp |  |  |  | Company Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | 4227 2208 01000 | 0 | VENTILATOR BIPAP HOME CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Pediatric to Adult (≥20 kg)
3 | CONTROL PARAMETER | dial control/push button, or similar
4 | Inspiration Time, Sec | 0.5 - 5
5 | Pressure Range, cmH2O | 0-25
6 | Expiration Time, Sec | 0.5 - 5
7 | Ramp Feature | At least 5 min
8 | Starting ramp | At least 4 cmH2O
9 | Trigger Sensitivity | high sensitivity
10 | Panel Lock | yes
11 | Rise Profile | yes, specify
12 | Humidification | yes, specify
13 | Data storage capacity | yes, min last 6 months
14 | APNEA alarm | preferred, specify
15 | EQUIPMENT ALARMS | yes, specify
16 | Mask Disconnection | yes, specify
17 | Power Failure | yes, specify
18 | Device Inoperative | yes, specify
19 | Overheat | yes, specify
20 | MONITORING PARAMETER
21 | Inspiration Time | preferred, specify
22 | Expiration Time | preferred, specify
23 | Inspiration Pressure | preferred, specify
24 | Expiration Pressure | preferred, specify
25 | Breath Rate | preferred, specify
26 | I:E Ratio | preferred, specify
27 | Mean Airway Pressure MAP | preferred, specify
28 | Minute Volume | preferred, specify
29 | Tidal Volume | preferred, specify
30 | Leak Percentage | preferred, specify
31 | Flow | preferred, specify
32 | PATIENT ALARMS | preferred, specify
33 | High Respiration Rate | preferred, specify
34 | APNEA | preferred, specify
35 | High Pressure Alarm | preferred, specify
36 | High/Low Tidal Volume (Auto mode only) | preferred, specify
37 | Rebreathing | preferred, specify
38 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION | preferred, specify
39 | Acceptable Noise ≤40dB | yes, specify
40 | Altitude compensation | yes
41 | Temperature : 15 to 45 C
42 | Dimension | yes, specify
43 | ELECTRICAL SPECIFICATION
44 | DEVICE must have dual voltage operational compatibility at automatic selection of frequency.
45 | Voltage Range : 100 ~ 240 Volts
46 | Frequency : 50Hz / 60Hz
47 | Maximum Current : ≤ 3 Amp
48 | VENTILATION MODES
49 | CPAP
50 | BiLEVEL
51 | SAFETY STANDARD
52 | The system must compliance with the following safety standards
53 | - ISO 80601-1-12
54 | - ISO 80601-1-55
55 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
56 | - IPX1
57 | OTHERS
58 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
59 | - The equipment or any of its accessories must not be presently listed in any recall
60 | - Vendor Should provide any cable for interfacing to laptop for calibration and maintenance FOC
Company Stamp |  |  |  | Company Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | 4227 2208 01100 | 0 | VENTILATOR CPAP AUTO HOME CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Pediatric to Adult (≥20 kg)
2 | CONTROL PARAMETER | dial control/push button, or similar
3 | Pressure Range, cmH2O | 0-25
4 | Expiration Time, Sec | 0.5 - 5
5 | Ramp Feature | At least 5 min
6 | Starting ramp | At least 4 cmH2O
7 | Trigger Sensitivity | high sensitivity
8 | Panel Lock | yes
9 | Rise Profile | yes, specify
10 | Humidification | yes, specify
11 | Data storage capacity | yes, min last 6 months
12 | APNEA alarm | preferred, specify
13 | EQUIPMENT ALARMS | yes, specify
14 | Mask Disconnection | yes, specify
15 | Power Failure | yes, specify
16 | Device Inoperative | yes, specify
17 | Overheat | yes, specify
18 | MONITORING PARAMETER
19 | Inspiration Pressure | yes, specify
20 | Leak Percentage | yes, specify
21 | Flow | yes, specify
22 | PATIENT ALARMS
23 | High Pressure Alarm | Preferred
24 | Low Minute Volume | Preferred
25 | High/Low Tidal Volume | Preferred
26 | Rebreathing | Preferred
27 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
28 | Acceptable Noise ≤50dB
29 | Altitude compensation | yes
30 | Temperature : 15 to 45 C
31 | Dimension | Specify
32 | ELECTRICAL SPECIFICATION
33 | DEVICE must have dual voltage operational compatibility at automatic selection of frequency.
34 | Voltage Range : 100 ~ 240 Volts
35 | Frequency : 50Hz / 60Hz
36 | Maximum Current : ≤ 3 Amp
37 | VENTILATION MODES
38 | CPAP AUTO
39 | CPAP
40 | SAFETY STANDARD
41 | The system must compliance with the following safety standards
42 | - ISO 80601-1-12
43 | - ISO 80601-1-55
44 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
45 | - IPX1 | Preferred
46 | OTHERS
47 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
48 | - The equipment or any of its accessories must not be presently listed in any recall
49 | - Vendor Should provide any cable for interfacing to laptop for calibration and maintenance FOC
Company Stamp |  |  |  | Company Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | 4219 2210 03400 | 0 | WHEELCHAIR LIGHT WEIGHT SIZE 20 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | set belt | Yes
2 | IV stand (fixed) | Yes
3 | Hospital logo (fixed) | Yes
4 | foot rests (adjustable) | Yes
Company Stamp |  |  |  | Company Signature