## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | ALARM & MONITOR RADIATION AREA DIGITAL
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | ANALYZER CO2
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | CABINET AUTOMATED MEDICATION MOBILE
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | CABINET DRYING & STORAGE FLEXIBLE ENDOSCOPE
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | CABINET DRYING ENDOSCOPY
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | CABINET UNIT DOSE PET LEAD LINED
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | CALIBRATOR DOSE 55T PET
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | COMPRESSOR AIR HOME CARE
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | CONCENTRATOR OXYGEN ADULT 10L
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | CONCENTRATOR OXYGEN ADULT 5L
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | DEFIBRILLATOR UNIT
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | DISINFECTOR ENDOSCOPY
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | ECG 12 LEAD WITH CART
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | ELECTROSURGERY WITH ARGON & VESSEL SEALING
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | ENDOSCOPY WORKSTATION
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | FLOWMETER AIR
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | HUMIDIFIER
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | IRRIGATION SYSTEM SURGCAL
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | LIFT PATIENT BARIATRIC
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | LIFT PATIENT MOBILE
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | METER SURVEY RADIATION
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | MONITOR PHYSIOLOGICAL TRANSPORT
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | MONITOR RADIATION SYSTEM EXTERNAL PROBE
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | NEBULIZER AIR
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | OXIMETER PULSE FINGER TIP
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | PUMP DUTY CHEMICAL
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | PUMP INFUSION
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | PUMP INFUSION EPIDURAL
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | PUMP INSULINE
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | PUMP PCA
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | PUMP SYRINGE
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | SCAN DEVICE BLADDER
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | STRETCHER EMERGENCY
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | SUCTION UNIT PORTABLE
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | TRANSILLUMINATOR VEIN PORTABLE
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | VASCULAR DOPPLER HAND HELD
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | VENTILATOR BIPAP AUTO HOME CARE
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | VENTILATOR CPAP AUTO HOME CARE
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | WHEELCHAIR LIGHT WEIGHT SIZE 20

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4227170200100 | OXYGEN CONCENTRATORS WITH ADULT FLOWMETER |  | MEDICAL GROUP | 100 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | MEDICAL GROUP | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4227170200100 | 0 | OXYGEN CONCENTRATORS WITH ADULT FLOWMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The unit should be Oxygen Concentrators to produce an oxygen-rich gas mixture by drawing in room air and extracting nitrogen.
2 | The unit should be designed to provide long-term oxygen therapy to patients at home and hospitals to deliver low-flow and low-pressure oxygen.
3 | The unit should be mobile with built in casters.
4 | The unit should have molded top handle.
5 | The concentrator should deliver an oxygen concentration of at least FiO2 of 80% at flow rate of 10 liter and at least 90% at a flow of 8 liter
6 | The flow setting should be up to 8-10 L/min.
7 | Outlet pressure at least 8 PSI (55 kPa).
8 | The device must have an Oxygen concentration status indicator (OCSI) to alert the user if the unit’s product gas has a low oxygen concentration; the OCSI should warn of concentrations below 80% and should have both an audible alarm and a visual indicator.
9 | The OCSI should be a built-in component.
10 | Alarms should be initiated for high and low pressures, battery (if available)- or line-power failures, low oxygen, and occlusions.
11 | The unit should be provided with a detailed service and operation manual at the delivery and installation.
12 | The unit should allow for freedom of movement.
13 | The unit should be capable to run on long oxygen tubing.
14 | The humidifier bottle should be included with tubing set.
15 | The concentrator should be light enough to permit a patient to move it with approximate weight 25Kg.
16 | The gross particle or bacteria filter should be included.
17 | The unit should have dust filter.
18 | The filter should be easy to be replaced and cleaned.
19 | The unit should have working hour meter.
20 | The unit should be user friendly and easy to use, operate, and maintain.
21 | The amount of noise produced by the concentrator should not exceed 60 dBA, and a noise level below 40 dBA is preferred.
22 | The Unit should have no sharp edges.
23 | The unit should be well constructed with durable material to withstand typical abuse & cleaning.
24 | The unit should be easy to clean and disinfect.
25 | The vendor should provide demonstration unit for at least two (2) weeks for evaluation.
26 | Switches, knobs, and other controls should be designed for conditions of heavy use.
27 | The controls should be visible and clearly identified, and their functions should be self-evident.
28 | Controls should be sealed against penetration of liquids.
29 | The unit should meet electrical safety standards.
30 | The unit shall comply with international respiratory standards.
31 | The unit should be Saudi FDA approved or CE marked.
32 | The vendor should provide plan for the routine inspection and preventive maintenance (e.g., filter replacement).
33 | The vendor shall provide itemized prices for all consumables and accessories not listed in the technical specifications and required to run the machine.
34 | Should include extra dust/and preferred bacterial filters.
35 | Should have extra humidifier tubings.
36 | The unit shall be single phase, 220VAC / 60Hz, British plug, cable at least 3 miters.
37 | All available options, consumables, or disposables deemed necessary for the intended function of the equipment and not listed in the Technical Specifications should be included and specified
Company Stamp |  |  |  | Company Signature