## Material code 139
Product ID | Customer ID | Attribute Name | Domain Name | Technical Parameteres | Specified | Mandatory | Comment Allowed
4219210301100 | 10005 | 4219210301100_A | BBP_YESNO | Description of Function: | YES | X | X
4219210301100 | 10005 | 4219210301100_A-1 | BBP_YESNO | Multi-functioanl, (mydruatic, & non-mydriatic) retinal camera to capture the fundus of the eye for glaucoma petients | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_B | BBP_YESNO | Technical Specifications: | YES | X | X
4219210301100 | 10005 | 4219210301100_B-1 | BBP_YESNO | Application | mydruatic, & non-mydriatic fundus camera | X | X
4219210301100 | 10005 | 4219210301100_B-2 | BBP_YESNO | Camera resoluation | 16 megapixel approximately | X | X
4219210301100 | 10005 | 4219210301100_B-3 | BBP_YESNO | Angle of view | wide with approximate to 44° | X | X
4219210301100 | 10005 | 4219210301100_B-4 | BBP_YESNO | Working distance | Approximate to 41 mm | X | X
4219210301100 | 10005 | 4219210301100_B-5 | BBP_YESNO | Pupil diameter for photography | 3.1 mm to 4.1 mm or more | X | X
4219210301100 | 10005 | 4219210301100_B-6 | BBP_YESNO | Photography type | color, red-free, and fluorescein angiography (FA) | X | X
4219210301100 | 10005 | 4219210301100_B-7 | BBP_YESNO | Should have the automated features that includes auto focus, auto shoot and auto pupil detection | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_B-8 | BBP_YESNO | Should have stereo photography feature to allow of capturing pair images from different positions | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_B-9 | BBP_YESNO | Correction range | -14 to 13 diopter (without lens), -11 to -34 diopter (with minus lens), & 8 to 41 diopter ( plus lens) | X | X
4219210301100 | 10005 | 4219210301100_B-10 | BBP_YESNO | Should have internal and external fixation target selection option for aligning the visual axis of the eye | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_B-11 | BBP_YESNO | Base movment | back-forth 47 mm, right-left 102 mm, & up-down 31 mm approximately | X | X
4219210301100 | 10005 | 4219210301100_B-12 | BBP_YESNO | Chinrest movment | 70 mm approximately | X | X
4219210301100 | 10005 | 4219210301100_B-13 | BBP_YESNO | Provide the system with high specs workstation(top of the range system with most reliable brand value at the time of supply, having at least following features- i7 or higher 8-Core processor, 1 TB SSD hard disk for data storeg and 1 TB SSD hard disk for operationg system and the softwares, 16 GB RAM ore more, with >18 inch LED monitor) that must be compatible with the installed software and applications | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_B-14 | BBP_YESNO | Lifetime license for acquiring, displaying, enhancing, analyzing and saving images | Multi-image management functions that facilitates image acquisition, enhancement, measurement and comparison | X | X
4219210301100 | 10005 | 4219210301100_C | BBP_YESNO | System Configuration Accessories, Spares and Consumables: | - | X | X
4219210301100 | 10005 | 4219210301100_C-1 | BBP_YESNO | All accessories, tubes, & tools that required operating the device should be provided. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_D | BBP_YESNO | Environmental Factors: | - | X | X
4219210301100 | 10005 | 4219210301100_D-1 | BBP_YESNO | The operating temperature should be approximately within the range of 5 ° C to 45 ° C. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_E | BBP_YESNO | Power Supply: | - | X | X
4219210301100 | 10005 | 4219210301100_E-1 | BBP_YESNO | The voltage of the unit must be 220 VAC, frequency must be 60 Hz. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_E-2 | BBP_YESNO | The plug of power cord should be in UK standard. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_F | BBP_YESNO | Standards, Safety and Training: | - | X | X
4219210301100 | 10005 | 4219210301100_F-1 | BBP_YESNO | Manufacturer level training for Biomedcal Engineers and End-users by certified trainer. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_G | BBP_YESNO | Documentations: | - | X | X
4219210301100 | 10005 | 4219210301100_G-1 | BBP_YESNO | Service Manual, Operation Manual, Spare Parts List, Checklist for Inspection and Preventive Maintenance and list of special tools required for maintenance and calibration. | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_H | BBP_YESNO | Warranty: | - | X | X
4219210301100 | 10005 | 4219210301100_H-1 | BBP_YESNO | Detailed Warranty Certificates including period of warranty and the coverage for labor and spare parts for at least 3 years | REQUIRED | X | X
4219210301100 | 10005 | 4219210301100_I | BBP_YESNO | Pre-Installation Requirments: | - | X | X
4219210301100 | 10005 | 4219210301100_I-1 | BBP_YESNO | All the required civil works and pre-installation should be clearly mentioned and included (if required) | REQUIRED | X | X